## Cover Sheet
Analytical Cover Sheet
QAPP: | In Vitro TK Assessments of PFAS (D-CED-0031343)
Analysis Set ID: | 2020_PFAS_PPB_UC_ALCOHOLS_A_ALK | *Ultracentrifugation assay to assess plasma protein binding with PFAS acids
Analytes: |  | PFAS
Dates Prepared: | 1/14/2020, 1/29/2020, 2/19/2020 | PPB_CC
Assay Preparer: | A. Kreutz
Dates Prepared: | 1/17/2020, 1/29/2020, 2/13/2020 (899 & 273 repeats; 273 was run but no QC) | PPB_UC
Dates of MS Run, resp.: | 2/3/2020, 2/19/2020, 3/4/2020
Assay Preparer: | A. Kreutz
Report Preparer: | A. Kreutz
Sample Type: | Standards, QCs, Unknown Analyte Samples | CC1-CC15
Internal Standard (conc) | 2-Perfluorobutyl-[1,1,2,2-2H4]-ethanol ; 2-Perfluorobutyl-[1,1,2,2-2H4]-ethanol; 2-Perfluorooctyl-[1,1-2H2]-[1,2-13C2]-ethanol (8:2) [lot ] (3 pg/uL); 13C6-4-Nitrotoluene [lot SDFK-011] (3 pg/uL) | Alcohols, Amides; Ref
Total # Samples (not including blanks): | 219
Actual # Samples (not including replicates): | 214
LOD Calculation | Minimum measured concentration of a substance that can be reported with 99% confidence that the measured concentration is distinguishable from method blank results
eLOQ Definition | Lowest curve concentration at which precision checks can be quantitated against curve ±30%
QC Samples and Acceptance Criteria:
Note: one of the three instrument blanks should be run in 1 of every six injections
QC Type | Description | Acceptance Criteria | Type
Calibration Curve Linearity | 7-point minimum | >0.98 | Linearity
CC Point Accuracy | Every curve sample | 80-120% of spike | Accuracy
Curve Precision Check | 2-3 each run | 80-120% | Accuracy
Acetonitrile Blank (ACN) | minimum of 3 each run | eLOQ | Specificity
Matrix Blank | 5 each run | NA (for LOD calc) | LOD
Crash reagent blank (FA/ACN) | minimum of 2 each run | eLOQ | Specificity
Replicate Assessment | Every Sample | 75-125% of each other | Accuracy
Internal Standard | Every Sample | w/in 3 StDev of relevant IS mean | Accuracy
Assay Reference Chemical | 1-2 each run | RSD ± +/-20% | Reproducibility/Precision
Analyte-IS Matching
Analyte | Name | Sample ID | IS | Reference Compound | Avg. MW
DTXSID10379991 | 3-(Perfluorooctyl)propanol | 267 | MFOET | 4-nitrotoluene | 478.149
DTXSID0059871 | Pentafluoropropionamide | 273 | MFOET |  | 163.047
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | 476 | MFHET |  | 299.264
DTXSID9059832 | Dodecafluoroheptanol | 906 | MFOET |  | 332.089
DTXSID80382093 | 2-Aminohexafluoropropan-2-ol | 899 | MFBET |  | 183.0119
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | 900 | MFOET |  | 232.0134
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | 913 | MFOET |  | 294.117
DTXSID4059914 | Heptafluorobutanol | 915 | MFBET |  | 200.056
DTXSID60379269 | 3-(Perfluoropropyl)propanol | 965 | MFHET |  | 228.11
DTXSID70381151 | Perfluorooctanamidine | 3125 | MFOET |  | 412.102
Sample ID Key
CC | Calibration Curve
QC | Quality Check
UC_S#letterAF | Ultracentrifugation Assay_Mix Replicate Aqueous Fraction
UC_S#letterT# | Ultracentrifugation Assay_Mix Replicate Stability Timepoint
Review of Data Timeline
Person | Action
ALK | Assays conducted 1/17, 1/29, 2/13/2020
MSC | Analytical data generated 2/13, 2/19, 3/4/2020
ALK | QC Report completed 9/9/2020
Additional Notes
Concentrations listed are at the instrument, unless otherwise specified

## Executive Summary
Executive Summary
|  | Quantitative Limits |  |  | Plasma Protein Binding
DTXSID | Name | Sample ID | LOD (nM) | eLOQ (nM) | 3-(Perfluorooctyl)propanol |  |  |  |  | Pentafluoropropionamide
DTXSID10379991 | 3-(Perfluorooctyl)propanol | 267 | 0.72 | 7.5 | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
DTXSID0059871 | Pentafluoropropionamide | 273 | 2.7 | 5 | AF | 0.35786 | 2020-01-29 | 0.035748 |  | AF | 5.5773 | 2020-01-29 | 0.80894
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | 476 | 0.57 | 1.75 | [8 empty] | 0.60148
DTXSID80382093 | 2-Aminohexafluoropropan-2-ol | 899 | 1.29 | 1.75 | [8 empty] | 0.45042
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | 900 | 0.53 | 1.75 | T1h | 11.367 |  |  |  | T1h | 5.5773
DTXSID9059832 | Dodecafluoroheptanol | 906 | 0.35 | 12.5
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | 913 | 2.14 | 1.75
DTXSID4059914 | Heptafluorobutanol | 915 | 1.92 | 1.75 | T5h | 10.1621 |  |  |  | T5h | 9.7147
DTXSID60379269 | 3-(Perfluoropropyl)propanol | 965 | 1.73 | 1.75
DTXSID70381151 | Perfluorooctanamidine | 3125 | 0.53 | 3
DTXSID5023792 | 4-nitrotoluene | 4-NT | 0.19 | 5
[5 empty] | N-Methyl-N-trimethylsilylheptafluorobutyramide |  |  |  |  | 2-Aminohexafluoropropan-2-ol
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 5.5111 | 2020-01-29 | 0.54507 |  | AF | 7.6588 | 2010-01-17 | 1.1507 | 1.0718
[8 empty] | 0.60092 |  |  |  | 2020-02-13 | 0.99293
NA
[5 empty] | T1h | 10.8327 |  |  |  | T1h | 9.8202
[5 empty] | T5h | 10.1294 |  |  |  | T5h | 7.4386
[5 empty] | 1H,1H,5H-Perfluoropentanol |  |  |  |  | Dodecafluoroheptanol
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 3.2026 | 2020-01-17 | 0.3616 |  | AF | 0.44228 | 2020-01-29 | 0.053683
[8 empty] | 0.66804 |  |  |  |  | NA
[8 empty] | 0.87014 |  |  |  |  | 0.02975
[5 empty] | T1h | 9.0715 |  |  |  | T1h | 9.5728
[5 empty] | T5h | 9.0383 |  |  |  | T5h | 8.3321
[5 empty] | 3-(Perfluoro-2-butyl)propane-1,2-diol |  |  |  |  | Heptafluorobutanol
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 3.1919 | 2020-01-29 | 0.36959 |  | AF | 3.203 | 2020-01-29 | 0.51051 | 0.17017
[8 empty] | NA |  |  |  |  | 0
[8 empty] | 0 |  |  |  |  | 0
[5 empty] | T1h | 9.0312 |  |  |  | T1h | 7.7211
[5 empty] | T5h | 8.6876 |  |  |  | T5h | 6.2845
[5 empty] | 3-(Perfluoropropyl)propanol |  |  |  |  | Perfluorooctanamidine
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 1.5116 | 2020-01-29 | 0.2033 |  | AF | 0.13206 | 2020-01-29 | 0.025258
0
0
[5 empty] | T1h | 8.3808 |  |  |  | T1h | 7.7654
[5 empty] | T5h | 7.4565 |  |  |  | T5h | 5.2322
4-nitrotoluene
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 0.46373 | 2020-01-29 | 0.060189 | 0.063122
[7 empty] | 2020-01-29 | 0.053392
[7 empty] | 2020-01-17 | 0.060696
[7 empty] | 2020-02-13 | 0.078211
[5 empty] | T1h | 9.136
[5 empty] | T5h | 7.4192
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: CC1 w/in 30%, but CC2 wasn't

## Sample ID
Namearray | Comment | Data File | Type | Level | Conc. pre-Crash (nM) | Conc. after Crash (nM) | Acq. Date-Time | Total # Samples: | 219
ACN | PFAS Group 2 UC | 7010_1_02042000.D | Blank |  |  |  | 2020-02-04 15:48 | Actual # Samples: | 214
SpikeBlank | PFAS Group 2 UC | 7010_1_02042001.D | MatrixBlank |  |  |  | 2020-02-04 16:09
SpikeBlank | PFAS Group 2 UC | 7010_1_02042002.D | MatrixBlank |  |  |  | 2020-02-04 16:30
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042003.D | Cal | 1 | 7 | 1.75 | 2020-02-04 16:51
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042004.D | Cal | 2 | 12 | 3 | 2020-02-04 17:13
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042005.D | Cal | 3 | 20 | 5 | 2020-02-04 17:34
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042006.D | Cal | 4 | 30 | 7.5 | 2020-02-04 17:55
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042007.D | Cal | 5 | 50 | 12.5 | 2020-02-04 18:17
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042008.D | Cal | 6 | 80 | 20 | 2020-02-04 18:38
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042009.D | Cal | 7 | 125 | 31.25 | 2020-02-04 18:59
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042010.D | Cal | 8 | 200 | 50 | 2020-02-04 19:20
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042011.D | Cal | 9 | 350 | 87.5 | 2020-02-04 19:42
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042012.D | Cal | 10 | 500 | 125 | 2020-02-04 20:03
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042013.D | Cal | 11 | 800 | 200 | 2020-02-04 20:24
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042014.D | Cal | 12 | 1500 | 375 | 2020-02-04 20:46
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042015.D | Cal | 13 | 2500 | 625 | 2020-02-04 21:07
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042016.D | Cal | 14 | 3500 | 875 | 2020-02-04 21:28
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042017.D | Cal | 15 | 5000 | 1250 | 2020-02-04 21:49
SpikeBlank | PFAS Group 2 UC | 7010_1_02042018.D | MatrixBlank |  |  |  | 2020-02-04 22:11
UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  |  |  | 2020-02-04 22:32
UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  |  |  | 2020-02-04 22:53
UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  |  |  | 2020-02-04 23:15
SpikeBlank | PFAS Group 2 UC | 7010_1_02042022.D | MatrixBlank |  |  |  | 2020-02-04 23:36
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042023.D | QC | 3 | 20 | 5 | 2020-02-04 23:57
UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  |  |  | 2020-02-05 00:19
UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  |  |  | 2020-02-05 00:40
UCGCAAT1hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042026.D | Sample |  |  |  | 2020-02-05 01:01
UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  |  |  | 2020-02-05 01:23
UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  |  |  | 2020-02-05 01:44
UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  |  |  | 2020-02-05 02:05
UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  |  |  | 2020-02-05 02:27
ACN | PFAS Group 2 UC | 7010_1_02042031.D | Blank |  |  |  | 2020-02-05 02:48
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042032.D | QC | 9 | 350 | 87.5 | 2020-02-05 03:09
QCUCGCAACC4_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042033.D | QC | 4 | 30 | 7.5 | 2020-02-05 03:30
QCUCGCAACC8_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042034.D | QC | 8 | 200 | 50 | 2020-02-05 03:52
QCUCGCAACC11_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042035.D | QC | 11 | 800 | 200 | 2020-02-05 04:13
UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  |  |  | 2020-02-05 04:35
UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  |  |  | 2020-02-05 04:56
UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  |  |  | 2020-02-05 05:17
UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  |  |  | 2020-02-05 05:39
UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  |  |  | 2020-02-05 06:00
UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  |  |  | 2020-02-05 06:21
SpikeBlank | PFAS Group 2 UC | 7010_1_02042042.D | MatrixBlank |  |  |  | 2020-02-05 06:42
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042043.D | QC | 13 | 2500 | 625 | 2020-02-05 07:04
UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  |  |  | 2020-02-05 07:25
UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  |  |  | 2020-02-05 07:47
UCGCAAT5hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042046.D | Sample |  |  |  | 2020-02-05 08:08
UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  |  |  | 2020-02-05 08:29
UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  |  |  | 2020-02-05 08:51
UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  |  |  | 2020-02-05 09:12
SpikeBlank | PFAS Group 2 UC | 7010_1_02042050.D | MatrixBlank |  |  |  | 2020-02-05 09:33
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042051.D | QC | 6 | 80 | 20 | 2020-02-05 09:55
UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  |  |  | 2020-02-05 10:16
UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  |  |  | 2020-02-05 10:37
UCGCAAT1hS1C_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042054.D | Sample |  |  |  | 2020-02-05 10:59
UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  |  |  | 2020-02-05 11:20
UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  |  |  | 2020-02-05 11:41
UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  |  |  | 2020-02-05 12:03
UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  |  |  | 2020-02-05 12:24
UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  |  |  | 2020-02-05 12:45
ACN | PFAS Group 2 UC | 7010_1_02042060.D | Blank |  |  |  | 2020-02-05 13:07
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042064.D | QC | 1 | 7 | 1.75 | 2020-02-05 14:32
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042065.D | QC | 2 | 12 | 3 | 2020-02-05 14:53
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042066.D | QC | 3 | 20 | 5 | 2020-02-05 15:15
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042067.D | QC | 4 | 30 | 7.5 | 2020-02-05 15:36
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042068.D | QC | 5 | 50 | 12.5 | 2020-02-05 15:58
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042069.D | QC | 6 | 80 | 20 | 2020-02-05 16:19
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042070.D | QC | 7 | 125 | 31.25 | 2020-02-05 16:40
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042071.D | QC | 8 | 200 | 50 | 2020-02-05 17:02
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042072.D | QC | 9 | 350 | 87.5 | 2020-02-05 17:23
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042073.D | QC | 10 | 500 | 125 | 2020-02-05 17:45
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042074.D | QC | 11 | 800 | 200 | 2020-02-05 18:06
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042075.D | QC | 12 | 1500 | 375 | 2020-02-05 18:27
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042076.D | QC | 13 | 2500 | 625 | 2020-02-05 18:49
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042077.D | QC | 14 | 3500 | 875 | 2020-02-05 19:10
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042078.D | QC | 15 | 5000 | 1250 | 2020-02-05 19:31
SpikeBlank | PFAS Group 2 UC | 7010_1_02042079.D | MatrixBlank |  |  |  | 2020-02-05 19:53
ACN | ACN | 7010_2_03032000.D | Blank |  |  |  | 2020-03-03 10:01
SpikeBlank | SpikeBlank | 7010_2_03032001.D | MatrixBlank |  |  |  | 2020-03-03 10:22
SpikeBlank | SpikeBlank | 7010_2_03032002.D | MatrixBlank |  |  |  | 2020-03-03 10:42
UCGCAACC1_1/29/20_GC 2/3 | UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032003.D | Cal | 1 | 7 | 1.75 | 2020-03-03 11:03
UCGCAACC2_1/29/20_GC 2/3 | UCGCAACC2_1/29/20_GC 2/3 | 7010_2_03032004.D | Cal | 2 | 12 | 3 | 2020-03-03 11:24
UCGCAACC3_1/29/20_GC 2/3 | UCGCAACC3_1/29/20_GC 2/3 | 7010_2_03032005.D | Cal | 3 | 20 | 5 | 2020-03-03 11:44
UCGCAACC4_1/29/20_GC 2/3 | UCGCAACC4_1/29/20_GC 2/3 | 7010_2_03032006.D | Cal | 4 | 30 | 7.5 | 2020-03-03 12:05
UCGCAACC5_1/29/20_GC 2/3 | UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032007.D | Cal | 5 | 50 | 12.5 | 2020-03-03 12:26
UCGCAACC6_1/29/20_GC 2/3 | UCGCAACC6_1/29/20_GC 2/3 | 7010_2_03032008.D | Cal | 6 | 80 | 20 | 2020-03-03 12:46
UCGCAACC7_1/29/20_GC 2/3 | UCGCAACC7_1/29/20_GC 2/3 | 7010_2_03032009.D | Cal | 7 | 125 | 31.25 | 2020-03-03 13:07
UCGCAACC8_1/29/20_GC 2/3 | UCGCAACC8_1/29/20_GC 2/3 | 7010_2_03032010.D | Cal | 8 | 200 | 50 | 2020-03-03 13:27
UCGCAACC9_1/29/20_GC 2/3 | UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032011.D | Cal | 9 | 350 | 87.5 | 2020-03-03 13:48
UCGCAACC10_1/29/20_GC 2/3 | UCGCAACC10_1/29/20_GC 2/3 | 7010_2_03032012.D | Cal | 10 | 500 | 125 | 2020-03-03 14:08
UCGCAACC11_1/29/20_GC 2/3 | UCGCAACC11_1/29/20_GC 2/3 | 7010_2_03032013.D | Cal | 11 | 800 | 200 | 2020-03-03 14:29
UCGCAACC12_1/29/20_GC 2/3 | UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032014.D | Cal | 12 | 1500 | 375 | 2020-03-03 14:49
UCGCAACC13_1/29/20_GC 2/3 | UCGCAACC13_1/29/20_GC 2/3 | 7010_2_03032015.D | Cal | 13 | 2500 | 625 | 2020-03-03 15:10
UCGCAACC14_1/29/20_GC 2/3 | UCGCAACC14_1/29/20_GC 2/3 | 7010_2_03032016.D | Cal | 14 | 3500 | 875 | 2020-03-03 15:31
UCGCAACC15_1/29/20_GC 2/3 | UCGCAACC15_1/29/20_GC 2/3 | 7010_2_03032017.D | Cal | 15 | 5000 | 1250 | 2020-03-03 15:51
QCUCGCAACC4_1/30/20_GC 2/3 | QCUCGCAACC4_1/30/20_GC 2/3 | 7010_2_03032018.D | QC | 4 | 30 | 7.5 | 2020-03-03 16:12
QCUCGCAACC8_1/30/20_GC 2/3 | QCUCGCAACC8_1/30/20_GC 2/3 | 7010_2_03032019.D | QC | 8 | 200 | 50 | 2020-03-03 16:32
QCUCGCAACC11_1/30/20_GC 2/3 | QCUCGCAACC11_1/30/20_GC 2/3 | 7010_2_03032020.D | QC | 11 | 800 | 200 | 2020-03-03 16:53
SpikeBlank | SpikeBlank | 7010_2_03032021.D | MatrixBlank |  |  |  | 2020-03-03 17:13
UCGCAAT1hS1A_GC 2/3 | UCGCAAT1hS1A_GC 2/3 | 7010_2_03032022.D | Sample |  |  |  | 2020-03-03 17:34
UCGCAAT1hS1B_GC 2/3 | UCGCAAT1hS1B_GC 2/3 | 7010_2_03032023.D | Sample |  |  |  | 2020-03-03 17:54
UCGCAAT1hS2B_GC 2/3 | UCGCAAT1hS2B_GC 2/3 | 7010_2_03032024.D | Sample |  |  |  | 2020-03-03 18:15
UCGCAAT1hS2C_GC 2/3 | UCGCAAT1hS2C_GC 2/3 | 7010_2_03032025.D | Sample |  |  |  | 2020-03-03 18:36
UCGCAAT1hS3A_GC 2/3 | UCGCAAT1hS3A_GC 2/3 | 7010_2_03032026.D | Sample |  |  |  | 2020-03-03 18:56
UCGCAAT1hS3B_GC 2/3 | UCGCAAT1hS3B_GC 2/3 | 7010_2_03032027.D | Sample |  |  |  | 2020-03-03 19:17
ACN | ACN | 7010_2_03032028.D | Blank |  |  |  | 2020-03-03 19:37
UCGCAAT5hS1B_GC 2/3 | UCGCAAT5hS1B_GC 2/3 | 7010_2_03032029.D | Sample |  |  |  | 2020-03-03 19:58
UCGCAAT5hS1C_GC 2/3 | UCGCAAT5hS1C_GC 2/3 | 7010_2_03032030.D | Sample |  |  |  | 2020-03-03 20:19
UCGCAACC5_1/29/20_GC 2/3 | UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032031.D | QC | 5 | 50 | 12.5 | 2020-03-03 20:39
UCGCAAT5hS3B_GC 2/3 | UCGCAAT5hS3B_GC 2/3 | 7010_2_03032032.D | Sample |  |  |  | 2020-03-03 21:00
UCGCAAT5hS3C_GC 2/3 | UCGCAAT5hS3C_GC 2/3 | 7010_2_03032033.D | Sample |  |  |  | 2020-03-03 21:20
UCGCAAT5hS2A_GC 2/3 | UCGCAAT5hS2A_GC 2/3 | 7010_2_03032034.D | Sample |  |  |  | 2020-03-03 21:41
SpikeBlank | SpikeBlank | 7010_2_03032035.D | MatrixBlank |  |  |  | 2020-03-03 22:01
UCGCAAAFS1A_GC 2/3 | UCGCAAAFS1A_GC 2/3 | 7010_2_03032036.D | Sample |  |  |  | 2020-03-03 22:22
UCGCAAAFS1B_GC 2/3 | UCGCAAAFS1B_GC 2/3 | 7010_2_03032037.D | Sample |  |  |  | 2020-03-03 22:42
UCGCAAAFS1B_GC 2/3_Dup | UCGCAAAFS1B_GC 2/3_Dup | 7010_2_03032038.D | Sample |  |  |  | 2020-03-03 23:03
UCGCAAAFS3A_GC 2/3 | UCGCAAAFS3A_GC 2/3 | 7010_2_03032039.D | Sample |  |  |  | 2020-03-03 23:24
SpikeBlank | SpikeBlank | 7010_2_03032040.D | MatrixBlank |  |  |  | 2020-03-03 23:44
UCGCAAAFS3B_GC 2/3 | UCGCAAAFS3B_GC 2/3 | 7010_2_03032041.D | Sample |  |  |  | 2020-03-04 00:05
UCGCAACC9_1/29/20_GC 2/3 | UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032042.D | QC | 9 | 350 | 87.5 | 2020-03-04 00:25
UCGCAAT5hS2C_GC 2/3 | UCGCAAT5hS2C_GC 2/3 | 7010_2_03032043.D | Sample |  |  |  | 2020-03-04 00:46
ACN | ACN | 7010_2_03032044.D | Blank |  |  |  | 2020-03-04 01:06
UCGCAAT5hS3A_GC 2/3 | UCGCAAT5hS3A_GC 2/3 | 7010_2_03032045.D | Sample |  |  |  | 2020-03-04 01:27
UCGCAAAFS1C_GC 2/3 | UCGCAAAFS1C_GC 2/3 | 7010_2_03032046.D | Sample |  |  |  | 2020-03-04 01:48
UCGCAAT1hS3C_GC 2/3 | UCGCAAT1hS3C_GC 2/3 | 7010_2_03032047.D | Sample |  |  |  | 2020-03-04 02:08
SpikeBlank | SpikeBlank | 7010_2_03032048.D | MatrixBlank |  |  |  | 2020-03-04 02:29
UCGCAAT5hS1A_GC 2/3 | UCGCAAT5hS1A_GC 2/3 | 7010_2_03032049.D | Sample |  |  |  | 2020-03-04 02:49
UCGCAAAFS3C_GC 2/3 | UCGCAAAFS3C_GC 2/3 | 7010_2_03032050.D | Sample |  |  |  | 2020-03-04 03:10
UCGCAAT1hS1C_GC 2/3 | UCGCAAT1hS1C_GC 2/3 | 7010_2_03032051.D | Sample |  |  |  | 2020-03-04 03:30
UCGCAAT1hS2A_GC 2/3 | UCGCAAT1hS2A_GC 2/3 | 7010_2_03032052.D | Sample |  |  |  | 2020-03-04 03:51
UCGCAAAFS2B_GC 2/3 | UCGCAAAFS2B_GC 2/3 | 7010_2_03032053.D | Sample |  |  |  | 2020-03-04 04:11
UCGCAACC12_1/29/20_GC 2/3 | UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032054.D | QC | 12 | 1500 | 375 | 2020-03-04 04:32
SpikeBlank | SpikeBlank | 7010_2_03032055.D | MatrixBlank |  |  |  | 2020-03-04 04:52
UCGCAAAFS2C_GC 2/3 | UCGCAAAFS2C_GC 2/3 | 7010_2_03032056.D | Sample |  |  |  | 2020-03-04 05:13
UCGCAAT5hS2B_GC 2/3 | UCGCAAT5hS2B_GC 2/3 | 7010_2_03032057.D | Sample |  |  |  | 2020-03-04 05:34
UCGCAAAFS2A_GC 2/3 | UCGCAAAFS2A_GC 2/3 | 7010_2_03032058.D | Sample |  |  |  | 2020-03-04 05:54
SpikeBlank | SpikeBlank | 7010_2_03032059.D | MatrixBlank |  |  |  | 2020-03-04 06:15
UCGCAACC1_1/29/20_GC 2/3 | UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032060.D | QC | 1 | 7 | 1.75 | 2020-03-04 06:35
UCGCAACC2_1/29/20_GC 2/3 | UCGCAACC2_1/29/20_GC 2/3 | 7010_2_03032061.D | QC | 2 | 12 | 3 | 2020-03-04 06:56
UCGCAACC3_1/29/20_GC 2/3 | UCGCAACC3_1/29/20_GC 2/3 | 7010_2_03032062.D | QC | 3 | 20 | 5 | 2020-03-04 07:16
UCGCAACC4_1/29/20_GC 2/3 | UCGCAACC4_1/29/20_GC 2/3 | 7010_2_03032063.D | QC | 4 | 30 | 7.5 | 2020-03-04 07:37
UCGCAACC5_1/29/20_GC 2/3 | UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032064.D | QC | 5 | 50 | 12.5 | 2020-03-04 07:57
UCGCAACC6_1/29/20_GC 2/3 | UCGCAACC6_1/29/20_GC 2/3 | 7010_2_03032065.D | QC | 6 | 80 | 20 | 2020-03-04 08:18
UCGCAACC7_1/29/20_GC 2/3 | UCGCAACC7_1/29/20_GC 2/3 | 7010_2_03032066.D | QC | 7 | 125 | 31.25 | 2020-03-04 08:39
UCGCAACC8_1/29/20_GC 2/3 | UCGCAACC8_1/29/20_GC 2/3 | 7010_2_03032067.D | QC | 8 | 200 | 50 | 2020-03-04 08:59
UCGCAACC9_1/29/20_GC 2/3 | UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032068.D | QC | 9 | 350 | 87.5 | 2020-03-04 09:20
UCGCAACC10_1/29/20_GC 2/3 | UCGCAACC10_1/29/20_GC 2/3 | 7010_2_03032069.D | QC | 10 | 500 | 125 | 2020-03-04 09:40
UCGCAACC11_1/29/20_GC 2/3 | UCGCAACC11_1/29/20_GC 2/3 | 7010_2_03032070.D | QC | 11 | 800 | 200 | 2020-03-04 10:01
UCGCAACC12_1/29/20_GC 2/3 | UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032071.D | QC | 12 | 1500 | 375 | 2020-03-04 10:21
UCGCAACC13_1/29/20_GC 2/3 | UCGCAACC13_1/29/20_GC 2/3 | 7010_2_03032072.D | QC | 13 | 2500 | 625 | 2020-03-04 10:42
UCGCAACC14_1/29/20_GC 2/3 | UCGCAACC14_1/29/20_GC 2/3 | 7010_2_03032073.D | QC | 14 | 3500 | 875 | 2020-03-04 11:03
UCGCAACC15_1/29/20_GC 2/3 | UCGCAACC15_1/29/20_GC 2/3 | 7010_2_03032074.D | QC | 15 | 5000 | 1250 | 2020-03-04 11:23
SpikeBlank | SpikeBlank | 7010_2_03032075.D | MatrixBlank |  |  |  | 2020-03-04 11:44
ACN |  | 7010_2_020420000.D | Blank |  |  |  | 2020-02-04 13:56
SpikeBlank |  | 7010_2_020420001.D | MatrixBlank |  |  |  | 2020-02-04 14:17
SpikeBlank |  | 7010_2_020420002.D | MatrixBlank |  |  |  | 2020-02-04 14:37
UCGC3CC1_1/17/20_GC 2/3 |  | 7010_2_020420003.D | Cal | 1 | 7 | 1.75 | 2020-02-04 14:58
UCGC3CC2_1/17/20_GC 2/3 |  | 7010_2_020420004.D | Cal | 2 | 12 | 3 | 2020-02-04 15:19
UCGC3CC3_1/17/20_GC 2/3 |  | 7010_2_020420005.D | Cal | 3 | 20 | 5 | 2020-02-04 15:39
UCGC3CC4_1/17/20_GC 2/3 |  | 7010_2_020420006.D | Cal | 4 | 30 | 7.5 | 2020-02-04 16:00
UCGC3CC5_1/17/20_GC 2/3 |  | 7010_2_020420007.D | Cal | 5 | 50 | 12.5 | 2020-02-04 16:20
UCGC3CC6_1/17/20_GC 2/3 |  | 7010_2_020420008.D | Cal | 6 | 80 | 20 | 2020-02-04 16:41
UCGC3CC7_1/17/20_GC 2/3 |  | 7010_2_020420009.D | Cal | 7 | 125 | 31.25 | 2020-02-04 17:01
UCGC3CC8_1/17/20_GC 2/3 |  | 7010_2_020420010.D | Cal | 8 | 200 | 50 | 2020-02-04 17:22
UCGC3CC9_1/17/20_GC 2/3 |  | 7010_2_020420011.D | Cal | 9 | 350 | 87.5 | 2020-02-04 17:42
UCGC3CC10_1/17/20_GC 2/3 |  | 7010_2_020420012.D | Cal | 10 | 500 | 125 | 2020-02-04 18:03
UCGC3CC11_1/17/20_GC 2/3 |  | 7010_2_020420013.D | Cal | 11 | 800 | 200 | 2020-02-04 18:23
UCGC3CC12_1/17/20_GC 2/3 |  | 7010_2_020420014.D | Cal | 12 | 1500 | 375 | 2020-02-04 18:44
UCGC3CC13_1/17/20_GC 2/3 |  | 7010_2_020420015.D | Cal | 13 | 2500 | 625 | 2020-02-04 19:05
UCGC3CC14_1/17/20_GC 2/3 |  | 7010_2_020420016.D | Cal | 14 | 3500 | 875 | 2020-02-04 19:25
UCGC3CC15_1/17/20_GC 2/3 |  | 7010_2_020420017.D | Cal | 15 | 5000 | 1250 | 2020-02-04 19:46
SpikeBlank |  | 7010_2_020420018.D | MatrixBlank |  |  |  | 2020-02-04 20:06
QCUCGC3CC4_1/17/20_GC 2/3 |  | 7010_2_020420019.D | QC | 4 | 30 | 7.5 | 2020-02-04 20:27
QCUCGC3CC8_1/17/20_GC 2/3 |  | 7010_2_020420020.D | QC | 8 | 200 | 50 | 2020-02-04 20:47
QCUCGC3CC11_1/17/20_GC 2/3 |  | 7010_2_020420021.D | QC | 11 | 800 | 200 | 2020-02-04 21:08
SpikeBlank |  | 7010_2_020420022.D | MatrixBlank |  |  |  | 2020-02-04 21:28
ACN |  | 7010_2_020420027.D | Blank |  |  |  | 2020-02-04 23:11
ACN |  | 7010_2_020420054.D | Blank |  |  |  | 2020-02-05 08:26
SpikeBlank |  | 7010_2_020420032.D | MatrixBlank |  |  |  | 2020-02-05 00:54
SpikeBlank |  | 7010_2_020420037.D | MatrixBlank |  |  |  | 2020-02-05 02:37
SpikeBlank |  | 7010_2_020420053.D | MatrixBlank |  |  |  | 2020-02-05 08:06
UCGC3AFS3A_GC 2/3 |  | 7010_2_020420030.D | Sample |  |  |  | 2020-02-05 00:13
UCGC3AFS3B_GC 2/3 |  | 7010_2_020420031.D | Sample |  |  |  | 2020-02-05 00:34
UCGC3AFS3C_GC 2/3 |  | 7010_2_020420036.D | Sample |  |  |  | 2020-02-05 02:16
UCGC3CC1_1/17/20_GC 2/3 |  | 7010_2_020420038.D | QC | 1 | 7 | 1.75 | 2020-02-05 02:57
UCGC3CC10_1/17/20_GC 2/3 |  | 7010_2_020420047.D | QC | 10 | 500 | 125 | 2020-02-05 06:02
UCGC3CC11_1/17/20_GC 2/3 |  | 7010_2_020420048.D | QC | 11 | 800 | 200 | 2020-02-05 06:23
UCGC3CC12_1/17/20_GC 2/3 |  | 7010_2_020420049.D | QC | 12 | 1500 | 375 | 2020-02-05 06:44
UCGC3CC13_1/17/20_GC 2/3 |  | 7010_2_020420033.D | QC | 13 | 2500 | 625 | 2020-02-05 01:15
UCGC3CC13_1/17/20_GC 2/3 |  | 7010_2_020420050.D | QC | 13 | 2500 | 625 | 2020-02-05 07:04
UCGC3CC14_1/17/20_GC 2/3 |  | 7010_2_020420051.D | QC | 14 | 3500 | 875 | 2020-02-05 07:25
UCGC3CC2_1/17/20_GC 2/3 |  | 7010_2_020420039.D | QC | 2 | 12 | 3 | 2020-02-05 03:18
UCGC3CC3_1/17/20_GC 2/3 |  | 7010_2_020420040.D | QC | 3 | 20 | 5 | 2020-02-05 03:38
UCGC3CC4_1/17/20_GC 2/3 |  | 7010_2_020420041.D | QC | 4 | 30 | 7.5 | 2020-02-05 03:59
UCGC3CC5_1/17/20_GC 2/3 |  | 7010_2_020420042.D | QC | 5 | 50 | 12.5 | 2020-02-05 04:20
UCGC3CC6_1/17/20_GC 2/3 |  | 7010_2_020420043.D | QC | 6 | 80 | 20 | 2020-02-05 04:40
UCGC3CC7_1/17/20_GC 2/3 |  | 7010_2_020420044.D | QC | 7 | 125 | 31.25 | 2020-02-05 05:01
UCGC3CC8_1/17/20_GC 2/3 |  | 7010_2_020420045.D | QC | 8 | 200 | 50 | 2020-02-05 05:21
UCGC3CC9_1/17/20_GC 2/3 |  | 7010_2_020420028.D | QC | 9 | 350 | 87.5 | 2020-02-04 23:32
UCGC3CC9_1/17/20_GC 2/3 |  | 7010_2_020420046.D | QC | 9 | 350 | 87.5 | 2020-02-05 05:42
UCGC3T1hS3A_GC 2/3 |  | 7010_2_020420055.D | Sample |  |  |  | 2020-02-05 08:47
UCGC3T1hS3B_GC 2/3 |  | 7010_2_020420024.D | Sample |  |  |  | 2020-02-04 22:10
UCGC3T1hS3B_GC 2/3_dup |  | 7010_2_020420025.D | Sample |  |  |  | 2020-02-04 22:30
UCGC3T1hS3C_GC 2/3 |  | 7010_2_020420035.D | Sample |  |  |  | 2020-02-05 01:56
UCGC3T5hS3A_GC 2/3 |  | 7010_2_020420034.D | Sample |  |  |  | 2020-02-05 01:35
UCGC3T5hS3B_GC 2/3 |  | 7010_2_020420026.D | Sample |  |  |  | 2020-02-04 22:51
ACN | PFAS Group 2 UC | 7010_1_020242000.D | Blank |  |  |  | 2020-02-24 15:25
SpikeBlank | PFAS Group 2 UC | 7010_1_020242001.D | MatrixBlank |  |  |  | 2020-02-24 15:46
SpikeBlank | PFAS Group 2 UC | 7010_1_020242002.D | MatrixBlank |  |  |  | 2020-02-24 16:07
UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242003.D | Cal | 1 | 7 | 1.75 | 2020-02-24 16:29
UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242004.D | Cal | 2 | 12 | 3 | 2020-02-24 16:50
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242005.D | Cal | 3 | 20 | 5 | 2020-02-24 17:11
UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242006.D | Cal | 4 | 30 | 7.5 | 2020-02-24 17:32
UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242007.D | Cal | 5 | 50 | 12.5 | 2020-02-24 17:54
UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242008.D | Cal | 6 | 80 | 20 | 2020-02-24 18:15
UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242009.D | Cal | 7 | 125 | 31.25 | 2020-02-24 18:36
UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242010.D | Cal | 8 | 200 | 50 | 2020-02-24 18:57
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242011.D | Cal | 9 | 350 | 87.5 | 2020-02-24 19:19
UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242012.D | Cal | 10 | 500 | 125 | 2020-02-24 19:40
UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242013.D | Cal | 11 | 800 | 200 | 2020-02-24 20:01
UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242014.D | Cal | 12 | 1500 | 375 | 2020-02-24 20:23
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242015.D | Cal | 13 | 2500 | 625 | 2020-02-24 20:44
UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242016.D | Cal | 14 | 3500 | 875 | 2020-02-24 21:05
UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242017.D | Cal | 15 | 5000 | 1250 | 2020-02-24 21:26
SpikeBlank | PFAS Group 2 UC | 7010_1_020242018.D | MatrixBlank |  |  |  | 2020-02-24 21:48
QCUCGCAACC4_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242019.D | QC | 4 | 30 | 7.5 | 2020-02-24 22:09
QCUCGCAACC8_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242020.D | QC | 8 | 200 | 50 | 2020-02-24 22:30
QCUCGCAACC11_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242021.D | QC | 11 | 800 | 200 | 2020-02-24 22:51
SpikeBlank | PFAS Group 2 UC | 7010_1_020242022.D | MatrixBlank |  |  |  | 2020-02-24 23:12
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242023.D | QC | 3 | 20 | 5 | 2020-02-24 23:34
UCGC0213T1hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242024.D | Sample |  |  |  | 2020-02-24 23:55
UCGC0213T1hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242025.D | Sample |  |  |  | 2020-02-25 00:16
UCGC0213T5hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242026.D | Sample |  |  |  | 2020-02-25 00:37
ACN | PFAS Group 2 UC | 7010_1_020242027.D | Blank |  |  |  | 2020-02-25 00:59
UCGC0213T5hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242028.D | Sample |  |  |  | 2020-02-25 01:20
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242029.D | QC | 9 | 350 | 87.5 | 2020-02-25 01:41
UCGC0213AFS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242030.D | Sample |  |  |  | 2020-02-25 02:03
UCGC0213AFS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242031.D | Sample |  |  |  | 2020-02-25 02:24
SpikeBlank | PFAS Group 2 UC | 7010_1_020242032.D | MatrixBlank |  |  |  | 2020-02-25 02:45
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242033.D | QC | 13 | 2500 | 625 | 2020-02-25 03:07
UCGC0213T5hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242034.D | Sample |  |  |  | 2020-02-25 03:28
UCGC0213T1hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242035.D | Sample |  |  |  | 2020-02-25 03:49
SpikeBlank | PFAS Group 2 UC | 7010_1_020242036.D | MatrixBlank |  |  |  | 2020-02-25 04:11
UCGC0213AFS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242037.D | Sample |  |  |  | 2020-02-25 04:32
ACN | PFAS Group 2 UC | 7010_1_020242038.D | Blank |  |  |  | 2020-02-25 04:53
UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242039.D | QC | 1 | 7 | 1.75 | 2020-02-25 05:15
UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242040.D | QC | 2 | 12 | 3 | 2020-02-25 05:36
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242041.D | QC | 3 | 20 | 5 | 2020-02-25 05:57
UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242042.D | QC | 4 | 30 | 7.5 | 2020-02-25 06:19
UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242043.D | QC | 5 | 50 | 12.5 | 2020-02-25 06:40
UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242044.D | QC | 6 | 80 | 20 | 2020-02-25 07:01
UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242045.D | QC | 7 | 125 | 31.25 | 2020-02-25 07:23
UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242046.D | QC | 8 | 200 | 50 | 2020-02-25 07:44
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242047.D | QC | 9 | 350 | 87.5 | 2020-02-25 08:05
UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242048.D | QC | 10 | 500 | 125 | 2020-02-25 08:27
UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242049.D | QC | 11 | 800 | 200 | 2020-02-25 08:48
UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242050.D | QC | 12 | 1500 | 375 | 2020-02-25 09:09
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242051.D | QC | 13 | 2500 | 625 | 2020-02-25 09:31
UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242052.D | QC | 14 | 3500 | 875 | 2020-02-25 09:52
UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242053.D | QC | 15 | 5000 | 1250 | 2020-02-25 10:13
SpikeBlank | PFAS Group 2 UC | 7010_1_020242054.D | MatrixBlank |  |  |  | 2020-02-25 10:35

## FractionUnbound_Adjusted
Ultracentrifugation Plasma protein binding assay - Experimental data and Fu calculations
|  |  |  | Mean | SD | CV | Mean |  | Time 1 hr |  |  | Time 5 hr |  |  | Aqueous Fraction |  |  | Fraction Unbound (fu) |  |  | Stability (T5hr/ T1hr)
DTXSID | Chemical | AbbrevSampleID | CAS# | Fu | Fu | Fu | Stability | UC Assay Date | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | t test
DTXSID10379991 | 3-(Perfluorooctyl)propanol | 267 | 1651-41-8 | 0.035748 | 0.0085689 | 23.9701 | 0.89348 | Run-specific Avg | 11367.0403 |  |  | 10162.1309 |  |  | 357.857 |  |  | 0.035748 |  |  | 0.894 |  |  | 0.10066
[8 empty] | Run-specific CV | [9 empty] | 23.9701
| 3-(Perfluorooctyl)propanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
| 3-(Perfluorooctyl)propanol | [6 empty] | 1/29 UC, 2/3 GC | 11954.2931 | 11059.4748 | 11087.353 | 10944.9958 | 10453.9462 | 9087.4506 | 303.3919 | 363.3254 | 406.8537 | 0.02772 | 0.034755 | 0.044771 | 0.91557 | 0.94525 | 0.81962
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID0059871 | Pentafluoropropionamide | 273 | 354-76-7 | 0.80894 | 0.037181 | 4.5962 | 0.71052 | Run-specific Avg | 5577.2522 |  |  | 9714.7036 |  |  | 5577.2522 |  |  | 0.80894 |  |  | 1.7418 |  |  | 0.0033341
[8 empty] | Run-specific CV | [9 empty] | 4.5962
DTXSID0059871 | Pentafluoropropionamide | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID0059871 | Pentafluoropropionamide | [6 empty] | 1/29 UC, 2/3 GC | 10054.2794 | 9452.2742 | 9637.5571 | 7021.7502 | 6955.9295 | 6720.0306 | 5532.4948 | 5474.6522 | 5724.6096 | 0.78791 | 0.78705 | 0.85187 | 0.69838 | 0.7359 | 0.69728
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | 476 | 532-64-3 | 0.54507 | 0.050973 | 9.3515 | 0.93518 | Run-specific Avg | 10832.6542 |  |  | 10129.3608 |  |  | 5511.0521 |  |  | 0.54507 |  |  | 0.93508 |  |  | 0.099207
[8 empty] | Run-specific CV | [9 empty] | 9.3515
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | [6 empty] | 1/29 UC, 2/3 GC | 10818.4284 | 11001.1589 | 10678.3754 | 10581.3851 | 9967.8963 | 9838.8011 | 5369.1706 | 5230.4309 | 5933.5547 | 0.50742 | 0.52473 | 0.60308 | 0.97809 | 0.90608 | 0.92138
[9 empty] | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13
DTXSID80382093 | 2-Aminohexafluoropropan-2-ol | 899 | 31253-34-6 | 1.0718 | 0.31304 | 29.206 | 0.76008 | Run-specific Avg | 9071.4846 | 10568.9788 |  | 6928.9009 | 7948.3632 |  | 7717.5797 | 7599.9235 |  | 1.1507 | 0.99293 |  | 0.76891 | 0.75125 |  | 0.0078233
[8 empty] | Run-specific CV | [9 empty] | 30.5718 | 32.2514
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 1/17 UC, 2/3 GC | 9554.7072 | 8950.4847 | 8709.2618 | 5753.511 | 6963.2465 | 8069.9451 | 8931.0427 | 6984.0611 | 7237.6354 | 1.5523 | 1.003 | 0.89686 | 0.60217 | 0.77797 | 0.92659
[8 empty] | 2/13 UC, 2/19 GC | 10360.3764 | 10861.5489 | 10485.011 | 6287.125 | 8366.095 | 9191.8697 | 8545.7525 | 6409.1261 | 7844.892 | 1.3592 | 0.76608 | 0.85346 | 0.60684 | 0.77025 | 0.87667
[9 empty] | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 | 2020-10-09 | 2020-12-06
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | 900 | 355-80-6 | 0.3616 | 0.072492 | 20.0475 | 0.85466 | Run-specific Avg | 9071.4846 |  |  | 9038.3121 |  |  | 3202.6313 |  |  | 0.3616 |  |  | 0.99634 |  |  | 0.048761
[8 empty] | Run-specific CV | [9 empty] | 20.0475
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | [6 empty] | 1/17 UC, 2/3 GC | 9694.4306 | 11637.5142 | 10264.6613 | 7515.6157 | 10498.393 | 9100.9274 | 3345.6335 | 3311.6624 | 2950.5981 | 0.44516 | 0.31544 | 0.32421 | 0.77525 | 0.90212 | 0.88663
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID9059832 | Dodecafluoroheptanol | 906 | 335-99-9 | 0.053683 | 0.0092232 | 17.1809 | 0.87489 | Run-specific Avg | 9572.8323 |  |  | 8332.0906 |  |  | 442.2818 |  |  | 0.053683 |  |  | 0.87039 |  |  | 0.18816
[8 empty] | Run-specific CV | [9 empty] | 17.1809
DTXSID9059832 | Dodecafluoroheptanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID9059832 | Dodecafluoroheptanol | [6 empty] | 1/29 UC, 2/3 GC | 10740.2258 | 8829.4108 | 9148.8602 | 8797.8725 | 8847.8794 | 7350.5199 | 453.3503 | 404.5873 | 468.9078 | 0.05153 | 0.045727 | 0.063792 | 0.81915 | 1.0021 | 0.80344
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | 913 | 125070-38-4 | 0.36959 | 0.044797 | 12.1208 | 0.96498 | Run-specific Avg | 9031.1979 |  |  | 8687.647 |  |  | 3191.8738 |  |  | 0.36959 |  |  | 0.96196 |  |  | 0.47061
[8 empty] | Run-specific CV | [9 empty] | 12.1208
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | [6 empty] | 1/29 UC, 2/3 GC | 9522.9171 | 9681.0035 | 7889.6732 | 8409.304 | 9622.7874 | 8030.8496 | 2936.0246 | 3258.9126 | 3380.6841 | 0.34914 | 0.33867 | 0.42096 | 0.88306 | 0.99399 | 1.0179
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID4059914 | Heptafluorobutanol | 915 | 375-01-9 | 0.51051 | 0.031035 | 6.0792 | 0.81457 | Run-specific Avg | 7721.1285 |  |  | 6284.5375 |  |  | 3202.9826 |  |  | 0.51051 |  |  | 0.81394 |  |  | 0.009632
[8 empty] | Run-specific CV | [9 empty] | 6.0792
DTXSID4059914 | Heptafluorobutanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID4059914 | Heptafluorobutanol | [6 empty] | 1/29 UC, 2/3 GC | 8203.6713 | 7703.2854 | 7256.4288 | 6637.9647 | 6112.1517 | 6103.4961 | 3187.135 | 3313.8036 | 3108.0091 | 0.48014 | 0.54217 | 0.50922 | 0.80915 | 0.79345 | 0.84112
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID60379269 | 3-(Perfluoropropyl)propanol | 965 | 679-02-7 | 0.2033 | 0.022537 | 11.086 | 0.89158 | Run-specific Avg | 8380.7575 |  |  | 7456.4672 |  |  | 1511.5601 |  |  | 0.2033 |  |  | 0.88971 |  |  | 0.049685
[8 empty] | Run-specific CV | [9 empty] | 11.086
DTXSID60379269 | 3-(Perfluoropropyl)propanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID60379269 | 3-(Perfluoropropyl)propanol | [6 empty] | 1/29 UC, 2/3 GC | 9125.9621 | 8240.63 | 7775.6803 | 7832.9313 | 7312.2343 | 7224.236 | 1402.8217 | 1635.5949 | 1496.2636 | 0.17909 | 0.22368 | 0.20712 | 0.85831 | 0.88734 | 0.92908
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID70381151 | Perfluorooctanamidine | 3125 | 307-31-3 | 0.025258 | 0.00066538 | 2.6344 | 0.67416 | Run-specific Avg | 7765.4095 |  |  | 5232.2167 |  |  | 132.0575 |  |  | 0.025258 |  |  | 0.67379 |  |  | 0.0050268
[8 empty] | Run-specific CV | [9 empty] | 2.6344
DTXSID70381151 | Perfluorooctanamidine | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID70381151 | Perfluorooctanamidine | [6 empty] | 1/29 UC, 3/4 GC | 8179.2804 | 7783.803 | 7333.145 | 5292.903 | 5490.0647 | 4913.6823 | 135.84 | 134.4506 | 125.8818 | 0.025665 | 0.02449 | 0.025619 | 0.64711 | 0.70532 | 0.67006
[9 empty] | 1/29, 2/3 | 1/29, 3/4 | 1/17, 2/3 | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-29 | 2020-01-29 | 2020-01-17 | 2020-01-29 | 2020-01-29 | 2020-01-17
DTXSID5023792 | 4-Nitrotoluene | 4NT | 99-99-0 | 0.058092 | 0.0040785 | 7.0207 | 0.81599 | Run-specific Avg | 9078.7805 | 9327.0662 | 9088.8629 | 7477.2508 | 7808.8846 | 7376.5868 | 448.1106 | 416.7227 | 447.7998 | 0.060189 | 0.053392 | 0.060696 | 0.82353 | 0.83713 | 0.8121 | 1.9639e-13
[8 empty] | Run-specific CV | [9 empty] | 11.9233 | 3.7354 | 3.6294
[9 empty] | 2020-02-13 |  |  | 2020-02-13 |  |  | 2020-02-13 |  |  | 2020-02-13 |  |  | 2020-02-13
[8 empty] | Run-specific Avg | 9049.4441 |  |  | 7013.9309 |  |  | 542.2686 |  |  | 0.078211 |  |  | 0.77614
[8 empty] | Run-specific CV | [9 empty] | 38.2582
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 1/29 UC, 2/3 GC | 9444.9707 | 9574.5942 | 9167.7485 | 8010.7889 | 7913.7988 | 7693.4274 | 445.027 | 424.9107 | 434.1477 | 0.055553 | 0.053692 | 0.056431 | 0.84815 | 0.82654 | 0.83918
[9 empty] | 9144.311 | 9134.2806 | 8645.8651 | 7315.3486 | 7530.7245 | 6675.7867 | 503.825 | 427.6334 | 469.962 | 0.068872 | 0.056785 | 0.070398 | 0.79999 | 0.82445 | 0.77214
[9 empty] | 9102.6985 | 9030.2931 | 8464.2628 | 7507.9112 | 7288.6739 | 7358.7974 | 403.0495 | 507.2471 | 417.1933 | 0.053683 | 0.069594 | 0.056693 | 0.8248 | 0.80714 | 0.8694
[8 empty] | 1/29 UC, 3/4 GC | 9433.7903 | 9393.7768 | 9153.6314 | 7998.0177 | 7851.7076 | 7576.9284 | 424.6067 | 404.8851 | 420.6762 | 0.053089 | 0.051567 | 0.055521 | 0.84781 | 0.83584 | 0.82775
[8 empty] | 1/17 UC, 2/3 GC | 8671.4536 | 9222.0326 | 9373.1027 | 7181.9139 | 7614.5132 | 7333.3332 | 421.4128 | 459.6536 | 462.3331 | 0.058677 | 0.060365 | 0.063045 | 0.82822 | 0.82569 | 0.78238
[8 empty] | 2/13 UC, 2/19 GC | 9120.2086 | 9242.5968 | 8785.5268 | 6619.286 | 6926.8796 | 7495.6271 | 745.541 | 404.5371 | 476.7277 | 0.11263 | 0.058401 | 0.063601 | 0.72578 | 0.74945 | 0.85318

## FractionUnbound_Old
Ultracentrifugation Plasma protein binding assay - Experimental data and Fu calculations
|  |  |  | Mean | SD | CV | Mean |  | Time 1 hr |  |  | Time 5 hr |  |  | Aqueous Fraction |  |  | Fraction Unbound (fu) |  |  | Stability (T5hr/ T1hr)
DTXSID | Chemical | AbbrevSampleID | CAS# | Fu | Fu | Fu | Stability | UC Assay Date | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | t test
DTXSID10379991 | 3-(Perfluorooctyl)propanol | 267 | 1651-41-8 | 0.17874 | 0.042845 | 23.9701 | 0.894 | Overall | 284.176 |  |  | 254.0533 |  |  | 44.7321 |  |  | 0.17874 |  |  | 0.894 |  |  | #DIV/0!
| 3-(Perfluorooctyl)propanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
| 3-(Perfluorooctyl)propanol | [6 empty] | 1/29 UC, 2/3 GC | 298.8573 | 276.4869 | 277.1838 | 273.6249 | 261.3487 | 227.1863 | 37.924 | 45.4157 | 50.8567 | 0.1386 | 0.17377 | 0.22385 | 0.91557 | 0.94525 | 0.81962
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID0059871 | Pentafluoropropionamide | 273 | 354-76-7 | 4.0447 | 0.1859 | 4.5962 | 0.34837 | Overall | 697.1565 |  |  | 242.8676 |  |  | #REF! |  |  | 4.0447 |  |  | 0.34837 |  |  | #DIV/0!
DTXSID0059871 | Pentafluoropropionamide | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID0059871 | Pentafluoropropionamide | [6 empty] | 1/29 UC, 2/3 GC | 251.357 | 236.3069 | 240.9389 | 175.5438 | 173.8982 | 168.0008 | 691.5618 | 684.3315 | 715.5762 | 3.9395 | 3.9352 | 4.2594 | 0.69838 | 0.7359 | 0.69728
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | 476 | 532-64-3 | 2.7254 | #DIV/0! | #DIV/0! | 0.93508 | Overall | 270.8164 |  |  | 253.234 |  |  | 688.8815 |  |  | 2.7254 |  |  | 0.93508 |  |  | #DIV/0!
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID40379666 | N-Methyl-N-trimethylsilylheptafluorobutyramide | [6 empty] | 1/29 UC, 2/3 GC | 270.4607 | 275.029 | 266.9594 | 264.5346 | 249.1974 | 245.97 | 671.1463 | 653.8039 | 741.6943 | 2.5371 | 2.6236 | 3.0154 | 0.97809 | 0.90608 | 0.92138
[9 empty] | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13 |  | 2020-01-17 | 2020-02-13
DTXSID80382093 | 2-Aminohexafluoropropan-2-ol | 899 | 31253-34-6 | 5.7535 | 0.55784 | 9.6955 | #REF! | Overall | #REF! | NC | 264.2245 | 173.2225 | NC | 198.7091 | 964.6975 | NA | 949.9904 | 5.7535 | 4.9646 |  | #REF! | NA | 0.75205 | #REF!
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 1/17 UC, 2/3 GC | 238.8677 | 223.7621 | 217.7315 | 143.8378 | 174.0812 | 201.7486 | 1116.3803 | 873.0076 | 904.7044 | 7.7614 | 5.0149 | 4.4843 | 0.60217 | 0.77797 | 0.92659
[8 empty] | 2/13 UC, 2/19 GC | 259.0094 | 271.5387 | 262.1253 | 157.1781 | 209.1524 | 229.7967 | 1068.2191 | 801.1408 | 980.6115 | 6.7962 | 3.8304 | 4.2673 | 0.60684 | 0.77025 | 0.87667
[9 empty] | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 |  |  | 2020-01-17 | 2020-10-09 | 2020-12-06
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | 900 | 355-80-6 | 1.808 | #DIV/0! | #DIV/0! | 0.99634 | Overall | 226.7871 |  |  | 225.9578 |  |  | 400.3289 |  |  | 1.808 |  |  | 0.99634 |  |  | #DIV/0!
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID0059879 | 1H,1H,5H-Perfluoropentanol | [6 empty] | 1/17 UC, 2/3 GC | 242.3608 | 290.9379 | 256.6165 | 187.8904 | 262.4598 | 227.5232 | 418.2042 | 413.9578 | 368.8248 | 2.2258 | 1.5772 | 1.621 | 0.77525 | 0.90212 | 0.88663
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID9059832 | Dodecafluoroheptanol | 906 | 335-99-9 | 0.26842 | #DIV/0! | #DIV/0! | 0.87039 | Overall | 239.3208 |  |  | 208.3023 |  |  | 55.2852 |  |  | 0.26842 |  |  | 0.87039 |  |  | #DIV/0!
DTXSID9059832 | Dodecafluoroheptanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID9059832 | Dodecafluoroheptanol | [6 empty] | 1/29 UC, 2/3 GC | 268.5056 | 220.7353 | 228.7215 | 219.9468 | 221.197 | 183.763 | 56.6688 | 50.5734 | 58.6135 | 0.25765 | 0.22864 | 0.31896 | 0.81915 | 1.0021 | 0.80344
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | 913 | 125070-38-4 | 1.8479 | #DIV/0! | #DIV/0! | 0.96196 | Overall | 225.7799 |  |  | 217.1912 |  |  | 398.9842 |  |  | 1.8479 |  |  | 0.96196 |  |  | #DIV/0!
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID10382147 | 3-(Perfluoro-2-butyl)propane-1,2-diol | [6 empty] | 1/29 UC, 2/3 GC | 238.0729 | 242.0251 | 197.2418 | 210.2326 | 240.5697 | 200.7712 | 367.0031 | 407.3641 | 422.5855 | 1.7457 | 1.6933 | 2.1048 | 0.88306 | 0.99399 | 1.0179
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID4059914 | Heptafluorobutanol | 915 | 375-01-9 | 2.5525 | #DIV/0! | #DIV/0! | 0.81394 | Overall | 193.0282 |  |  | 157.1134 |  |  | 400.3728 |  |  | 2.5525 |  |  | 0.81394 |  |  | #DIV/0!
DTXSID4059914 | Heptafluorobutanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID4059914 | Heptafluorobutanol | [6 empty] | 1/29 UC, 2/3 GC | 205.0918 | 192.5821 | 181.4107 | 165.9491 | 152.8038 | 152.5874 | 398.3919 | 414.2254 | 388.5011 | 2.4007 | 2.7108 | 2.5461 | 0.80915 | 0.79345 | 0.84112
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID60379269 | 3-(Perfluoropropyl)propanol | 965 | 679-02-7 | 1.0165 | #DIV/0! | #DIV/0! | 0.88971 | Overall | 209.5189 |  |  | 186.4117 |  |  | 188.945 |  |  | 1.0165 |  |  | 0.88971 |  |  | #DIV/0!
DTXSID60379269 | 3-(Perfluoropropyl)propanol | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID60379269 | 3-(Perfluoropropyl)propanol | [6 empty] | 1/29 UC, 2/3 GC | 228.1491 | 206.0158 | 194.392 | 195.8233 | 182.8059 | 180.6059 | 175.3527 | 204.4494 | 187.033 | 0.89546 | 1.1184 | 1.0356 | 0.85831 | 0.88734 | 0.92908
[9 empty] | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29 |  |  | 2020-01-29
DTXSID70381151 | Perfluorooctanamidine | 3125 | 307-31-3 | 0.12629 | #DIV/0! | #DIV/0! | 0.67379 | Overall | 194.1352 |  |  | 130.8054 |  |  | 16.5072 |  |  | 0.12629 |  |  | 0.67379 |  |  | #DIV/0!
DTXSID70381151 | Perfluorooctanamidine | [7 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
DTXSID70381151 | Perfluorooctanamidine | [6 empty] | 1/29 UC, 3/4 GC | 204.482 | 194.5951 | 183.3286 | 132.3226 | 137.2516 | 122.8421 | 16.98 | 16.8063 | 15.7352 | 0.12832 | 0.12245 | 0.12809 | 0.64711 | 0.70532 | 0.67006
[9 empty] | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-17 | 2020-01-29 | 2020-02-13 | 2020-01-17 | 2020-01-29 | 2020-02-13
DTXSID5023792 | 4-Nitrotoluene | 4NT |  | 0.32288 | 0.060387 | 18.7029 | #REF! | Overall | #REF! | 229.6324 | 226.2361 | #REF! | 188.0828 | 175.3483 | #REF! | 56.3647 | 67.7836 | 0.27613 | 0.30144 | 0.39106 | #REF! | 0.81841 | 0.77614 | #REF!
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 1/29 UC, 2/3 GC | 236.1243 | 239.3649 | 229.1937 | 200.2697 | 197.845 | 192.3357 | 55.6284 | 53.1138 | 54.2685 | 0.27777 | 0.26846 | 0.28215 | 0.84815 | 0.82654 | 0.83918
[9 empty] | 228.6078 | 228.357 | 216.1466 | 182.8837 | 188.2681 | 166.8947 | 62.9781 | 53.4542 | 58.7453 | 0.34436 | 0.28393 | 0.35199 | 0.79999 | 0.82445 | 0.77214
[9 empty] | 227.5675 | 225.7573 | 211.6066 | 187.6978 | 182.2168 | 183.9699 | 50.3812 | 63.4059 | 52.1492 | 0.26842 | 0.34797 | 0.28347 | 0.8248 | 0.80714 | 0.8694
[8 empty] | 1/29 UC, 3/4 GC | 235.8448 | 234.8444 | 228.8408 | 199.9504 | 196.2927 | 189.4232 | 53.0758 | 50.6106 | 52.5845 | 0.26544 | 0.25783 | 0.2776 | 0.84781 | 0.83584 | 0.82775
[8 empty] | 1/17 UC, 2/3 GC | 216.7863 | 230.5508 | 234.3276 | 179.5478 | 190.3628 | 183.3333 | 52.6766 | 57.4567 | 57.7916 | 0.29338 | 0.30183 | 0.31523 | 0.82822 | 0.82569 | 0.78238
[8 empty] | 2/13 UC, 2/19 GC | 228.0052 | 231.0649 | 219.6382 | 165.4822 | 173.172 | 187.3907 | 93.1926 | 50.5671 | 59.591 | 0.56316 | 0.29201 | 0.318 | 0.72578 | 0.74945 | 0.85318

## UC Data
Sample | [6 empty] | 915 Method |  | 915 Results |  |  |  | MFBET (ISTD) Results |  | 965 Method |  | 965 Results |  |  |  | MFHET (ISTD) Results |  | 476 Method |  | 476 Results |  |  |  | MFHET (ISTD) Results |  | 267 Method |  | 267 Results |  |  |  | MFOET (ISTD) Results |  | 3125 Method |  | 3125 Results |  |  |  | MFOET (ISTD) Results |  | 906 Method |  | 906 Results |  |  |  | MFOET (ISTD) Results |  | 273 Method |  | 273 Results |  |  |  | MFOET (ISTD) Results |  | 913 Method |  | 913 Results |  |  |  | MFOET (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| ACN | PFAS Group 2 UC | 7010_1_02042000.D | Blank |  | 2020-02-04 15:48 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 4.6821 |  |  | 40366.6486 | 5.6374 | 0 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1028 | 0 | 0 | 0 | 6.3404 | 0 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0942 | 0 | 0 | 0 | 6.3404 | 0 | y = 484.688584 * x | 0.99801 | 7.7173 | 0 | 0 | 0 | 7.1098 | 0 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2348 | 0 | 0 | 0 | 7.1098 | 0 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3665 | 0 | 0 | 0 | 7.1098 | 0 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6607 | 0 | 0 | 0 | 7.1098 | 0 | y = 4960.639652 * x | 0.99845 | 11.1927 | 0 | 0 | 0 | 7.1098 | 0 | y = 153.621749 * x | 0.99954 | 11.7094 | 0 | 0 | 0 | 11.475 | 0
| ACN | PFAS Group 2 UC | 7010_1_02042031.D | Blank |  | 2020-02-05 02:48 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 4.6897 | 0 | 0 | 0 | 5.9912 | 0 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1209 | 0 | 0 | 0 | 6.2614 | 0 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.89 | 0 | 0 | 0 | 6.2614 | 0 | y = 484.688584 * x | 0.99801 | 7.735 | 0 | 0 | 0 | 7.0893 | 0 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 0 | 7.0893 | 0 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.2746 | 0 | 0 | 0 | 7.0893 | 0 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6401 | 0 | 0 | 0 | 7.0893 | 0 | y = 4960.639652 * x | 0.99845 | 11.4261 | 0 | 0 | 0 | 7.0893 | 0 | y = 153.621749 * x | 0.99954 | 11.468 | 0 | 0 | 0 | 11.2259 | 0
| ACN | PFAS Group 2 UC | 7010_1_02042060.D | Blank |  | 2020-02-05 13:07 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5103 | 0 | 0 | 0 | 5.6334 | 0 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0815 | 0 | 0 | 0 | 6.2827 | 0 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 0 | 0 | 0 | 6.2827 | 0 | y = 484.688584 * x | 0.99801 | 7.787 | 0 | 0 | 0 | 7.1275 | 0 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.1781 | 0 | 0 | 0 | 7.1275 | 0 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3588 | 0 | 0 | 0 | 7.1275 | 0 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6886 | 0 | 0 | 0 | 7.1275 | 0 | y = 4960.639652 * x | 0.99845 | 11.4494 | 0 | 0 | 0 | 7.1275 | 0 | y = 153.621749 * x | 0.99954 | 11.2112 | 0 | 0 | 0 | 11.1636 | 0
| ACN | PFAS Group 2 UC | 7010_1_02042080.D | Blank |  | 2020-02-05 20:14 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 4.7053 | 0 | 0 | 0 | 5.6373 | 0 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0845 | 0 | 0 | 0 | 6.316 | 0 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0876 | 0 | 0 | 0 | 6.316 | 0 | y = 484.688584 * x | 0.99801 | 7.7003 | 0 | 0 | 0 | 7.0755 | 0 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2193 |  |  | 434.0766 | 7.0755 | 0 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3613 | 0 | 0 | 0 | 7.0755 | 0 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.7549 | 0 | 0 | 0 | 7.0755 | 0 | y = 4960.639652 * x | 0.99845 | 10.9202 | 0 | 0 | 0 | 7.0755 | 0 | y = 153.621749 * x | 0.99954 | 12.184 | 0 | 0 | 0 | 10.7823 | 0
| QCUCGCAACC11_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042035.D | QC | 11 | 2020-02-05 04:13 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5841 | 165.6581 | 82.829 | 205339.9684 | 5.6606 | 37215.3988 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1057 | 162.7312 | 81.3656 | 426276.8148 | 6.3008 | 28996.3692 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 186.1059 | 93.053 | 186909.6193 | 6.3008 | 28996.3692 | y = 484.688584 * x | 0.99801 | 7.7904 | 115.6368 | 57.8184 | 61410.7936 | 7.1136 | 13148.2235 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 160.0331 | 80.0166 | 41373.2785 | 7.1136 | 13148.2235 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 184.1054 | 92.0527 | 1531699.5618 | 7.1136 | 13148.2235 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 196.9925 | 98.4962 | 76328.9693 | 7.1136 | 13148.2235 | y = 4960.639652 * x | 0.99845 | 11.3249 | 180.8328 | 90.4164 | 982880.6616 | 7.1136 | 13148.2235 | y = 153.621749 * x | 0.99954 | 11.2189 | 190.9872 | 95.4936 | 954607.6467 | 11.2181 | 390435.4446
| QCUCGCAACC4_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042033.D | QC | 4 | 2020-02-05 03:30 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5958 | 1.7629 | 23.5048 | 2194.517 | 5.6684 | 37408.4757 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 1.5009 | 20.0122 | 4031.6268 | 6.3039 | 29211.8116 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 2.5932 | 34.5765 | 2076.7738 | 6.3039 | 29211.8116 | y = 484.688584 * x | 0.99801 | 7.7904 | 2.7519 | 36.6919 | 1492.8834 | 7.1136 | 13431.1384 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 | 0 | 583.6231 | 7.1136 | 13431.1384 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 | 0 | 18736.3005 | 7.1136 | 13431.1384 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 2.034 | 27.1195 | 812.1352 | 7.1136 | 13431.1384 | y = 4960.639652 * x | 0.99845 | 11.3482 | 2.2551 | 30.0675 | 12520.687 | 7.1136 | 13431.1384 | y = 153.621749 * x | 0.99954 | 11.2189 | 2.6486 | 35.3149 | 13353.4567 | 11.2181 | 393825.2296
| QCUCGCAACC8_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042034.D | QC | 8 | 2020-02-05 03:52 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 41.3405 | 82.6811 | 52422.276 | 5.6645 | 38097.3649 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 41.4005 | 82.801 | 106005.5361 | 6.3039 | 27967.0874 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 48.7293 | 97.4586 | 46810.4798 | 6.3039 | 27967.0874 | y = 484.688584 * x | 0.99801 | 7.7904 | 29.2179 | 58.4359 | 15662.8953 | 7.1171 | 13272.138 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 31.1812 | 62.3625 | 10925.6195 | 7.1171 | 13272.138 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 43.4496 | 86.8992 | 381397.6324 | 7.1171 | 13272.138 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 46.8744 | 93.7489 | 18457.757 | 7.1171 | 13272.138 | y = 4960.639652 * x | 0.99845 | 11.3249 | 39.3378 | 78.6756 | 215827.9211 | 7.1171 | 13272.138 | y = 153.621749 * x | 0.99954 | 11.219 | 48.7628 | 97.5257 | 237554.8759 | 11.2181 | 380542.8038
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042001.D | MatrixBlank |  | 2020-02-04 16:09 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1641 | 95.955 |  | 131536.0959 | 5.6528 | 41172.1507 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 0 | 0 | 0 | 6.2978 | 32044.9747 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0979 | 0 | 0 | 0 | 6.2978 | 32044.9747 | y = 484.688584 * x | 0.99801 | 7.7835 | 2.7259 |  | 1604.2079 | 7.1066 | 14570.3396 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 |  | 1629.3366 | 7.1066 | 14570.3396 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 6089.3436 | 7.1066 | 14570.3396 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6503 | 0 | 0 | 0 | 7.1066 | 14570.3396 | y = 4960.639652 * x | 0.99845 | 11.2159 | 3.5232 |  | 21220.6685 | 7.1066 | 14570.3396 | y = 153.621749 * x | 0.99954 | 11.2345 | 1.0375 |  | 5827.2768 | 11.2181 | 438721.4315
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042002.D | MatrixBlank |  | 2020-02-04 16:30 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1097 | 100.1017 |  | 136141.6344 | 5.6373 | 40847.531 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0875 | 0 | 0 | 0 | 6.2857 | 31968.7845 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 0 | 0 | 0 | 6.2857 | 31968.7845 | y = 484.688584 * x | 0.99801 | 7.7801 | 2.8637 |  | 1674.8842 | 7.1032 | 14480.1653 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 1201.8005 | 7.1032 | 14480.1653 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 4544.4565 | 7.1032 | 14480.1653 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6172 | 0 | 0 | 0 | 7.1032 | 14480.1653 | y = 4960.639652 * x | 0.99845 | 11.1692 | 0 | 0 | 0 | 7.1032 | 14480.1653 | y = 153.621749 * x | 0.99954 | 11.2345 | 1.2761 |  | 7577.747 | 11.2181 | 463849.0886
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042018.D | MatrixBlank |  | 2020-02-04 22:11 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1992 | 0 |  | 0 | 5.6645 | 39337.1317 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1452 | 0 | 0 | 0 | 6.2978 | 29058.2865 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 0 | 0 | 0 | 6.2978 | 29058.2865 | y = 484.688584 * x | 0.99801 | 7.7835 | 2.2658 |  | 1379.8781 | 7.1067 | 15077.7524 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 749.9133 | 7.1067 | 15077.7524 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 4244.9426 | 7.1067 | 15077.7524 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6657 | 0.72281 |  | 324.0092 | 7.1067 | 15077.7524 | y = 4960.639652 * x | 0.99845 | 11.1537 | 3.3648 |  | 20972.7817 | 7.1067 | 15077.7524 | y = 153.621749 * x | 0.99954 | 11.2268 | 2.0339 |  | 12127.2711 | 11.2181 | 465763.8198
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042022.D | MatrixBlank |  | 2020-02-04 23:36 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1525 | 0 |  | 0 | 5.6451 | 38859.6518 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0967 | 0 | 0 | 0 | 6.2978 | 30314.9518 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9316 | 0 | 0 | 0 | 6.2978 | 30314.9518 | y = 484.688584 * x | 0.99801 | 7.787 | 1.306 |  | 764.9152 | 7.1136 | 14500.284 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 595.053 | 7.1136 | 14500.284 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 2890.9212 | 7.1136 | 14500.284 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6707 | 0 | 0 | 0 | 7.1136 | 14500.284 | y = 4960.639652 * x | 0.99845 | 11.1926 | 0 | 0 | 0 | 7.1136 | 14500.284 | y = 153.621749 * x | 0.99954 | 11.2267 | 1.5462 |  | 8529.4496 | 11.2181 | 430906.5473
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042042.D | MatrixBlank |  | 2020-02-05 06:42 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1914 | 0 |  | 0 | 5.6568 | 43295.0213 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 0 | 0 | 0 | 6.3039 | 33984.1072 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0872 | 0 | 0 | 0 | 6.3039 | 33984.1072 | y = 484.688584 * x | 0.99801 | 7.7935 | 0.86289 |  | 535.622 | 7.1166 | 15368.0945 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2322 | 0 |  | 292.8998 | 7.1166 | 15368.0945 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1403.3844 | 7.1166 | 15368.0945 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.7957 | 0 | 0 | 0 | 7.1166 | 15368.0945 | y = 4960.639652 * x | 0.99845 | 11.395 | 2.3771 |  | 15101.6347 | 7.1166 | 15368.0945 | y = 153.621749 * x | 0.99954 | 11.2268 | 1.7166 |  | 10873.6586 | 11.2181 | 494808.5795
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042050.D | MatrixBlank |  | 2020-02-05 09:33 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.2147 | 0 |  | 0 | 5.6645 | 41200.1748 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0997 | 0 | 0 | 0 | 6.3039 | 35888.428 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9104 | 0 | 0 | 0 | 6.3039 | 35888.428 | y = 484.688584 * x | 0.99801 | 7.7969 | 0.64673 |  | 431.4648 | 7.1132 | 16517.2482 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 |  | 451.8663 | 7.1132 | 16517.2482 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 0 |  | 1439.0521 | 7.1132 | 16517.2482 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6606 | 0 | 0 | 0 | 7.1132 | 16517.2482 | y = 4960.639652 * x | 0.99845 | 11.2315 | 0 | 0 | 0 | 7.1132 | 16517.2482 | y = 153.621749 * x | 0.99954 | 11.219 | 1.6475 |  | 10450.2785 | 11.2181 | 495498.5055
| SpikeBlank | PFAS Group 2 UC | 7010_1_02042079.D | MatrixBlank |  | 2020-02-05 19:53 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1719 | 0 |  | 0 | 5.6451 | 41562.5619 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0996 | 0 | 0 | 0 | 6.3038 | 33089.5027 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9246 | 0 | 0 | 0 | 6.3038 | 33089.5027 | y = 484.688584 * x | 0.99801 | 7.8008 | 1.2958 |  | 752.9156 | 7.124 | 14386.0657 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 |  | 533.701 | 7.124 | 14386.0657 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 0 |  | 7614.0453 | 7.124 | 14386.0657 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 0 | 0 | 0 | 7.124 | 14386.0657 | y = 4960.639652 * x | 0.99845 | 11.7062 | 5.2749 |  | 31369.8159 | 7.124 | 14386.0657 | y = 153.621749 * x | 0.99954 | 11.2189 | 2.1616 |  | 13128.1978 | 11.2181 | 474411.87
| UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5608 | 58.5227 |  | 86254.2943 | 5.6451 | 44276.266 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0997 | 23.6592 |  | 75854.1305 | 6.3039 | 34951.2412 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 671.1463 | 97.4255 | 814200.3783 | 6.3039 | 34951.2412 | y = 484.688584 * x | 0.99801 | 7.7935 | 2.6544 |  | 1642.759 | 7.1201 | 15322.4289 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2322 | 0 |  | 1004.3118 | 7.1201 | 15322.4289 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.0945 |  | 44962.0124 | 7.1201 | 15322.4289 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 691.5618 | 99.1975 | 305308.0013 | 7.1201 | 15322.4289 | y = 4960.639652 * x | 0.99845 | 11.0992 | 0 |  | 0 | 7.1201 | 15322.4289 | y = 153.621749 * x | 0.99954 | 11.219 | 55.6284 | 102.3768 | 324929.2471 | 11.2181 | 456268.9968
| UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.4675 | 57.8433 |  | 53199.4156 | 5.5712 | 27629.305 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0663 | 20.8275 |  | 51626.4969 | 6.2735 | 27013.7501 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9108 | 653.8039 | 94.908 | 613019.7317 | 6.2735 | 27013.7501 | y = 484.688584 * x | 0.99801 | 7.7939 | 4.1633 |  | 2055.8706 | 7.117 | 12225.8416 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2347 | 0 |  | 1452.8442 | 7.117 | 12225.8416 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 3.1237 |  | 43725.6026 | 7.117 | 12225.8416 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 684.3315 | 98.1604 | 241140.2062 | 7.117 | 12225.8416 | y = 4960.639652 * x | 0.99845 | 11.465 | 0 |  | 0 | 7.117 | 12225.8416 | y = 153.621749 * x | 0.99954 | 11.2189 | 53.1138 | 97.7491 | 272083.1737 | 11.2181 | 400149.7862
| UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.4908 | 50.6497 |  | 57926.3776 | 5.5867 | 34358.3476 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0784 | 19.1734 |  | 47990.0789 | 6.2887 | 27272.3765 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 741.6943 | 107.6665 | 702155.1796 | 6.2887 | 27272.3765 | y = 484.688584 * x | 0.99801 | 7.7939 | 1.9084 |  | 1044.7429 | 7.124 | 13553.8369 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2373 | 0 | 0 | 0 | 7.124 | 13553.8369 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.2344 |  | 40955.056 | 7.124 | 13553.8369 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 715.5762 | 102.6421 | 279136.4205 | 7.124 | 13553.8369 | y = 4960.639652 * x | 0.99845 | 11.3171 | 1.2867 |  | 7209.4103 | 7.124 | 13553.8369 | y = 153.621749 * x | 0.99954 | 11.219 | 54.2685 | 99.8741 | 259199.7576 | 11.2181 | 373091.7078
| UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5919 | 66.0308 |  | 95664.8815 | 5.6645 | 43521.4244 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1087 | 22.547 |  | 74086.72 | 6.3069 | 35816.3161 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9142 | 40.4984 |  | 49707.4753 | 6.3069 | 35816.3161 | y = 484.688584 * x | 0.99801 | 7.7939 | 37.924 | 84.7802 | 26629.0199 | 7.117 | 17384.4001 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2372 | 0 | 0 | 0 | 7.117 | 17384.4001 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 56.6688 | 102.5026 | 642952.0649 | 7.117 | 17384.4001 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 4.1677 |  | 2153.7077 | 7.117 | 17384.4001 | y = 4960.639652 * x | 0.99845 | 11.1536 | 0 |  | 0 | 7.117 | 17384.4001 | y = 153.621749 * x | 0.99954 | 11.2189 | 62.9781 | 107.8531 | 406758.0232 | 11.2181 | 504516.0123
| UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5958 | 54.7694 |  | 86246.4642 | 5.6723 | 47307.1562 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 20.7372 |  | 71781.8794 | 6.3039 | 37723.2278 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9173 | 47.017 |  | 60896.1099 | 6.3039 | 37723.2278 | y = 484.688584 * x | 0.99801 | 7.7935 | 45.4157 | 101.5281 | 32217.5211 | 7.1166 | 17563.247 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 | 0 | 0 | 7.1166 | 17563.247 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 50.5734 | 91.4773 | 582772.9937 | 7.1166 | 17563.247 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 5.587 |  | 2916.6985 | 7.1166 | 17563.247 | y = 4960.639652 * x | 0.99845 | 11.4105 | 0 |  | 0 | 7.1166 | 17563.247 | y = 153.621749 * x | 0.99954 | 11.219 | 53.4542 | 91.5429 | 349758.6536 | 11.2181 | 511111.1467
| UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.4714 | 50.5526 |  | 59336.7053 | 5.5751 | 35262.4856 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0663 | 18.2397 |  | 45605.1221 | 6.2705 | 27240.8843 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9073 | 38.5676 |  | 35979.0563 | 6.2705 | 27240.8843 | y = 484.688584 * x | 0.99801 | 7.7939 | 50.8567 | 113.6917 | 28252.9584 | 7.1205 | 13754.1616 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.245 | 0 | 0 | 0 | 7.1205 | 13754.1616 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 58.6135 | 106.0201 | 525378.1677 | 7.1205 | 13754.1616 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 4.3203 |  | 1766.3357 | 7.1205 | 13754.1616 | y = 4960.639652 * x | 0.99845 | 11.3716 | 0 |  | 0 | 7.1205 | 13754.1616 | y = 153.621749 * x | 0.99954 | 11.2189 | 58.7453 | 100.6041 | 307408.5122 | 11.2181 | 408763.0851
| UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5725 | 398.3919 | 99.5052 | 572966.5654 | 5.6529 | 43125.2652 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 175.3527 | 92.8062 | 581711.0964 | 6.3009 | 36772.6699 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 33.5377 |  | 42143.0458 | 6.3009 | 36772.6699 | y = 484.688584 * x | 0.99801 | 7.79 | 3.3016 |  | 2297.8631 | 7.1201 | 17231.3883 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2373 | 0 |  | 397.1058 | 7.1201 | 17231.3883 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.4938 |  | 54856.83 | 7.1201 | 17231.3883 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 3.6451 |  | 1867.1011 | 7.1201 | 17231.3883 | y = 4960.639652 * x | 0.99845 | 11.3794 | 367.0031 | 91.9844 | 2614245.7392 | 7.1201 | 17231.3883 | y = 153.621749 * x | 0.99954 | 11.219 | 50.3812 | 91.0853 | 327375.3083 | 11.2181 | 507581.8712
| UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 414.2254 | 103.4599 | 633501.796 | 5.6723 | 45855.005 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 204.4494 | 108.2057 | 660811.0812 | 6.3069 | 35943.9604 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 38.909 |  | 47900.0928 | 6.3069 | 35943.9604 | y = 484.688584 * x | 0.99801 | 7.7939 | 2.5203 |  | 1658.8397 | 7.117 | 16295.4315 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 |  | 320.1417 | 7.117 | 16295.4315 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.526 |  | 52204.0001 | 7.117 | 16295.4315 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 4.1388 |  | 2004.8167 | 7.117 | 16295.4315 | y = 4960.639652 * x | 0.99845 | 11.3794 | 407.3641 | 102.1003 | 2744132.1562 | 7.117 | 16295.4315 | y = 153.621749 * x | 0.99954 | 11.2189 | 63.4059 | 114.633 | 403640.3802 | 11.2181 | 497271.4447
| UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 388.5011 | 97.0348 | 624713.6456 | 5.6723 | 48219.7539 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 187.033 | 98.988 | 642376.6085 | 6.3069 | 38121.0607 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 38.6247 |  | 50424.9148 | 6.3069 | 38121.0607 | y = 484.688584 * x | 0.99801 | 7.7904 | 2.339 |  | 1649.137 | 7.1171 | 17456.357 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2347 | 0 |  | 426.6992 | 7.1171 | 17456.357 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.2938 |  | 53393.8173 | 7.1171 | 17456.357 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 4.1529 |  | 2154.925 | 7.1171 | 17456.357 | y = 4960.639652 * x | 0.99845 | 11.4261 | 422.5855 | 105.9153 | 3049472.0421 | 7.1171 | 17456.357 | y = 153.621749 * x | 0.99954 | 11.219 | 52.1492 | 94.2817 | 376158.2799 | 11.2181 | 563445.3712
| UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042003.D | Cal | 1 | 2020-02-04 16:51 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5803 | 1.3764 | 78.6533 | 1924.8778 | 5.6606 | 42024.0239 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 1.424 | 81.3705 | 4298.432 | 6.3008 | 32827.3281 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 2.2358 | 127.7587 | 1926.1814 | 6.3008 | 32827.3281 | y = 484.688584 * x | 0.99801 | 7.78 | 3.3556 | 191.7494 | 1991.1675 | 7.1032 | 14691.0928 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 1333.9443 | 7.1032 | 14691.0928 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 | 0 | 17574.9326 | 7.1032 | 14691.0928 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6682 | 1.7303 | 98.8771 | 755.7303 | 7.1032 | 14691.0928 | y = 4960.639652 * x | 0.99845 | 11.1303 | 1.9852 | 113.438 | 12056.1079 | 7.1032 | 14691.0928 | y = 153.621749 * x | 0.99954 | 11.2267 | 2.3386 | 133.6331 | 13303.6892 | 11.2181 | 444373.7693
| UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042064.D | QC | 1 | 2020-02-05 14:32 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.553 | 1.5005 | 85.7409 | 1363.0423 | 5.5868 | 27298.1014 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0784 | 1.4551 | 83.1481 | 2704.5327 | 6.2978 | 20213.1454 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9247 | 2.5058 | 143.1886 | 1375.6295 | 6.2978 | 20213.1454 | y = 484.688584 * x | 0.99801 | 7.8008 | 1.4935 | 85.3408 | 700.5609 | 7.1275 | 11613.6729 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2347 | 0 | 0 | 0 | 7.1275 | 11613.6729 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 | 0 | 11274.4939 | 7.1275 | 11613.6729 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 3.1473 | 179.8477 | 1086.5864 | 7.1275 | 11613.6729 | y = 4960.639652 * x | 0.99845 | 11.6518 | 1.3378 | 76.4437 | 6422.5246 | 7.1275 | 11613.6729 | y = 153.621749 * x | 0.99954 | 11.2267 | 3.3532 | 191.6096 | 14496.396 | 11.2181 | 337701.7486
| UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042012.D | Cal | 10 | 2020-02-04 20:03 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.4753 | 114.6275 | 91.702 | 119699.0692 | 5.5751 | 31360.5911 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0663 | 111.1731 | 88.9385 | 247909.0494 | 6.2705 | 24543.8103 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9139 | 125.4633 | 100.3707 | 106504.4545 | 6.2705 | 24543.8103 | y = 484.688584 * x | 0.99801 | 7.7866 | 113.4019 | 90.7215 | 53223.009 | 7.1132 | 11619.7704 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 146.4702 | 117.1761 | 33721.877 | 7.1132 | 11619.7704 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.369 | 123.7442 | 98.9953 | 916034.9636 | 7.1132 | 11619.7704 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6682 | 126.174 | 100.9392 | 43343.5407 | 7.1132 | 11619.7704 | y = 4960.639652 * x | 0.99845 | 11.177 | 110.4869 | 88.3895 | 530719.216 | 7.1132 | 11619.7704 | y = 153.621749 * x | 0.99954 | 11.219 | 128.9488 | 103.159 | 573775.8807 | 11.2181 | 347579.0019
| UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042073.D | QC | 10 | 2020-02-05 17:45 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5297 | 122.6101 | 98.0881 | 140554.9228 | 5.6178 | 34425.7382 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0845 | 120.8104 | 96.6483 | 279525.1422 | 6.3038 | 25493.3563 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9281 | 129.5724 | 103.6579 | 114263.8002 | 6.3038 | 25493.3563 | y = 484.688584 * x | 0.99801 | 7.7973 | 114.0086 | 91.2069 | 52024.1479 | 7.124 | 11297.5873 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 2.5167 | 2.0133 | 3460.6149 | 7.124 | 11297.5873 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 137.4977 | 109.9982 | 987582.0896 | 7.124 | 11297.5873 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 136.8831 | 109.5065 | 45696.5658 | 7.124 | 11297.5873 | y = 4960.639652 * x | 0.99845 | 11.4961 | 112.6073 | 90.0859 | 525906.7929 | 7.124 | 11297.5873 | y = 153.621749 * x | 0.99954 | 11.2189 | 126.671 | 101.3368 | 610701.9329 | 11.2181 | 376600.4022
| UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042013.D | Cal | 11 | 2020-02-04 20:24 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5841 | 176.9886 | 88.4943 | 237248.2017 | 5.6606 | 40243.2095 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 170.0797 | 85.0398 | 474543.5724 | 6.3039 | 30910.0715 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 191.9748 | 95.9874 | 205547.0322 | 6.3039 | 30910.0715 | y = 484.688584 * x | 0.99801 | 7.7835 | 182.9148 | 91.4574 | 105439.5116 | 7.1136 | 14271.6128 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 183.6933 | 91.8466 | 50997.1277 | 7.1136 | 14271.6128 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 199.161 | 99.5805 | 1796629.761 | 7.1136 | 14271.6128 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 186.7185 | 93.3593 | 78565.9276 | 7.1136 | 14271.6128 | y = 4960.639652 * x | 0.99845 | 11.1537 | 189.7053 | 94.8527 | 1119203.2633 | 7.1136 | 14271.6128 | y = 153.621749 * x | 0.99954 | 11.2189 | 194.5538 | 97.2769 | 1086407.0578 | 11.2181 | 436195.6576
| UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042074.D | QC | 11 | 2020-02-05 18:06 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 185.6288 | 92.8144 | 241654.0013 | 5.6723 | 39080.7871 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 189.3426 | 94.6713 | 518216.9183 | 6.31 | 30385.6208 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 212.5181 | 106.2591 | 223743.7026 | 6.31 | 30385.6208 | y = 484.688584 * x | 0.99801 | 7.7974 | 171.2601 | 85.6301 | 107208.166 | 7.1205 | 15498.5133 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 11.0172 | 5.5086 | 7123.0836 | 7.1205 | 15498.5133 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 196.9061 | 98.4531 | 1929277.6331 | 7.1205 | 15498.5133 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 193.5385 | 96.7692 | 88409.2405 | 7.1205 | 15498.5133 | y = 4960.639652 * x | 0.99845 | 11.5895 | 170.1702 | 85.0851 | 1090259.9597 | 7.1205 | 15498.5133 | y = 153.621749 * x | 0.99954 | 11.219 | 197.5967 | 98.7983 | 1215524.7694 | 11.2103 | 480521.37
| UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042014.D | Cal | 12 | 2020-02-04 20:46 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6153 | 381.3088 | 101.6824 | 505376.6982 | 5.6917 | 39745.8323 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1239 | 353.1856 | 94.1828 | 1029604.8018 | 6.3069 | 32964.8471 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 390.0355 | 104.0095 | 446016.9965 | 6.3069 | 32964.8471 | y = 484.688584 * x | 0.99801 | 7.7835 | 354.4674 | 94.5246 | 221857.5907 | 7.1067 | 15495.8861 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 409.5556 | 109.2148 | 118483.4156 | 7.1067 | 15495.8861 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 386.1436 | 102.9716 | 3758538.5905 | 7.1067 | 15495.8861 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 377.3024 | 100.614 | 170896.766 | 7.1067 | 15495.8861 | y = 4960.639652 * x | 0.99845 | 11.1459 | 362.717 | 96.7245 | 2323489.4384 | 7.1067 | 15495.8861 | y = 153.621749 * x | 0.99954 | 11.219 | 393.2735 | 104.8729 | 2350939.0595 | 11.2181 | 466955.2069
| UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042075.D | QC | 12 | 2020-02-05 18:27 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 412.907 | 110.1085 | 524999.6435 | 5.6645 | 38122.8677 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1087 | 380.3587 | 101.429 | 1053847.2564 | 6.3069 | 31427.1749 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9142 | 419.6983 | 111.9196 | 457595.6901 | 6.3069 | 31427.1749 | y = 484.688584 * x | 0.99801 | 7.7973 | 313.9097 | 83.7092 | 199197.2965 | 7.1205 | 15710.7598 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 34.2289 | 9.1277 | 13796.491 | 7.1205 | 15710.7598 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 397.2524 | 105.934 | 3919547.2645 | 7.1205 | 15710.7598 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 394.2968 | 105.1458 | 180930.8573 | 7.1205 | 15710.7598 | y = 4960.639652 * x | 0.99845 | 11.5817 | 344.6497 | 91.9066 | 2238368.4775 | 7.1205 | 15710.7598 | y = 153.621749 * x | 0.99954 | 11.2112 | 395.8822 | 105.5686 | 2409381.8049 | 11.2103 | 475409.8669
| UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042015.D | Cal | 13 | 2020-02-04 21:07 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6075 | 629.2968 | 100.6875 | 915731.6523 | 5.68 | 43579.4674 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1179 | 675.2472 | 108.0395 | 1871143.8157 | 6.3039 | 32520.1082 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 648.8004 | 103.8081 | 732322.4678 | 6.3039 | 32520.1082 | y = 484.688584 * x | 0.99801 | 7.7835 | 612.5912 | 98.0146 | 372776.1231 | 7.1032 | 15065.9234 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 632.9318 | 101.2691 | 175880.8445 | 7.1032 | 15065.9234 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 646.6277 | 103.4604 | 6102787.4466 | 7.1032 | 15065.9234 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 623.0228 | 99.6837 | 271299.9212 | 7.1032 | 15065.9234 | y = 4960.639652 * x | 0.99845 | 11.1537 | 637.047 | 101.9275 | 3967561.664 | 7.1032 | 15065.9234 | y = 153.621749 * x | 0.99954 | 11.2189 | 638.9849 | 102.2376 | 3885213.4544 | 11.2181 | 474955.3685
| UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042043.D | QC | 13 | 2020-02-05 07:04 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5958 | 576.9675 | 92.3148 | 889916.3354 | 5.6723 | 46205.1348 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 630.7058 | 100.9129 | 1782570.5907 | 6.3069 | 32996.0187 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 696.7451 | 111.4792 | 797994.5784 | 6.3069 | 32996.0187 | y = 484.688584 * x | 0.99801 | 7.7904 | 627.6241 | 100.4198 | 379629.5693 | 7.117 | 14975.4163 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 250.322 | 40.0515 | 71504.4556 | 7.117 | 14975.4163 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 612.6394 | 98.0223 | 5748556.7522 | 7.117 | 14975.4163 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 616.1424 | 98.5828 | 266776.3183 | 7.117 | 14975.4163 | y = 4960.639652 * x | 0.99845 | 11.356 | 579.9451 | 92.7912 | 3590229.4625 | 7.117 | 14975.4163 | y = 153.621749 * x | 0.99954 | 11.2189 | 625.9803 | 100.1569 | 3665112.2414 | 11.2181 | 457356.7859
| UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042076.D | QC | 13 | 2020-02-05 18:49 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 719.7433 | 115.1589 | 851152.7935 | 5.6645 | 35398.619 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 712.6151 | 114.0184 | 1651767.13 | 6.31 | 27321.942 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 726.9544 | 116.3127 | 689441.4927 | 6.31 | 27321.942 | y = 484.688584 * x | 0.99801 | 7.7974 | 585.6476 | 93.7036 | 300705.2021 | 7.124 | 12712.2673 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 62.285 | 9.9656 | 17594.6493 | 7.124 | 12712.2673 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 723.733 | 115.7973 | 5760943.7389 | 7.124 | 12712.2673 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6606 | 719.5302 | 115.1248 | 263203.1673 | 7.124 | 12712.2673 | y = 4960.639652 * x | 0.99845 | 11.6907 | 528.7522 | 84.6004 | 2778636.0818 | 7.124 | 12712.2673 | y = 153.621749 * x | 0.99954 | 11.219 | 628.3617 | 100.5379 | 3477236.9552 | 11.2103 | 432268.0107
| UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042016.D | Cal | 14 | 2020-02-04 21:28 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 886.3451 | 101.2966 | 1227738.0531 | 5.6762 | 41425.4777 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 873.801 | 99.863 | 2317726.6713 | 6.3069 | 31871.6607 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 852.7496 | 97.4571 | 943524.3564 | 6.3069 | 31871.6607 | y = 484.688584 * x | 0.99801 | 7.787 | 926.2638 | 105.8587 | 528588.3273 | 7.1101 | 14128.672 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 990.5703 | 113.208 | 256055.2175 | 7.1101 | 14128.672 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 888.0624 | 101.4928 | 7851422.3846 | 7.1101 | 14128.672 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 886.0868 | 101.2671 | 357473.5745 | 7.1101 | 14128.672 | y = 4960.639652 * x | 0.99845 | 11.1537 | 914.9039 | 104.5604 | 5343591.7002 | 7.1101 | 14128.672 | y = 153.621749 * x | 0.99954 | 11.2189 | 881.5464 | 100.7482 | 5120010.8074 | 11.2181 | 453684.8102
| UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042077.D | QC | 14 | 2020-02-05 19:10 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5919 | 955.5984 | 109.2112 | 1209272.3539 | 5.6645 | 37831.2485 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 952.9115 | 108.9042 | 2342810.2447 | 6.3069 | 29825.7109 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 956.0429 | 109.262 | 989977.0774 | 6.3069 | 29825.7109 | y = 484.688584 * x | 0.99801 | 7.7973 | 874.2159 | 99.9104 | 489085.2495 | 7.124 | 13851.1011 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 80.9583 | 9.2524 | 23834.8373 | 7.124 | 13851.1011 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 947.5314 | 108.2893 | 8211104.7281 | 7.124 | 13851.1011 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 987.1941 | 112.8222 | 388602.177 | 7.124 | 13851.1011 | y = 4960.639652 * x | 0.99845 | 11.5739 | 780.0529 | 89.1489 | 4466473.8645 | 7.124 | 13851.1011 | y = 153.621749 * x | 0.99954 | 11.2189 | 882.676 | 100.8773 | 4969720.3056 | 11.2103 | 439804.0144
| UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042017.D | Cal | 15 | 2020-02-04 21:49 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5841 | 1243.561 | 99.4849 | 1732896.7351 | 5.6606 | 41594.0843 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 1239.6407 | 99.1713 | 3243254.4977 | 6.3039 | 32883.5617 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 1251.0188 | 100.0815 | 1428427.5789 | 6.3039 | 32883.5617 | y = 484.688584 * x | 0.99801 | 7.787 | 1230.9185 | 98.4735 | 705806.2659 | 7.1136 | 14196.284 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 1156.7021 | 92.5362 | 299808.6655 | 7.1136 | 14196.284 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 1227.4675 | 98.1974 | 10895228.053 | 7.1136 | 14196.284 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 1245.0594 | 99.6047 | 496267.5337 | 7.1136 | 14196.284 | y = 4960.639652 * x | 0.99845 | 11.1225 | 1224.0663 | 97.9253 | 7183499.2429 | 7.1136 | 14196.284 | y = 153.621749 * x | 0.99954 | 11.2189 | 1233.6403 | 98.6912 | 6612093.0517 | 11.2181 | 418676.8552
| UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042078.D | QC | 15 | 2020-02-05 19:31 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 1276.7089 | 102.1367 | 1811006.9776 | 5.6723 | 42332.7454 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 1410.6779 | 112.8542 | 3479023.3044 | 6.3069 | 31678.7751 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9143 | 1381.0278 | 110.4822 | 1519162.4989 | 6.3069 | 31678.7751 | y = 484.688584 * x | 0.99801 | 7.7974 | 1230.8924 | 98.4714 | 705547.8565 | 7.124 | 14191.3873 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 135.845 | 10.8676 | 38465.9934 | 7.124 | 14191.3873 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3613 | 1325.2951 | 106.0236 | 11757666.2241 | 7.124 | 14191.3873 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 1461.7026 | 116.9362 | 576447.5807 | 7.124 | 14191.3873 | y = 4960.639652 * x | 0.99845 | 11.6673 | 1212.2202 | 96.9776 | 7111526.0873 | 7.124 | 14191.3873 | y = 153.621749 * x | 0.99954 | 11.2112 | 1245.8138 | 99.6651 | 7513845.8065 | 11.2103 | 471126.658
| UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042004.D | Cal | 2 | 2020-02-04 17:13 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5842 | 2.1367 | 71.224 | 2485.995 | 5.6023 | 34962.3275 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0784 | 3.5182 | 117.2746 | 8077.2133 | 6.2827 | 24972.7331 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9247 | 3.7119 | 123.7293 | 2745.8266 | 6.2827 | 24972.7331 | y = 484.688584 * x | 0.99801 | 7.7835 | 5.3749 | 179.1648 | 2381.2628 | 7.1101 | 10968.6152 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 2421.8271 | 7.1101 | 10968.6152 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 1.1342 | 37.8051 | 25613.902 | 7.1101 | 10968.6152 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 3.9828 | 132.7608 | 1298.6083 | 7.1101 | 10968.6152 | y = 4960.639652 * x | 0.99845 | 11.177 | 3.3179 | 110.5952 | 15044.0887 | 7.1101 | 10968.6152 | y = 153.621749 * x | 0.99954 | 11.2267 | 3.4301 | 114.3372 | 15923.5231 | 11.2181 | 362625.8197
| UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042065.D | QC | 2 | 2020-02-05 14:53 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5608 | 1.9579 | 65.264 | 1964.0935 | 5.6139 | 30145.0196 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0845 | 2.9629 | 98.7645 | 5839.6559 | 6.3038 | 21437.2156 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9281 | 4.0467 | 134.8884 | 2606.3878 | 6.3038 | 21437.2156 | y = 484.688584 * x | 0.99801 | 7.7973 | 3.1358 | 104.5251 | 1375.9693 | 7.1274 | 10863.8929 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 0 | 7.1274 | 10863.8929 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0.58427 | 19.4757 | 21642.1356 | 7.1274 | 10863.8929 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 4.329 | 144.3005 | 1397.9875 | 7.1274 | 10863.8929 | y = 4960.639652 * x | 0.99845 | 11.5661 | 2.57 | 85.6661 | 11541.761 | 7.1274 | 10863.8929 | y = 153.621749 * x | 0.99954 | 11.2189 | 4.0782 | 135.9413 | 17406.6363 | 11.2181 | 333403.9132
| UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042005.D | Cal | 3 | 2020-02-04 17:34 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5802 | 4.1846 | 83.6919 | 5789.2701 | 5.6606 | 41573.1688 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1087 | 5.249 | 104.9807 | 14974.7299 | 6.3008 | 31037.7893 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 6.3342 | 126.6839 | 6240.0571 | 6.3008 | 31037.7893 | y = 484.688584 * x | 0.99801 | 7.78 | 7.4961 | 149.9219 | 4194.0148 | 7.1067 | 13852.0238 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 2955.6073 | 7.1067 | 13852.0238 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 2.8616 | 57.2328 | 47277.1643 | 7.1067 | 13852.0238 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 5.6642 | 113.284 | 2332.1396 | 7.1067 | 13852.0238 | y = 4960.639652 * x | 0.99845 | 11.1303 | 5.4141 | 108.2811 | 31002.1989 | 7.1067 | 13852.0238 | y = 153.621749 * x | 0.99954 | 11.2189 | 5.962 | 119.2396 | 33192.094 | 11.2181 | 434883.2637
| UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042023.D | QC | 3 | 2020-02-04 23:57 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 4.5875 | 91.7508 | 6073.5825 | 5.6762 | 39783.8694 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 4.7994 | 95.9888 | 13639.5016 | 6.3039 | 30917.0632 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 5.7548 | 115.0964 | 5593.5388 | 6.3039 | 30917.0632 | y = 484.688584 * x | 0.99801 | 7.787 | 6.5028 | 130.0563 | 3666.8465 | 7.1067 | 13960.7816 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 1118.1915 | 7.1067 | 13960.7816 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 2.4454 | 48.9076 | 44022.5538 | 7.1067 | 13960.7816 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 5.7429 | 114.8583 | 2383.1053 | 7.1067 | 13960.7816 | y = 4960.639652 * x | 0.99845 | 11.2081 | 5.521 | 110.4201 | 31862.8245 | 7.1067 | 13960.7816 | y = 153.621749 * x | 0.99954 | 11.2189 | 6.3864 | 127.7282 | 36809.9541 | 11.2181 | 450232.7156
| UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042066.D | QC | 3 | 2020-02-05 15:15 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.553 | 6.018 | 120.3603 | 6982.7263 | 5.6412 | 34866.664 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0967 | 5.7719 | 115.4377 | 13232.2199 | 6.3069 | 24943.155 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9316 | 8.0407 | 160.8147 | 6493.4267 | 6.3069 | 24943.155 | y = 484.688584 * x | 0.99801 | 7.7974 | 4.9202 | 98.4049 | 2480.7714 | 7.124 | 12482.98 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 | 0 | 338.4467 | 7.124 | 12482.98 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 3.1872 | 63.7431 | 45139.8415 | 7.124 | 12482.98 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6606 | 6.5627 | 131.2532 | 2434.913 | 7.124 | 12482.98 | y = 4960.639652 * x | 0.99845 | 11.4338 | 4.6826 | 93.653 | 24163.8618 | 7.124 | 12482.98 | y = 153.621749 * x | 0.99954 | 11.219 | 6.1035 | 122.071 | 32303.8274 | 11.2181 | 413428.1999
| UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042006.D | Cal | 4 | 2020-02-04 17:55 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6153 | 7.9142 | 105.5222 | 12053.4071 | 5.6879 | 45765.5533 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.124 | 6.9622 | 92.829 | 21085.9763 | 6.3069 | 32956.5043 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 7.8679 | 104.9051 | 8381.646 | 6.3069 | 32956.5043 | y = 484.688584 * x | 0.99801 | 7.7835 | 9.4758 | 126.3445 | 5754.8286 | 7.0997 | 15036.032 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 3165.0566 | 7.0997 | 15036.032 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 4.9979 | 66.6383 | 71358.8197 | 7.0997 | 15036.032 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 6.7359 | 89.812 | 3010.3065 | 7.0997 | 15036.032 | y = 4960.639652 * x | 0.99845 | 11.177 | 7.3039 | 97.3851 | 45398.721 | 7.0997 | 15036.032 | y = 153.621749 * x | 0.99954 | 11.219 | 7.9964 | 106.6188 | 47080.1892 | 11.2181 | 459908.6328
| UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042067.D | QC | 4 | 2020-02-05 15:36 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5686 | 6.9137 | 92.1829 | 9121.8938 | 5.6528 | 39646.9403 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 6.7556 | 90.0742 | 18795.9301 | 6.3099 | 30275.051 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9142 | 8.0532 | 107.3756 | 7894.5489 | 6.3099 | 30275.051 | y = 484.688584 * x | 0.99801 | 7.7973 | 6.9345 | 92.4594 | 4295.4841 | 7.124 | 15336.2093 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 382.9384 | 7.124 | 15336.2093 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 5.2449 | 69.9326 | 75147.5734 | 7.124 | 15336.2093 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 7.5262 | 100.3494 | 3430.5264 | 7.124 | 15336.2093 | y = 4960.639652 * x | 0.99845 | 11.4105 | 5.5561 | 74.0812 | 35224.4004 | 7.124 | 15336.2093 | y = 153.621749 * x | 0.99954 | 11.2189 | 7.8906 | 105.2086 | 46554.309 | 11.2181 | 460867.2116
| UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042007.D | Cal | 5 | 2020-02-04 18:17 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6113 | 13.2575 | 106.06 | 19725.4415 | 5.6878 | 44708.1517 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1239 | 10.7181 | 85.7445 | 32057.6607 | 6.3038 | 32560.0815 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 13.328 | 106.6238 | 14456.5849 | 6.3038 | 32560.0815 | y = 484.688584 * x | 0.99801 | 7.7835 | 12.967 | 103.7356 | 8247.8008 | 7.0997 | 15747.7567 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 0 | 0 | 4060.1317 | 7.0997 | 15747.7567 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 8.8585 | 70.8681 | 112668.7312 | 7.0997 | 15747.7567 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 10.792 | 86.3358 | 5050.3674 | 7.0997 | 15747.7567 | y = 4960.639652 * x | 0.99845 | 11.1848 | 11.2783 | 90.2267 | 73420.9929 | 7.0997 | 15747.7567 | y = 153.621749 * x | 0.99954 | 11.2189 | 12.0792 | 96.634 | 74076.4525 | 11.2181 | 479036.7769
| UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042068.D | QC | 5 | 2020-02-05 15:58 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 12.8271 | 102.6169 | 16980.4693 | 5.6645 | 39778.031 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 11.5323 | 92.2581 | 33954.7173 | 6.3069 | 32054.8952 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9143 | 14.0349 | 112.2795 | 15019.5154 | 6.3069 | 32054.8952 | y = 484.688584 * x | 0.99801 | 7.7939 | 11.4703 | 91.7623 | 7302.697 | 7.1171 | 15762.5805 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 707.3405 | 7.1171 | 15762.5805 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 10.0727 | 80.5813 | 124715.5162 | 7.1171 | 15762.5805 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6606 | 12.7557 | 102.0453 | 5974.4143 | 7.1171 | 15762.5805 | y = 4960.639652 * x | 0.99845 | 11.4105 | 9.8313 | 78.6505 | 64061.2521 | 7.1171 | 15762.5805 | y = 153.621749 * x | 0.99954 | 11.219 | 13.0847 | 104.6774 | 83043.0255 | 11.2181 | 495756.9008
| UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042008.D | Cal | 6 | 2020-02-04 18:38 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5725 | 15.7423 | 78.7113 | 20613.0995 | 5.6529 | 39345.1842 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1057 | 17.0266 | 85.133 | 46449.379 | 6.2978 | 29718.0841 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 19.5619 | 97.8094 | 19630.3182 | 6.2978 | 29718.0841 | y = 484.688584 * x | 0.99801 | 7.7835 | 19.6888 | 98.4441 | 10757.5045 | 7.1101 | 13527.271 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 14.7462 | 73.7311 | 7126.7064 | 7.1101 | 13527.271 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 17.5255 | 87.6273 | 169930.6101 | 7.1101 | 13527.271 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 17.794 | 88.9701 | 7150.7451 | 7.1101 | 13527.271 | y = 4960.639652 * x | 0.99845 | 11.1303 | 18.7941 | 93.9707 | 105096.6566 | 7.1101 | 13527.271 | y = 153.621749 * x | 0.99954 | 11.219 | 19.5135 | 97.5677 | 107821.4624 | 11.2181 | 431616.0593
| UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042051.D | QC | 6 | 2020-02-05 09:55 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5841 | 17.9745 | 89.8723 | 26275.7915 | 5.6606 | 43924.8288 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1057 | 16.5867 | 82.9335 | 53483.6638 | 6.3039 | 35124.4078 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9143 | 20.377 | 101.8852 | 24196.0934 | 6.3039 | 35124.4078 | y = 484.688584 * x | 0.99801 | 7.7904 | 17.7858 | 88.9291 | 11804.3406 | 7.1136 | 16431.8493 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 1765.3153 | 7.1136 | 16431.8493 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 17.7019 | 88.5094 | 208226.9403 | 7.1136 | 16431.8493 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 19.2883 | 96.4414 | 9414.958 | 7.1136 | 16431.8493 | y = 4960.639652 * x | 0.99845 | 11.4027 | 18.2137 | 91.0683 | 123720.0958 | 7.1136 | 16431.8493 | y = 153.621749 * x | 0.99954 | 11.2189 | 19.8254 | 99.127 | 123569.8913 | 11.2181 | 486877.1563
| UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042069.D | QC | 6 | 2020-02-05 16:19 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 20.5871 | 102.9357 | 27657.149 | 5.6606 | 40365.9838 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1057 | 18.6463 | 93.2315 | 51456.5563 | 6.3069 | 30067.2479 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9143 | 21.2542 | 106.2708 | 21628.4993 | 6.3069 | 30067.2479 | y = 484.688584 * x | 0.99801 | 7.7973 | 20.8962 | 104.4811 | 11300.3788 | 7.1205 | 13388.8649 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 | 0 | 1065.9844 | 7.1205 | 13388.8649 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 20.0683 | 100.3414 | 189433.6383 | 7.1205 | 13388.8649 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 22.1541 | 110.7706 | 8810.0994 | 7.1205 | 13388.8649 | y = 4960.639652 * x | 0.99845 | 11.4027 | 16.5686 | 82.8428 | 91703.2683 | 7.1205 | 13388.8649 | y = 153.621749 * x | 0.99954 | 11.2189 | 19.7105 | 98.5525 | 119132.3322 | 11.2181 | 472128.959
| UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042009.D | Cal | 7 | 2020-02-04 18:59 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6036 | 27.7998 | 88.9594 | 38651.931 | 5.6801 | 41774.9961 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1209 | 26.9485 | 86.2352 | 81429.7279 | 6.3039 | 32952.4017 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 28.9161 | 92.5314 | 32474.5134 | 6.3039 | 32952.4017 | y = 484.688584 * x | 0.99801 | 7.7835 | 31.942 | 102.2145 | 18858.7095 | 7.1066 | 14617.3179 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 17.2311 | 55.1394 | 8355.9507 | 7.1066 | 14617.3179 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 27.9128 | 89.3211 | 278357.3719 | 7.1066 | 14617.3179 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 28.6876 | 91.8004 | 12451.395 | 7.1066 | 14617.3179 | y = 4960.639652 * x | 0.99845 | 11.177 | 27.1538 | 86.892 | 164079.3874 | 7.1066 | 14617.3179 | y = 153.621749 * x | 0.99954 | 11.2189 | 30.5489 | 97.7564 | 177042.5341 | 11.2181 | 452700.6733
| UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042070.D | QC | 7 | 2020-02-05 16:40 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5725 | 28.3056 | 90.5778 | 37937.1058 | 5.6529 | 40269.6793 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 31.0614 | 99.3966 | 81842.1581 | 6.3069 | 28746.775 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 36.8934 | 118.0589 | 36296.1002 | 6.3069 | 28746.775 | y = 484.688584 * x | 0.99801 | 7.7939 | 24.4065 | 78.101 | 15402.3791 | 7.1205 | 15624.2761 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 1345.2635 | 7.1205 | 15624.2761 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 27.8293 | 89.0537 | 296718.3207 | 7.1205 | 15624.2761 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 29.3327 | 93.8648 | 13608.0402 | 7.1205 | 15624.2761 | y = 4960.639652 * x | 0.99845 | 11.5039 | 26.2897 | 84.127 | 169801.643 | 7.1205 | 15624.2761 | y = 153.621749 * x | 0.99954 | 11.219 | 30.736 | 98.3551 | 190220.4247 | 11.2181 | 483435.5795
| UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042010.D | Cal | 8 | 2020-02-04 19:20 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6075 | 44.1379 | 88.2758 | 58979.9177 | 5.684 | 40145.8739 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1209 | 42.4548 | 84.9095 | 131579.8875 | 6.3039 | 33856.1378 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 44.9769 | 89.9538 | 52254.2264 | 6.3039 | 33856.1378 | y = 484.688584 * x | 0.99801 | 7.7835 | 46.9547 | 93.9093 | 28477.5555 | 7.1032 | 15015.5888 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 42.5134 | 85.0269 | 15429.1763 | 7.1032 | 15015.5888 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 46.7498 | 93.4997 | 462416.8791 | 7.1032 | 15015.5888 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 46.1738 | 92.3477 | 20570.9364 | 7.1032 | 15015.5888 | y = 4960.639652 * x | 0.99845 | 11.2004 | 44.1978 | 88.3957 | 274346.685 | 7.1032 | 15015.5888 | y = 153.621749 * x | 0.99954 | 11.219 | 49.7147 | 99.4294 | 283653.6312 | 11.2181 | 445689.3513
| UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042071.D | QC | 8 | 2020-02-05 17:02 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5841 | 52.2273 | 104.4546 | 68570.5169 | 5.6645 | 39442.9265 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 46.8528 | 93.7055 | 135742.2627 | 6.3099 | 31663.801 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 53.8189 | 107.6377 | 58596.1317 | 6.3099 | 31663.801 | y = 484.688584 * x | 0.99801 | 7.7939 | 43.1534 | 86.3069 | 29386.7027 | 7.124 | 16859.8566 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 1927.8257 | 7.124 | 16859.8566 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 45.5409 | 91.0818 | 506495.5691 | 7.124 | 16859.8566 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6605 | 46.3045 | 92.609 | 23162.747 | 7.124 | 16859.8566 | y = 4960.639652 * x | 0.99845 | 11.4728 | 44.7786 | 89.5571 | 312090.4724 | 7.124 | 16859.8566 | y = 153.621749 * x | 0.99954 | 11.219 | 50.2157 | 100.4314 | 323899.2897 | 11.2181 | 503847.545
| UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042011.D | Cal | 9 | 2020-02-04 19:42 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5919 | 81.1917 | 92.7905 | 107328.7844 | 5.6684 | 39706.8932 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 74.4576 | 85.0944 | 213351.5386 | 6.3008 | 31411.0942 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 82.353 | 94.1177 | 89263.7616 | 6.3008 | 31411.0942 | y = 484.688584 * x | 0.99801 | 7.7835 | 80.7788 | 92.3186 | 47593.4172 | 7.1066 | 14587.0677 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2244 | 78.6815 | 89.9218 | 24502.3981 | 7.1066 | 14587.0677 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 84.303 | 96.3463 | 790998.548 | 7.1066 | 14587.0677 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 83.8457 | 95.8237 | 36226.9094 | 7.1066 | 14587.0677 | y = 4960.639652 * x | 0.99845 | 11.1614 | 74.8291 | 85.519 | 451226.9336 | 7.1066 | 14587.0677 | y = 153.621749 * x | 0.99954 | 11.2189 | 84.7464 | 96.8531 | 463082.635 | 11.2181 | 426840.4477
| UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042032.D | QC | 9 | 2020-02-05 03:09 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5842 | 87.8897 | 100.4454 | 105681.257 | 5.6606 | 36116.4741 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1057 | 74.6243 | 85.2849 | 204019.8123 | 6.3039 | 29970.6563 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 82.8031 | 94.6321 | 85638.9248 | 6.3039 | 29970.6563 | y = 484.688584 * x | 0.99801 | 7.7904 | 76.2612 | 87.1556 | 45596.4639 | 7.1171 | 14802.8796 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 38.88 | 44.4343 | 14240.7394 | 7.1171 | 14802.8796 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 76.2257 | 87.1151 | 728100.5971 | 7.1171 | 14802.8796 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 77.9332 | 89.0666 | 34179.5641 | 7.1171 | 14802.8796 | y = 4960.639652 * x | 0.99845 | 11.4494 | 64.2902 | 73.4745 | 393411.7728 | 7.1171 | 14802.8796 | y = 153.621749 * x | 0.99954 | 11.219 | 84.0148 | 96.017 | 437289.3082 | 11.2181 | 406575.6376
| UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042072.D | QC | 9 | 2020-02-05 17:23 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5764 | 81.6981 | 93.3693 | 109984.7665 | 5.6529 | 40437.1372 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 90.0383 | 102.9009 | 220929.3725 | 6.3069 | 26944.2104 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9143 | 100.6233 | 114.998 | 93670.6283 | 6.3069 | 26944.2104 | y = 484.688584 * x | 0.99801 | 7.7939 | 69.3701 | 79.2802 | 43651.2042 | 7.124 | 15579.0894 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 2929.0637 | 7.124 | 15579.0894 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 76.895 | 87.88 | 772784.7526 | 7.124 | 15579.0894 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6606 | 76.7134 | 87.6724 | 35410.6909 | 7.124 | 15579.0894 | y = 4960.639652 * x | 0.99845 | 11.465 | 61.2103 | 69.9547 | 394205.9529 | 7.124 | 15579.0894 | y = 153.621749 * x | 0.99954 | 11.219 | 86.55 | 98.9143 | 471960.6683 | 11.2181 | 425958.6316
| UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1797 | 0 |  | 0 | 5.6606 | 42279.1923 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0936 | 0 | 0 | 0 | 6.2948 | 34299.1518 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 270.4607 | 99.8687 | 321598.6557 | 6.2948 | 34299.1518 | y = 484.688584 * x | 0.99801 | 7.7835 | 1.3726 |  | 908.5191 | 7.1032 | 16387.1308 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2218 | 122.0789 |  | 40349.7739 | 7.1032 | 16387.1308 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1907.6094 | 7.1032 | 16387.1308 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 251.357 | 103.4955 | 121087.9193 | 7.1032 | 16387.1308 | y = 4960.639652 * x | 0.99845 | 11.2937 | 0 | 0 | 0 | 7.1032 | 16387.1308 | y = 153.621749 * x | 0.99954 | 11.2189 | 236.1243 | 100.5236 | 1559477.053 | 11.2181 | 515901.6354
| UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1642 | 0 |  | 0 | 5.6529 | 45252.3426 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0967 | 0 | 0 | 0 | 6.2978 | 32412.4363 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 275.029 | 101.5555 | 309051.8833 | 6.2978 | 32412.4363 | y = 484.688584 * x | 0.99801 | 7.79 | 1.3002 |  | 826.6672 | 7.1132 | 15740.8529 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 120.5983 |  | 38338.1948 | 7.1132 | 15740.8529 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1807.4914 | 7.1132 | 15740.8529 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 236.3069 | 97.2986 | 109422.5473 | 7.1132 | 15740.8529 | y = 4960.639652 * x | 0.99845 | 11.1848 | 0 | 0 | 0 | 7.1132 | 15740.8529 | y = 153.621749 * x | 0.99954 | 11.219 | 239.3649 | 101.9032 | 1494748.7556 | 11.2181 | 487793.8886
| UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.2225 | 0 |  | 0 | 5.6684 | 48899.4333 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 0 | 0 | 0 | 6.3069 | 38512.1517 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 266.9594 | 98.5758 | 356416.7453 | 6.3069 | 38512.1517 | y = 484.688584 * x | 0.99801 | 7.7939 | 0.9351 |  | 627.4642 | 7.117 | 16613.0802 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 28.9607 |  | 13010.678 | 7.117 | 16613.0802 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 12993.5192 | 7.117 | 16613.0802 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 240.9389 | 99.2059 | 117724.9479 | 7.117 | 16613.0802 | y = 4960.639652 * x | 0.99845 | 11.4183 | 1.9831 |  | 13619.0514 | 7.117 | 16613.0802 | y = 153.621749 * x | 0.99954 | 11.2189 | 229.1937 | 97.5731 | 1583252.3854 | 11.2181 | 539605.0373
| UCGCAAT1hS1C_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042054.D | Sample |  | 2020-02-05 10:59 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.2303 | 0 |  | 0 | 5.6723 | 48880.2301 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0997 | 0 | 0 | 0 | 6.3069 | 36219.3641 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 286.0904 | 105.64 | 359268.2578 | 6.3069 | 36219.3641 | y = 484.688584 * x | 0.99801 | 7.7939 | 0.77705 |  | 506.8878 | 7.1171 | 16150.3833 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 31.2286 |  | 13308.774 | 7.1171 | 16150.3833 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 0 |  | 1442.7784 | 7.1171 | 16150.3833 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 237.8881 | 97.9497 | 113012.5987 | 7.1171 | 16150.3833 | y = 4960.639652 * x | 0.99845 | 11.1848 | 0 | 0 | 0 | 7.1171 | 16150.3833 | y = 153.621749 * x | 0.99954 | 11.219 | 228.6078 | 101.8885 | 1527886.6507 | 11.2181 | 522069.9328
| UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1952 | 0 |  | 0 | 5.6567 | 45429.6893 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 0 | 0 | 0 | 6.3039 | 33035.4225 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 0 | 0 | 0 | 6.3039 | 33035.4225 | y = 484.688584 * x | 0.99801 | 7.7939 | 298.8573 | 105.1663 | 182057.7361 | 7.1205 | 15082.165 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 0 | 7.1205 | 15082.165 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 268.5056 | 112.1949 | 2551207.2326 | 7.1205 | 15082.165 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3655 | 0 | 0 | 0 | 7.1205 | 15082.165 | y = 4960.639652 * x | 0.99845 | 11.2315 | 0 | 0 | 0 | 7.1205 | 15082.165 | y = 153.621749 * x | 0.99954 | 11.2189 | 228.357 | 101.7768 | 1268681.9773 | 11.2181 | 433977.2416
| UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1991 | 0 |  | 0 | 5.6645 | 42273.4175 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0966 | 0 | 0 | 0 | 6.3008 | 35240.2399 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 0 | 0 | 0 | 6.3008 | 35240.2399 | y = 484.688584 * x | 0.99801 | 7.7835 | 276.4869 | 97.2942 | 183688.8077 | 7.1066 | 16448.5129 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 476.9612 | 7.1066 | 16448.5129 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 220.7353 | 92.234 | 2292083.3033 | 7.1066 | 16448.5129 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3655 | 0 | 0 | 0 | 7.1066 | 16448.5129 | y = 4960.639652 * x | 0.99845 | 11.1614 | 0 | 0 | 0 | 7.1066 | 16448.5129 | y = 153.621749 * x | 0.99954 | 11.2189 | 216.1466 | 96.3347 | 1341527.1913 | 11.2181 | 484818.9148
| UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6386 | 0 |  | 0 | 5.6451 | 40772.8976 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0936 | 0 | 0 | 0 | 6.2948 | 33685.8517 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 0 | 0 | 0 | 6.2948 | 33685.8517 | y = 484.688584 * x | 0.99801 | 7.7866 | 277.1838 | 97.5395 | 181926.1609 | 7.1132 | 16249.7142 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 280.6394 | 7.1132 | 16249.7142 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 228.7215 | 95.5711 | 2345349.6761 | 7.1132 | 16249.7142 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3656 | 0 | 0 | 0 | 7.1132 | 16249.7142 | y = 4960.639652 * x | 0.99845 | 11.0837 | 0 | 0 | 0 | 7.1132 | 16249.7142 | y = 153.621749 * x | 0.99954 | 11.219 | 211.3997 | 94.1801 | 1347459.2993 | 11.2181 | 497897.4822
| UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5919 | 205.0918 | 106.2496 | 262299.4346 | 5.6684 | 38389.9729 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 228.1491 | 108.8919 | 596959.8831 | 6.3039 | 29174.7978 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.89 | 0 | 0 | 0 | 6.3039 | 29174.7978 | y = 484.688584 * x | 0.99801 | 7.7904 | 1.3468 |  | 836.4007 | 7.1136 | 15375.4789 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 |  | 380.6254 | 7.1136 | 15375.4789 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1968.5507 | 7.1136 | 15375.4789 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.5364 | 0 | 0 | 0 | 7.1136 | 15375.4789 | y = 4960.639652 * x | 0.99845 | 11.2159 | 238.0729 | 105.4447 | 1513195.7042 | 7.1136 | 15375.4789 | y = 153.621749 * x | 0.99954 | 11.2189 | 227.5675 | 101.2137 | 1251495.903 | 11.2181 | 429583.7135
| UCGCAAT1hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042026.D | Sample |  | 2020-02-05 01:01 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 179.8831 | 93.1901 | 270871.3112 | 5.6762 | 45206.4963 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 207.9551 | 99.2536 | 620020.0005 | 6.3039 | 33169.5863 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0807 | 0 | 0 | 0 | 6.3039 | 33169.5863 | y = 484.688584 * x | 0.99801 | 7.787 | 1.1653 |  | 731.2054 | 7.1101 | 15535.976 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 |  | 211.2431 | 7.1101 | 15535.976 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1509.7103 | 7.1101 | 15535.976 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.5338 | 0 | 0 | 0 | 7.1101 | 15535.976 | y = 4960.639652 * x | 0.99845 | 11.2237 | 234.7505 | 103.9732 | 1507653.5931 | 7.1101 | 15535.976 | y = 153.621749 * x | 0.99954 | 11.219 | 234.4226 | 104.2627 | 1286768.304 | 11.2181 | 428774.9495
| UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 192.5821 | 99.7689 | 299018.5217 | 5.6762 | 46610.1305 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 206.0158 | 98.328 | 651363.2661 | 6.3008 | 35166.8058 | y = 416.856543 * x - 0.018990 | 0.99928 | 7.0876 | 0 | 0 | 0 | 6.3008 | 35166.8058 | y = 484.688584 * x | 0.99801 | 7.787 | 1.2346 |  | 821.5587 | 7.1067 | 16474.9308 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 |  | 306.5811 | 7.1067 | 16474.9308 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1203.3863 | 7.1067 | 16474.9308 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.5313 | 0 | 0 | 0 | 7.1067 | 16474.9308 | y = 4960.639652 * x | 0.99845 | 11.3171 | 242.0251 | 107.1951 | 1648315.7882 | 7.1067 | 16474.9308 | y = 153.621749 * x | 0.99954 | 11.2189 | 225.7573 | 100.4087 | 1438599.7058 | 11.2181 | 497767.6323
| UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5492 | 181.4107 | 93.9815 | 281804.2735 | 5.6295 | 46634.702 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0936 | 194.392 | 92.7802 | 623589.8648 | 6.2918 | 35634.3775 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9281 | 0 | 0 | 0 | 6.2918 | 35634.3775 | y = 484.688584 * x | 0.99801 | 7.7904 | 0.86635 |  | 637.1618 | 7.1205 | 18208.5429 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 0 | 7.1205 | 18208.5429 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 9095.6218 | 7.1205 | 18208.5429 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.7931 | 0 | 0 | 0 | 7.1205 | 18208.5429 | y = 4960.639652 * x | 0.99845 | 11.3716 | 197.2418 | 87.3602 | 1484672.4604 | 7.1205 | 18208.5429 | y = 153.621749 * x | 0.99954 | 11.219 | 211.6066 | 94.1149 | 1420407.5714 | 11.2181 | 524339.2574
| UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.0786 | 0 |  | 0 | 5.6256 | 49525.8927 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1391 | 0 | 0 | 0 | 6.2735 | 38036.2844 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 264.5346 | 104.4625 | 348808.9109 | 6.2735 | 38036.2844 | y = 484.688584 * x | 0.99801 | 7.7973 | 0.47251 |  | 360.3432 | 7.1205 | 18880.9947 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 2.3472 |  | 5725.8095 | 7.1205 | 18880.9947 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 10898.2122 | 7.1205 | 18880.9947 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 175.5438 | 101.7758 | 97769.3431 | 7.1205 | 18880.9947 | y = 4960.639652 * x | 0.99845 | 11.3015 | 0 | 0 | 0 | 7.1205 | 18880.9947 | y = 153.621749 * x | 0.99954 | 11.2189 | 200.2697 | 101.7544 | 1446510.0873 | 11.2181 | 564202.1179
| UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.2264 | 0 |  | 0 | 5.6684 | 45385.4767 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0997 | 0 | 0 | 0 | 6.3009 | 35474.482 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 249.1974 | 98.406 | 306415.8203 | 6.3009 | 35474.482 | y = 484.688584 * x | 0.99801 | 7.7904 | 0.77134 |  | 533.1338 | 7.1101 | 17112.4072 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 13.6111 |  | 8665.2176 | 7.1101 | 17112.4072 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 0 |  | 1009.324 | 7.1101 | 17112.4072 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6657 | 173.8982 | 100.8217 | 87787.144 | 7.1101 | 17112.4072 | y = 4960.639652 * x | 0.99845 | 11.3249 | 0 | 0 | 0 | 7.1101 | 17112.4072 | y = 153.621749 * x | 0.99954 | 11.219 | 197.845 | 100.5224 | 1199756.8571 | 11.2181 | 473692.7868
| UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.0824 | 0 |  | 0 | 5.6256 | 44371.295 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0845 | 0 | 0 | 0 | 6.2826 | 31675.7212 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9246 | 245.97 | 97.1315 | 270052.2341 | 6.2826 | 31675.7212 | y = 484.688584 * x | 0.99801 | 7.7939 | 1.065 |  | 679.5872 | 7.117 | 15797.7788 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 18.8334 |  | 9487.2043 | 7.117 | 15797.7788 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 0 |  | 3914.4483 | 7.117 | 15797.7788 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 168.0008 | 97.4025 | 78315.4156 | 7.117 | 15797.7788 | y = 4960.639652 * x | 0.99845 | 11.177 | 0 | 0 | 0 | 7.117 | 15797.7788 | y = 153.621749 * x | 0.99954 | 11.2189 | 192.3357 | 97.7232 | 1079328.812 | 11.2181 | 438351.4479
| UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6502 | 0 |  | 0 | 5.6528 | 46946.472 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0966 | 0 | 0 | 0 | 6.2978 | 38576.7512 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9697 | 0 | 0 | 0 | 6.2978 | 38576.7512 | y = 484.688584 * x | 0.99801 | 7.7904 | 273.6249 | 107.7037 | 199725.3272 | 7.1136 | 18071.5743 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 159.9197 | 7.1136 | 18071.5743 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 219.9468 | 105.5902 | 2509365.1856 | 7.1136 | 18071.5743 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3655 | 0 | 0 | 0 | 7.1136 | 18071.5743 | y = 4960.639652 * x | 0.99845 | 11.3482 | 0 | 0 | 0 | 7.1136 | 18071.5743 | y = 153.621749 * x | 0.99954 | 11.2189 | 182.8837 | 101.971 | 1335265.836 | 11.2181 | 570323.466
| UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.1681 | 0 |  | 0 | 5.6529 | 44216.1553 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0997 | 0.31201 |  | 1040.7094 | 6.3009 | 36269.1837 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9247 | 0 | 0 | 0 | 6.3009 | 36269.1837 | y = 484.688584 * x | 0.99801 | 7.7939 | 261.3487 | 102.8716 | 184554.768 | 7.1171 | 17483.3019 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 288.742 | 7.1171 | 17483.3019 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 221.197 | 106.1904 | 2441316.5818 | 7.1171 | 17483.3019 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3656 | 0 | 0 | 0 | 7.1171 | 17483.3019 | y = 4960.639652 * x | 0.99845 | 11.2082 | 0 | 0 | 0 | 7.1171 | 17483.3019 | y = 153.621749 * x | 0.99954 | 11.219 | 188.2681 | 104.9731 | 1211653.6671 | 11.2181 | 502724.7548
| UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.273 | 0 |  | 0 | 5.684 | 48159.5878 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 0 | 0 | 0 | 6.3069 | 36298.6945 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9242 | 0 | 0 | 0 | 6.3069 | 36298.6945 | y = 484.688584 * x | 0.99801 | 7.79 | 227.1863 | 89.4247 | 161252.8244 | 7.1132 | 17572.9112 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 232.3474 | 7.1132 | 17572.9112 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 183.763 | 88.2194 | 2043398.9891 | 7.1132 | 17572.9112 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.3656 | 0 | 0 | 0 | 7.1132 | 17572.9112 | y = 4960.639652 * x | 0.99845 | 11.3171 | 1.9487 |  | 14155.8163 | 7.1132 | 17572.9112 | y = 153.621749 * x | 0.99954 | 11.219 | 166.8947 | 93.0559 | 1046225.116 | 11.2181 | 489678.7122
| UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6036 | 165.9491 | 105.6238 | 249607.3666 | 5.6762 | 45158.9158 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 195.8233 | 105.0488 | 593672.5283 | 6.3069 | 33682.1857 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 0 | 0 | 0 | 6.3069 | 33682.1857 | y = 484.688584 * x | 0.99801 | 7.7939 | 0.85785 |  | 553.7791 | 7.1136 | 15982.4355 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2398 | 0 | 0 | 0 | 7.1136 | 15982.4355 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 0 |  | 2154.8638 | 7.1136 | 15982.4355 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.5364 | 0 | 0 | 0 | 7.1136 | 15982.4355 | y = 4960.639652 * x | 0.99845 | 11.3482 | 210.2326 | 96.7961 | 1388991.0846 | 7.1136 | 15982.4355 | y = 153.621749 * x | 0.99954 | 11.2189 | 187.6978 | 101.2923 | 1119892.7533 | 11.2181 | 466064.303
| UCGCAAT5hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042046.D | Sample |  | 2020-02-05 08:08 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.588 | 174.488 | 111.0586 | 244841.3933 | 5.6645 | 42126.9768 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1088 | 178.636 | 95.8288 | 562220.569 | 6.3069 | 34900.0627 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.7097 | 0 | 0 | 0 | 6.3069 | 34900.0627 | y = 484.688584 * x | 0.99801 | 7.7974 | 0.67992 |  | 439.4395 | 7.1171 | 16001.3802 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 | 0 | 0 | 7.1171 | 16001.3802 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1227.8028 | 7.1171 | 16001.3802 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.5338 | 0 | 0 | 0 | 7.1171 | 16001.3802 | y = 4960.639652 * x | 0.99845 | 11.3482 | 212.2557 | 97.7276 | 1404019.5064 | 7.1171 | 16001.3802 | y = 153.621749 * x | 0.99954 | 11.219 | 187.3281 | 101.0928 | 1104268.2588 | 11.2181 | 460468.8876
| UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5686 | 152.8038 | 79.1614 | 228990.1946 | 5.649 | 44996.0952 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 182.8059 | 87.2503 | 595165.7019 | 6.2948 | 36119.1322 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9139 | 0 | 0 | 0 | 6.2948 | 36119.1322 | y = 484.688584 * x | 0.99801 | 7.7866 | 0.96546 |  | 630.86 | 7.1097 | 16177.7404 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 311.0805 | 7.1097 | 16177.7404 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 846.6326 | 7.1097 | 16177.7404 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.7906 | 0 | 0 | 0 | 7.1097 | 16177.7404 | y = 4960.639652 * x | 0.99845 | 11.3949 | 240.5697 | 106.5505 | 1608848.6635 | 7.1097 | 16177.7404 | y = 153.621749 * x | 0.99954 | 11.219 | 182.2168 | 98.3345 | 1154335.3982 | 11.2181 | 494848.2367
| UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5764 | 152.5874 | 79.0493 | 222741.8559 | 5.6529 | 43830.4305 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 180.6059 | 86.2003 | 531934.4626 | 6.3009 | 32667.0296 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9073 | 0 | 0 | 0 | 6.3009 | 32667.0296 | y = 484.688584 * x | 0.99801 | 7.7939 | 0.73126 |  | 488.1454 | 7.1171 | 16527.163 | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 177.5834 | 7.1171 | 16527.163 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 0 |  | 1462.4472 | 7.1171 | 16527.163 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 0 | 0 | 0 | 7.1171 | 16527.163 | y = 4960.639652 * x | 0.99845 | 11.3171 | 200.7712 | 88.9234 | 1371690.8616 | 7.1171 | 16527.163 | y = 153.621749 * x | 0.99954 | 11.219 | 183.9699 | 99.2805 | 1044674.4162 | 11.2181 | 443570.4903

## 3125Data
[5 empty] | 3125 Method |  | 3125 Results |  |  |  | 4NT Method |  | 4NT Results
Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy
ACN | 7010_2_03032000.D | Blank |  | 2020-03-03 10:01 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.2734 | 0 |  |  | y = 153.340196 * x + 0.007796 | 0.99968
SpikeBlank | 7010_2_03032001.D | MatrixBlank |  | 2020-03-03 10:22 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 198.932 | 0.98065 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 1742.944 | 1.362
SpikeBlank | 7010_2_03032002.D | MatrixBlank |  | 2020-03-03 10:42 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 220.5795 | 1.2234 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 1845.7349 | 1.483
UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032003.D | Cal | 1 | 2020-03-03 11:03 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 414.6233 | 2.5998 | 148.5601 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 3288.1632 | 3.6199 | 206.8491
UCGCAACC2_1/29/20_GC 2/3 | 7010_2_03032004.D | Cal | 2 | 2020-03-03 11:24 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 582.3344 | 3.6656 | 122.187 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 2605.1216 | 2.7479 | 91.5959
UCGCAACC3_1/29/20_GC 2/3 | 7010_2_03032005.D | Cal | 3 | 2020-03-03 11:44 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 932.6245 | 5.5942 | 111.8838 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 4565.5174 | 5.4605 | 109.2099
UCGCAACC4_1/29/20_GC 2/3 | 7010_2_03032006.D | Cal | 4 | 2020-03-03 12:05 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 1142.5911 | 7.9643 | 106.1906 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 5766.0826 | 6.6174 | 88.232
UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032007.D | Cal | 5 | 2020-03-03 12:26 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 2229.9551 | 14.0215 | 112.1724 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 11744.3518 | 12.4817 | 99.8532
UCGCAACC6_1/29/20_GC 2/3 | 7010_2_03032008.D | Cal | 6 | 2020-03-03 12:46 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 3009.4331 | 16.5986 | 82.9931 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 14103.3439 | 17.3912 | 86.9558
UCGCAACC7_1/29/20_GC 2/3 | 7010_2_03032009.D | Cal | 7 | 2020-03-03 13:07 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 4711.3505 | 38.4269 | 122.9661 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 24140.9873 | 30.5331 | 97.7058
UCGCAACC8_1/29/20_GC 2/3 | 7010_2_03032010.D | Cal | 8 | 2020-03-03 13:27 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 7377.3875 | 55.3364 | 110.6728 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 39401.3683 | 49.1537 | 98.3073
UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032011.D | Cal | 9 | 2020-03-03 13:48 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 12246.7588 | 70.8202 | 80.9374 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 61412.4905 | 80.9113 | 92.47
UCGCAACC10_1/29/20_GC 2/3 | 7010_2_03032012.D | Cal | 10 | 2020-03-03 14:08 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 17644.0643 | 141.021 | 112.8168 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 88925.2618 | 124.3844 | 99.5075
UCGCAACC11_1/29/20_GC 2/3 | 7010_2_03032013.D | Cal | 11 | 2020-03-03 14:29 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 28820.8624 | 176.8582 | 88.4291 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 148972.0372 | 197.8912 | 98.9456
UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032014.D | Cal | 12 | 2020-03-03 14:49 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 65149.3239 | 347.6115 | 92.6964 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 327950.4317 | 384.9569 | 102.6552
UCGCAACC13_1/29/20_GC 2/3 | 7010_2_03032015.D | Cal | 13 | 2020-03-03 15:10 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 104338.0012 | 724.1135 | 115.8582 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 508510.1834 | 623.3484 | 99.7357
UCGCAACC14_1/29/20_GC 2/3 | 7010_2_03032016.D | Cal | 14 | 2020-03-03 15:31 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 138411.0692 | 795.5382 | 90.9187 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 687340.632 | 892.3621 | 101.9842
UCGCAACC15_1/29/20_GC 2/3 | 7010_2_03032017.D | Cal | 15 | 2020-03-03 15:51 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 190214.6066 | 1266.146 | 101.2917 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 965328.509 | 1236.6406 | 98.9313
QCUCGCAACC4_1/30/20_GC 2/3 | 7010_2_03032018.D | QC | 4 | 2020-03-03 16:12 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 450.7911 | 3.4013 | 45.3504 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 4165.4096 | 5.7132 | 76.1758
QCUCGCAACC8_1/30/20_GC 2/3 | 7010_2_03032019.D | QC | 8 | 2020-03-03 16:32 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 5407.5313 | 38.5502 | 77.1004 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 32578.8152 | 46.1331 | 92.2662
QCUCGCAACC11_1/30/20_GC 2/3 | 7010_2_03032020.D | QC | 11 | 2020-03-03 16:53 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 20187.5767 | 184.8353 | 92.4177 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 124587.4178 | 188.7493 | 94.3746
SpikeBlank | 7010_2_03032021.D | MatrixBlank |  | 2020-03-03 17:13 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 220.6521 | 1.3144 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 1698.3707 | 1.3063
UCGCAAT1hS1A_GC 2/3 | 7010_2_03032022.D | Sample |  | 2020-03-03 17:34 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 31973.0127 | 204.482 | 105.3297 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 183985.2055 | 235.8448 | 101.1442
UCGCAAT1hS1B_GC 2/3 | 7010_2_03032023.D | Sample |  | 2020-03-03 17:54 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 33892.3029 | 194.5951 | 100.2369 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 194178.0654 | 234.8444 | 100.7152
UCGCAAT1hS2B_GC 2/3 | 7010_2_03032024.D | Sample |  | 2020-03-03 18:15 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 53.6062 | 0.19307 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 176186.5503 | 218.4364
UCGCAAT1hS2C_GC 2/3 | 7010_2_03032025.D | Sample |  | 2020-03-03 18:36 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 64.3884 | 0.22146 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 176869.109 | 211.0078
UCGCAAT1hS3A_GC 2/3 | 7010_2_03032026.D | Sample |  | 2020-03-03 18:56 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3657 | 145.2608 | 0.83991 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 169751.9152 | 220.6379
UCGCAAT1hS3B_GC 2/3 | 7010_2_03032027.D | Sample |  | 2020-03-03 19:17 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 74.1469 | 0.35631 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 185613.2951 | 227.2941
ACN | 7010_2_03032028.D | Blank |  | 2020-03-03 19:37 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.4161 | 0 |  |  | y = 153.340196 * x + 0.007796 | 0.99968
UCGCAAT5hS1B_GC 2/3 | 7010_2_03032029.D | Sample |  | 2020-03-03 19:58 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 20779.909 | 137.2516 | 104.9281 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 156886.9825 | 196.2927 | 100.5484
UCGCAAT5hS1C_GC 2/3 | 7010_2_03032030.D | Sample |  | 2020-03-03 20:19 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 22140.7158 | 122.8421 | 93.9121 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 164350.6487 | 189.4232 | 97.0296
UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032031.D | QC | 5 | 2020-03-03 20:39 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 2389.7828 | 13.5276 | 108.2206 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 11767.6091 | 11.6853 | 93.4822
UCGCAAT5hS3B_GC 2/3 | 7010_2_03032032.D | Sample |  | 2020-03-03 21:00 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.4245 | 22.99 | 0 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 152931.7244 | 180.4072
UCGCAAT5hS3C_GC 2/3 | 7010_2_03032033.D | Sample |  | 2020-03-03 21:20 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 89.5547 | 0.55573 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 153855.8897 | 186.6651
UCGCAAT5hS2A_GC 2/3 | 7010_2_03032034.D | Sample |  | 2020-03-03 21:41 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 144.0705 | 0.61588 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 157788.899 | 182.1726
SpikeBlank | 7010_2_03032035.D | MatrixBlank |  | 2020-03-03 22:01 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 260.7145 | 1.1991 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 2867.2199 | 2.4889
UCGCAAAFS1A_GC 2/3 | 7010_2_03032036.D | Sample |  | 2020-03-03 22:22 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2460.2076 | 16.98 | 102.8643 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 43377.5179 | 53.0758 | 101.8919
UCGCAAAFS1B_GC 2/3 | 7010_2_03032037.D | Sample |  | 2020-03-03 22:42 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 3018.0244 | 16.8063 | 101.8122 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41931.0929 | 50.6106 | 97.1594
UCGCAAAFS1B_GC 2/3_Dup | 7010_2_03032038.D | Sample |  | 2020-03-03 23:03 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2879.7851 | 17.8025 | 107.8471 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 43856.493 | 52.9525 | 101.655
UCGCAAAFS3A_GC 2/3 | 7010_2_03032039.D | Sample |  | 2020-03-03 23:24 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 155.0438 | 0.74838 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41569.3342 | 48.5766
SpikeBlank | 7010_2_03032040.D | MatrixBlank |  | 2020-03-03 23:44 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3657 | 145.7599 | 0.60846 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 2381.8933 | 1.9598
UCGCAAAFS3B_GC 2/3 | 7010_2_03032041.D | Sample |  | 2020-03-04 00:05 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 118.167 | 0.53612 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 48574.0434 | 59.4733
UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032042.D | QC | 9 | 2020-03-04 00:25 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 12456.4627 | 80.1614 | 91.613 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 63248.2678 | 79.9768 | 91.4021
UCGCAAT5hS2C_GC 2/3 | 7010_2_03032043.D | Sample |  | 2020-03-04 00:46 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 30.635 | 0.053418 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 124649.2223 | 163.1673
ACN | 7010_2_03032044.D | Blank |  | 2020-03-04 01:06 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3068 | 0 |  |  | y = 153.340196 * x + 0.007796 | 0.99968
UCGCAAT5hS3A_GC 2/3 | 7010_2_03032045.D | Sample |  | 2020-03-04 01:27 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3404 | 61.0611 | 0.40169 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 137120.2383 | 178.1027
UCGCAAAFS1C_GC 2/3 | 7010_2_03032046.D | Sample |  | 2020-03-04 01:48 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2284.9813 | 15.7352 | 95.3235 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41590.239 | 52.5845 | 100.9487
UCGCAAT1hS3C_GC 2/3 | 7010_2_03032047.D | Sample |  | 2020-03-04 02:08 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3656 | 13.6719 | 0 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 180499.8548 | 206.7185
SpikeBlank | 7010_2_03032048.D | MatrixBlank |  | 2020-03-04 02:29 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 94.3905 | 0.35351 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 3487.6365 | 3.0674
UCGCAAT5hS1A_GC 2/3 | 7010_2_03032049.D | Sample |  | 2020-03-04 02:49 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 25828.324 | 132.3226 | 101.1599 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 198589.9057 | 199.9504 | 102.422
UCGCAAAFS3C_GC 2/3 | 7010_2_03032050.D | Sample |  | 2020-03-04 03:10 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3657 | 193.3585 | 0.99623 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 42783.1405 | 51.1611
UCGCAAT1hS1C_GC 2/3 | 7010_2_03032051.D | Sample |  | 2020-03-04 03:30 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 32317.108 | 183.3286 | 94.4335 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 186786.148 | 228.8408 | 98.1405
UCGCAAT1hS2A_GC 2/3 | 7010_2_03032052.D | Sample |  | 2020-03-04 03:51 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 59.8406 | 0.2126 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 199135.7251 | 224.0855
UCGCAAAFS2B_GC 2/3 | 7010_2_03032053.D | Sample |  | 2020-03-04 04:11 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 310.805 | 1.78 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 44811.2231 | 51.5876
UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032054.D | QC | 12 | 2020-03-04 04:32 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 70221.7054 | 335.062 | 89.3499 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 359334.3791 | 385.608 | 102.8288
SpikeBlank | 7010_2_03032055.D | MatrixBlank |  | 2020-03-04 04:52 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 233.241 | 1.2021 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 2407.7011 | 2.0302
UCGCAAAFS2C_GC 2/3 | 7010_2_03032056.D | Sample |  | 2020-03-04 05:13 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 265.3131 | 1.5928 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 48789.4294 | 56.3312
UCGCAAT5hS2B_GC 2/3 | 7010_2_03032057.D | Sample |  | 2020-03-04 05:34 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 50.0304 | 0.19739 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 146504.7313 | 177.2152
UCGCAAAFS2A_GC 2/3 | 7010_2_03032058.D | Sample |  | 2020-03-04 05:54 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 128.2625 | 0.66788 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 48859.9988 | 59.4676
SpikeBlank | 7010_2_03032059.D | MatrixBlank |  | 2020-03-04 06:15 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 145.659 | 0.757 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.4896 | 2591.8192 | 2.2373
UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032060.D | QC | 1 | 2020-03-04 06:35 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 399.2569 | 2.0721 | 118.403 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 3049.9303 | 2.3058 | 131.7622
UCGCAACC2_1/29/20_GC 2/3 | 7010_2_03032061.D | QC | 2 | 2020-03-04 06:56 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 406.2685 | 2.4771 | 82.5694 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 3694.5937 | 3.8096 | 126.9874
UCGCAACC3_1/29/20_GC 2/3 | 7010_2_03032062.D | QC | 3 | 2020-03-04 07:16 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 956.9107 | 5.8852 | 117.7042 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 6341.3548 | 6.0272 | 120.5448
UCGCAACC4_1/29/20_GC 2/3 | 7010_2_03032063.D | QC | 4 | 2020-03-04 07:37 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 1641.6153 | 7.7882 | 103.8426 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 10139.9583 | 8.5509 | 114.0121
UCGCAACC5_1/29/20_GC 2/3 | 7010_2_03032064.D | QC | 5 | 2020-03-04 07:57 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2232.3705 | 12.8894 | 103.1151 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 10093.239 | 10.1792 | 81.4338
UCGCAACC6_1/29/20_GC 2/3 | 7010_2_03032065.D | QC | 6 | 2020-03-04 08:18 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 4474.6213 | 21.2059 | 106.0297 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 22659.2843 | 18.7936 | 93.9679
UCGCAACC7_1/29/20_GC 2/3 | 7010_2_03032066.D | QC | 7 | 2020-03-04 08:39 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 4759.5875 | 26.4755 | 84.7216 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 27337.2277 | 31.3433 | 100.2987
UCGCAACC8_1/29/20_GC 2/3 | 7010_2_03032067.D | QC | 8 | 2020-03-04 08:59 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 8624.2398 | 57.8816 | 115.7632 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41359.4884 | 44.7717 | 89.5433
UCGCAACC9_1/29/20_GC 2/3 | 7010_2_03032068.D | QC | 9 | 2020-03-04 09:20 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 13386.9727 | 86.9266 | 99.3447 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 66489.6628 | 83.205 | 95.0914
UCGCAACC10_1/29/20_GC 2/3 | 7010_2_03032069.D | QC | 10 | 2020-03-04 09:40 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 18619.6374 | 133.339 | 106.6712 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 90726.7922 | 121.5736 | 97.2589
UCGCAACC11_1/29/20_GC 2/3 | 7010_2_03032070.D | QC | 11 | 2020-03-04 10:01 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 30262.3173 | 189.3343 | 94.6671 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 155394.0301 | 193.2386 | 96.6193
UCGCAACC12_1/29/20_GC 2/3 | 7010_2_03032071.D | QC | 12 | 2020-03-04 10:21 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 64181.5261 | 419.2723 | 111.8059 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 320847.0647 | 383.8128 | 102.3501
UCGCAACC13_1/29/20_GC 2/3 | 7010_2_03032072.D | QC | 13 | 2020-03-04 10:42 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 107252.507 | 710.0162 | 113.6026 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 543793.7684 | 630.9246 | 100.9479
UCGCAACC14_1/29/20_GC 2/3 | 7010_2_03032073.D | QC | 14 | 2020-03-04 11:03 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 191486.6559 | 877.6558 | 100.3035 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 948621.8792 | 886.1897 | 101.2788
UCGCAACC15_1/29/20_GC 2/3 | 7010_2_03032074.D | QC | 15 | 2020-03-04 11:23 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 204239.078 | 1254.0275 | 100.3222 | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 1033676.7568 | 1270.3856 | 101.6308
SpikeBlank | 7010_2_03032075.D | MatrixBlank |  | 2020-03-04 11:44 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 523.7886 | 3.7028 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 3117.9408 | 2.8421
ACN | 7010_2_03032076.D | Blank |  | 2020-03-04 12:04 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3237 | 0 |  |  | y = 153.340196 * x + 0.007796 | 0.99968

## 899&900 Data
Sample | [5 empty] | 899 Method |  | 899 Results |  |  |  | 900 Method |  | 900 Results |  |  |  | 4NT Method |  | 4NT Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy
| ACN | 7010_2_020420000.D | Blank |  | 2020-02-04 13:56 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4747 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9628 | 0 |  |  | y = 159.065225 * x | 0.99958 | 12.1035 | 0
| SpikeBlank | 7010_2_020420001.D | MatrixBlank |  | 2020-02-04 14:17 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.1828 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 15839.3649 | 24.6721 |  | y = 159.065225 * x | 0.99958 | 11.4185 | 0
| SpikeBlank | 7010_2_020420002.D | MatrixBlank |  | 2020-02-04 14:37 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.1783 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 15390.4769 | 26.1031 |  | y = 159.065225 * x | 0.99958 | 11.4306 | 0
| UCGC3CC1_1/17/20_GC 2/3 | 7010_2_020420003.D | Cal | 1 | 2020-02-04 14:58 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4675 | 128.3693 | 2.2049 | 125.9952 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 1056.6832 | 2.1835 | 124.7722 | y = 159.065225 * x | 0.99958 | 11.5026 | 986.229 | 1.3558 | 77.4769
| UCGC3CC2_1/17/20_GC 2/3 | 7010_2_020420004.D | Cal | 2 | 2020-02-04 15:19 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4778 | 166.6586 | 3.1153 | 103.8446 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 1576.329 | 3.2675 | 108.9158 | y = 159.065225 * x | 0.99958 | 11.4967 | 1397.1031 | 2.0259 | 67.5313
| UCGC3CC3_1/17/20_GC 2/3 | 7010_2_020420005.D | Cal | 3 | 2020-02-04 15:39 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4823 | 304.0698 | 5.0253 | 100.5065 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 2645.784 | 5.2077 | 104.1537 | y = 159.065225 * x | 0.99958 | 11.4966 | 4420.282 | 6.218 | 124.3598
| UCGC3CC4_1/17/20_GC 2/3 | 7010_2_020420006.D | Cal | 4 | 2020-02-04 16:00 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 493.3071 | 7.6609 | 102.1453 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 3330.8483 | 6.55 | 87.3332 | y = 159.065225 * x | 0.99958 | 11.4967 | 4440.4834 | 6.465 | 86.1998
| UCGC3CC5_1/17/20_GC 2/3 | 7010_2_020420007.D | Cal | 5 | 2020-02-04 16:20 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4971 | 725.7007 | 10.6142 | 84.914 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 6670.0427 | 13.0299 | 104.2393 | y = 159.065225 * x | 0.99958 | 11.4966 | 7882.622 | 10.8909 | 87.1268
| UCGC3CC6_1/17/20_GC 2/3 | 7010_2_020420008.D | Cal | 6 | 2020-02-04 16:41 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 1035.8058 | 16.9881 | 84.9403 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 9562.4145 | 22.8421 | 114.2103 | y = 159.065225 * x | 0.99958 | 11.4967 | 10366.3512 | 15.2187 | 76.0935
| UCGC3CC7_1/17/20_GC 2/3 | 7010_2_020420009.D | Cal | 7 | 2020-02-04 17:01 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.5008 | 1870.4197 | 28.968 | 92.6976 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 16061.1743 | 28.8265 | 92.2447 | y = 159.065225 * x | 0.99958 | 11.4966 | 18413.7984 | 24.8181 | 79.418
| UCGC3CC8_1/17/20_GC 2/3 | 7010_2_020420010.D | Cal | 8 | 2020-02-04 17:22 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4889 | 2049.5703 | 33.1367 | 66.2734 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 23627.1803 | 45.3688 | 90.7377 | y = 159.065225 * x | 0.99958 | 11.4967 | 28076.6755 | 40.1088 | 80.2177
| UCGC3CC9_1/17/20_GC 2/3 | 7010_2_020420011.D | Cal | 9 | 2020-02-04 17:42 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4712 | 4785.6908 | 85.4867 | 97.6991 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 52508.8137 | 94.3605 | 107.8406 | y = 159.065225 * x | 0.99958 | 11.5026 | 62029.6447 | 89.2109 | 101.9553
| UCGC3CC10_1/17/20_GC 2/3 | 7010_2_020420012.D | Cal | 10 | 2020-02-04 18:03 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4815 | 9809.9719 | 115.2773 | 92.2218 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9482 | 90522.1041 | 148.7304 | 118.9843 | y = 159.065225 * x | 0.99958 | 11.5027 | 110060.6054 | 123.8572 | 99.0858
| UCGC3CC11_1/17/20_GC 2/3 | 7010_2_020420013.D | Cal | 11 | 2020-02-04 18:23 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4971 | 11544.7516 | 150.1797 | 75.0899 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 124957.6681 | 190.6807 | 95.3404 | y = 159.065225 * x | 0.99958 | 11.4966 | 163258.4832 | 186.7823 | 93.3912
| UCGC3CC12_1/17/20_GC 2/3 | 7010_2_020420014.D | Cal | 12 | 2020-02-04 18:44 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4815 | 22636.1264 | 369.4193 | 98.5118 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 225455.8294 | 347.2004 | 92.5868 | y = 159.065225 * x | 0.99958 | 11.4967 | 278430.7681 | 357.7792 | 95.4078
| UCGC3CC13_1/17/20_GC 2/3 | 7010_2_020420015.D | Cal | 13 | 2020-02-04 19:05 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.486 | 38711.6519 | 588.6371 | 94.1819 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 382089.4493 | 761.9213 | 121.9074 | y = 159.065225 * x | 0.99958 | 11.4966 | 501330.5361 | 638.9275 | 102.2284
| UCGC3CC14_1/17/20_GC 2/3 | 7010_2_020420016.D | Cal | 14 | 2020-02-04 19:25 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 51273.1269 | 981.2447 | 112.1422 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 504085.1085 | 885.252 | 101.1717 | y = 159.065225 * x | 0.99958 | 11.4967 | 656392.1371 | 879.6014 | 100.5259
| UCGC3CC15_1/17/20_GC 2/3 | 7010_2_020420017.D | Cal | 15 | 2020-02-04 19:46 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4453 | 69638.1126 | 1204.8083 | 96.3847 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9673 | 1170364.2748 | 1545.7639 | 123.6611 | y = 159.065225 * x | 0.99958 | 11.5026 | 1589342.9006 | 1247.7441 | 99.8195
| SpikeBlank | 7010_2_020420018.D | MatrixBlank |  | 2020-02-04 20:06 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.5739 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 16258.854 | 24.9899 |  | y = 159.065225 * x | 0.99958 | 11.4967 | 0
| QCUCGC3CC4_1/17/20_GC 2/3 | 7010_2_020420019.D | QC | 4 | 2020-02-04 20:27 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4823 | 345.4695 | 4.6258 | 61.6777 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 2865.8731 | 5.39 | 71.8662 | y = 159.065225 * x | 0.99958 | 11.5026 | 2801.8231 | 3.2655 | 43.5401
| QCUCGC3CC8_1/17/20_GC 2/3 | 7010_2_020420020.D | QC | 8 | 2020-02-04 20:47 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 2787.359 | 42.9047 | 85.8094 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 29806.9076 | 51.921 | 103.842 | y = 159.065225 * x | 0.99958 | 11.4967 | 35250.5009 | 40.5895 | 81.179
| QCUCGC3CC11_1/17/20_GC 2/3 | 7010_2_020420021.D | QC | 11 | 2020-02-04 21:08 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4786 | 11296.2856 | 193.0941 | 96.5471 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 114259.0595 | 205.546 | 102.773 | y = 159.065225 * x | 0.99958 | 11.4966 | 142609.4827 | 175.9776 | 87.9888
| SpikeBlank | 7010_2_020420022.D | MatrixBlank |  | 2020-02-04 21:28 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.2707 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 15416.7741 | 21.1726 |  | y = 159.065225 * x | 0.99958 | 11.9952 | 0
| ACN | 7010_2_020420027.D | Blank |  | 2020-02-04 23:11 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.1939 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9817 | 0 |  |  | y = 159.065225 * x | 0.99958 | 11.7249 | 0
| ACN | 7010_2_020420054.D | Blank |  | 2020-02-05 08:26 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 4.7383 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.8811 | 0 |  |  | y = 159.065225 * x | 0.99958 | 12.1754 | 0
| SpikeBlank | 7010_2_020420032.D | MatrixBlank |  | 2020-02-05 00:54 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.2301 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 14097.9408 | 25.2738 |  | y = 159.065225 * x | 0.99958 | 11.4246 | 0
| SpikeBlank | 7010_2_020420037.D | MatrixBlank |  | 2020-02-05 02:37 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.3861 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 14696.7614 | 25.4512 |  | y = 159.065225 * x | 0.99958 | 11.4966 | 0
| SpikeBlank | 7010_2_020420053.D | MatrixBlank |  | 2020-02-05 08:06 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.1384 | 0 |  |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 14552.7397 | 22.7924 |  | y = 159.065225 * x | 0.99958 | 11.4245 | 0
| UCGC3AFS3A_GC 2/3 | 7010_2_020420030.D | Sample |  | 2020-02-05 00:13 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 55689.0895 | 1116.3803 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 247327.9372 | 418.2042 | 104.4651 | y = 159.065225 * x | 0.99958 | 11.4967 | 41410.9431 | 52.6766 | 94.1074
| UCGC3AFS3B_GC 2/3 | 7010_2_020420031.D | Sample |  | 2020-02-05 00:34 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4749 | 46234.7342 | 873.0076 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 203632.3294 | 413.9578 | 103.4044 | y = 159.065225 * x | 0.99958 | 11.4966 | 42391.9698 | 57.4567 | 102.6471
| UCGC3AFS3C_GC 2/3 | 7010_2_020420036.D | Sample |  | 2020-02-05 02:16 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.5037 | 58070.3301 | 904.7044 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 266727.0549 | 368.8248 | 92.1304 | y = 159.065225 * x | 0.99958 | 11.4967 | 53254.8846 | 57.7916 | 103.2455
| UCGC3CC1_1/17/20_GC 2/3 | 7010_2_020420038.D | QC | 1 | 2020-02-05 02:57 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4741 | 111.0311 | 1.6613 | 94.9304 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 960.5029 | 2.2058 | 126.048 | y = 159.065225 * x | 0.99958 | 11.4967 | 1059.5213 | 1.4083 | 80.473
| UCGC3CC10_1/17/20_GC 2/3 | 7010_2_020420047.D | QC | 10 | 2020-02-05 06:02 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4638 | 8190.5143 | 111.4363 | 89.149 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9578 | 82835.1614 | 129.6482 | 103.7186 | y = 159.065225 * x | 0.99958 | 11.4966 | 99993.9875 | 123.8763 | 99.101
| UCGC3CC11_1/17/20_GC 2/3 | 7010_2_020420048.D | QC | 11 | 2020-02-05 06:23 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 11692.224 | 157.2738 | 78.6369 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 121907.8409 | 213.7278 | 106.8639 | y = 159.065225 * x | 0.99958 | 11.4967 | 152072.8573 | 190.2907 | 95.1453
| UCGC3CC12_1/17/20_GC 2/3 | 7010_2_020420049.D | QC | 12 | 2020-02-05 06:44 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4786 | 22842.0493 | 348.4238 | 92.913 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 223885.4728 | 404.328 | 107.8208 | y = 159.065225 * x | 0.99958 | 11.4966 | 278692.2252 | 361.6934 | 96.4516
| UCGC3CC13_1/17/20_GC 2/3 | 7010_2_020420033.D | QC | 13 | 2020-02-05 01:15 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4712 | 36269.479 | 666.2971 | 106.6075 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 348632.538 | 647.6961 | 103.6314 | y = 159.065225 * x | 0.99958 | 11.4966 | 437007.3466 | 639.0712 | 102.2514
| UCGC3CC13_1/17/20_GC 2/3 | 7010_2_020420050.D | QC | 13 | 2020-02-05 07:04 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4741 | 37324.8987 | 661.9771 | 105.9163 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 365338.113 | 674.5854 | 107.9337 | y = 159.065225 * x | 0.99958 | 11.4967 | 468400.9535 | 637.4921 | 101.9987
| UCGC3CC14_1/17/20_GC 2/3 | 7010_2_020420051.D | QC | 14 | 2020-02-05 07:25 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4823 | 52355.0638 | 937.3292 | 107.1233 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 496286.9047 | 846.8221 | 96.7797 | y = 159.065225 * x | 0.99958 | 11.4966 | 638601.9841 | 861.8958 | 98.5024
| UCGC3CC2_1/17/20_GC 2/3 | 7010_2_020420039.D | QC | 2 | 2020-02-05 03:18 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.5082 | 211.47 | 3.8124 | 127.0787 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 1564.3277 | 3.9651 | 132.1693 | y = 159.065225 * x | 0.99958 | 11.4966 | 1533.2091 | 2.3364 | 77.8802
| UCGC3CC3_1/17/20_GC 2/3 | 7010_2_020420040.D | QC | 3 | 2020-02-05 03:38 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4926 | 224.6806 | 3.5959 | 71.9188 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 2647.7265 | 6.354 | 127.0808 | y = 159.065225 * x | 0.99958 | 11.4967 | 2795.1069 | 3.9664 | 79.3274
| UCGC3CC4_1/17/20_GC 2/3 | 7010_2_020420041.D | QC | 4 | 2020-02-05 03:59 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4897 | 520.6992 | 9.1929 | 122.5714 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 3809.6465 | 7.8345 | 104.4596 | y = 159.065225 * x | 0.99958 | 11.4966 | 4185.3585 | 5.9635 | 79.5131
| UCGC3CC5_1/17/20_GC 2/3 | 7010_2_020420042.D | QC | 5 | 2020-02-05 04:20 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 751.4408 | 12.093 | 96.744 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 6483.6673 | 12.285 | 98.2798 | y = 159.065225 * x | 0.99958 | 11.4967 | 7800.6906 | 11.2567 | 90.0539
| UCGC3CC6_1/17/20_GC 2/3 | 7010_2_020420043.D | QC | 6 | 2020-02-05 04:40 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4823 | 788.6135 | 13.3377 | 66.6885 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 9488.8345 | 22.8355 | 114.1775 | y = 159.065225 * x | 0.99958 | 11.4966 | 10758.5357 | 15.6331 | 78.1655
| UCGC3CC7_1/17/20_GC 2/3 | 7010_2_020420044.D | QC | 7 | 2020-02-05 05:01 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4963 | 1561.7789 | 26.4694 | 84.7022 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 15505.8926 | 30.5377 | 97.7208 | y = 159.065225 * x | 0.99958 | 11.4967 | 17870.7297 | 24.7379 | 79.1612
| UCGC3CC8_1/17/20_GC 2/3 | 7010_2_020420045.D | QC | 8 | 2020-02-05 05:21 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4749 | 2198.0739 | 41.1849 | 82.3699 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 22166.0443 | 45.461 | 90.9221 | y = 159.065225 * x | 0.99958 | 11.4966 | 27453.5753 | 42.7605 | 85.521
| UCGC3CC9_1/17/20_GC 2/3 | 7010_2_020420028.D | QC | 9 | 2020-02-04 23:32 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4889 | 4604.0533 | 69.203 | 79.0891 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 57990.069 | 99.7226 | 113.9687 | y = 159.065225 * x | 0.99958 | 11.4967 | 66092.6083 | 87.8687 | 100.4214
| UCGC3CC9_1/17/20_GC 2/3 | 7010_2_020420046.D | QC | 9 | 2020-02-05 05:42 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 4938.3038 | 69.0831 | 78.9521 | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 56587.4011 | 98.0984 | 112.1125 | y = 159.065225 * x | 0.99958 | 11.4967 | 66333.2243 | 87.3069 | 99.7793
| UCGC3T1hS3A_GC 2/3 | 7010_2_020420055.D | Sample |  | 2020-02-05 08:47 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4934 | 7153.0608 | 238.8677 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 88984.6614 | 242.3608 | 92.9992 | y = 159.065225 * x | 0.99958 | 11.4966 | 113121.2744 | 216.7863 | 93.4784
| UCGC3T1hS3B_GC 2/3 | 7010_2_020420024.D | Sample |  | 2020-02-04 22:10 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4852 | 14368.6223 | 223.7621 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 150735.3745 | 290.9379 | 111.6392 | y = 159.065225 * x | 0.99958 | 11.4967 | 196836.21 | 230.5508 | 99.4137
| UCGC3T1hS3B_GC 2/3_dup | 7010_2_020420025.D | Sample |  | 2020-02-04 22:30 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4823 | 13400.0623 | 203.3976 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 148666.1484 | 252.5062 | 96.8922 | y = 159.065225 * x | 0.99958 | 11.4966 | 198369.3771 | 245.9777 | 106.0657
| UCGC3T1hS3C_GC 2/3 | 7010_2_020420035.D | Sample |  | 2020-02-05 01:56 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4749 | 13954.0623 | 217.7315 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 139925.4033 | 256.6165 | 98.4694 | y = 159.065225 * x | 0.99958 | 11.4966 | 185840.5881 | 234.3276 | 101.0422
| UCGC3T5hS3A_GC 2/3 | 7010_2_020420034.D | Sample |  | 2020-02-05 01:35 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4778 | 9462.9131 | 143.8378 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 121245.9556 | 187.8904 | 83.1529 | y = 159.065225 * x | 0.99958 | 11.4967 | 152842.322 | 179.5478 | 97.3609
| UCGC3T5hS3B_GC 2/3 | 7010_2_020420026.D | Sample |  | 2020-02-04 22:51 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4889 | 11952.3312 | 174.0812 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 136331.5861 | 262.4598 | 116.1544 | y = 159.065225 * x | 0.99958 | 11.4967 | 154230.2733 | 190.3628 | 103.2254
| UCGC3T5hS3C_GC 2/3 | 7010_2_020420029.D | Sample |  | 2020-02-04 23:52 | y = -238.476049 * x ^ 2 + 126.595216 * x | 0.99312 | 5.4934 | 15194.4309 | 201.7486 |  | y = 4652.837198 * x - 0.130559 | 0.99779 | 8.9626 | 162459.5585 | 227.5232 | 100.6928 | y = 159.065225 * x | 0.99958 | 11.4966 | 179825.1188 | 183.3333 | 99.4136

## 899_Data0225
Sample | [6 empty] | 899 Method |  | MFBET (ISTD) Results |  | 899 Results |  |  |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Area | RT | Final Conc. | Accuracy | Area | CF Formula | CF R2 | 4 | Final Conc. | Accuracy | Area | RT | Area
| ACN | PFAS Group 2 UC | 7010_1_020242000.D | Blank |  | 2020-02-24 15:25 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.5869 | 0 | 5.3159 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.6082 | 0 | 0 | 0 | 10.8524 | 0
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242001.D | MatrixBlank |  | 2020-02-24 15:46 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 48592.6222 | 5.1097 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 0 |  | 16744.0624 | 11.2181 | 610704.7863
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242002.D | MatrixBlank |  | 2020-02-24 16:07 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 45056.2853 | 5.308 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2267 | 0 |  | 15229.2335 | 11.2103 | 547939.7252
| UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242003.D | Cal | 1 | 2020-02-24 16:29 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6178 | 45159.1111 | 5.2769 | 1.9786 | 113.062 | 47.6438 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 1.5609 | 89.1947 | 21490.6389 | 11.2181 | 416578.0193
| UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242004.D | Cal | 2 | 2020-02-24 16:50 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6334 | 41986.7862 | 5.3041 | 2.664 | 88.7993 | 190.8916 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 3.4644 | 115.4808 | 33257.574 | 11.2181 | 432368.7952
| UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242005.D | Cal | 3 | 2020-02-24 17:11 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 47553.3329 | 5.3119 | 4.8336 | 96.6728 | 741.7786 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 4.6743 | 93.487 | 47273.3755 | 11.2181 | 508203.9671
| UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242006.D | Cal | 4 | 2020-02-24 17:32 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 53346.989 | 5.2964 | 7.3414 | 97.8852 | 1513.6419 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 7.1235 | 94.98 | 63504.7036 | 11.2103 | 505560.7045
| UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242007.D | Cal | 5 | 2020-02-24 17:54 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 50078.2096 | 5.3041 | 14.4823 | 115.8583 | 3242.5531 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 13.2238 | 105.7903 | 111659.5147 | 11.2103 | 539961.3813
| UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242008.D | Cal | 6 | 2020-02-24 18:15 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6567 | 43827.3109 | 5.3158 | 22.2636 | 111.3181 | 4575.0631 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 19.7778 | 98.8888 | 149818.1114 | 11.2103 | 509571.151
| UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242009.D | Cal | 7 | 2020-02-24 18:36 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 43628.2816 | 5.3002 | 31.8225 | 101.8321 | 6678.7121 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 32.5888 | 104.2841 | 226358.1451 | 11.2103 | 487326.1154
| UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242010.D | Cal | 8 | 2020-02-24 18:57 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6334 | 48761.7652 | 5.2886 | 47.8198 | 95.6396 | 11438.2136 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 47.9279 | 95.8559 | 337295.8545 | 11.2181 | 504469.2157
| UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242011.D | Cal | 9 | 2020-02-24 19:19 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6489 | 52843.4966 | 5.3119 | 89.0793 | 101.8049 | 23502.2592 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 91.9556 | 105.0922 | 693463.098 | 11.2103 | 552775.3856
| UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242012.D | Cal | 10 | 2020-02-24 19:40 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 50318.8843 | 5.2925 | 100.3852 | 80.3082 | 25277.4592 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 120.0336 | 96.0268 | 884914.7148 | 11.2104 | 543507.0176
| UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242013.D | Cal | 11 | 2020-02-24 20:01 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6528 | 49969.8291 | 5.308 | 192.29 | 96.145 | 48496.4758 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 199.2471 | 99.6236 | 1403950.5367 | 11.2103 | 523415.4418
| UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242014.D | Cal | 12 | 2020-02-24 20:23 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 50683.9632 | 5.2769 | 369.6345 | 98.5692 | 94977.7173 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 375.2444 | 100.0652 | 2647826.9055 | 11.2103 | 526997.7151
| UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242015.D | Cal | 13 | 2020-02-24 20:44 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 51176.2387 | 5.2925 | 618.1802 | 98.9088 | 160694.9345 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 639.6841 | 102.3495 | 4668943.8911 | 11.2103 | 546498.4929
| UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242016.D | Cal | 14 | 2020-02-24 21:05 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6528 | 51434.5275 | 5.2925 | 873.2095 | 99.7954 | 228326.4745 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 872.7176 | 99.7392 | 6371230.8243 | 11.2103 | 547146.9709
| UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242017.D | Cal | 15 | 2020-02-24 21:26 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 47441.556 | 5.2886 | 1292.5155 | 103.4012 | 311934.9841 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 1239.2761 | 99.1421 | 8970061.8581 | 11.2103 | 542902.5257
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242018.D | MatrixBlank |  | 2020-02-24 21:48 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.614 | 52230.0349 | 5.0708 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 1.6211 |  | 30359.5493 | 11.2103 | 579489.9307
| QCUCGCAACC4_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242019.D | QC | 4 | 2020-02-24 22:09 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6528 | 52147.8654 | 5.3002 | 9.1621 | 122.1608 | 1963.2701 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 8.1441 | 108.5886 | 74858.7215 | 11.2103 | 537799.0956
| QCUCGCAACC8_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242020.D | QC | 8 | 2020-02-24 22:30 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 52724.8583 | 5.2886 | 44.4126 | 88.8251 | 11452.7187 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 50.0784 | 100.1568 | 389036.7282 | 11.2104 | 557972.9033
| QCUCGCAACC11_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242021.D | QC | 11 | 2020-02-24 22:51 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6178 | 53816.5575 | 5.2964 | 191.7286 | 95.8643 | 52075.8643 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 215.7814 | 107.8907 | 1701103.4203 | 11.2103 | 586119.3129
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242022.D | MatrixBlank |  | 2020-02-24 23:12 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 56865.3096 | 5.2964 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 0.82749 |  | 25343.1657 | 11.2103 | 605881.7141
| UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242023.D | QC | 3 | 2020-02-24 23:34 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6373 | 47278.5083 | 5.2964 | 5.6209 | 112.4173 | 927.087 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 5.6983 | 113.965 | 57233.744 | 11.2103 | 536670.3306
| UCGC0213T1hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242024.D | Sample |  | 2020-02-24 23:55 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 52552.468 | 5.2925 | 259.0094 | 98.0263 | 68864.1454 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 228.0052 | 100.782 | 1711899.4298 | 11.2103 | 558534.5479
| UCGC0213T1hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242025.D | Sample |  | 2020-02-25 00:16 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6528 | 52540.6327 | 5.2847 | 271.5387 | 102.7682 | 72202.0537 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 231.0649 | 102.1344 | 1914008.384 | 11.2103 | 616288.7193
| UCGC0213T5hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242026.D | Sample |  | 2020-02-25 00:37 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6529 | 57340.2311 | 5.2964 | 209.1524 | 105.2556 | 60574.9721 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 173.172 | 98.7589 | 1498168.4841 | 11.2103 | 641533.1995
| ACN | PFAS Group 2 UC | 7010_1_020242027.D | Blank |  | 2020-02-25 00:59 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.5167 | 0 | 4.6625 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.7092 | 0 | 0 | 0 | 11.078 | 0
| UCGC0213T5hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242028.D | Sample |  | 2020-02-25 01:20 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 56418.084 | 5.3003 | 229.7967 | 115.6448 | 65533.9485 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 187.3907 | 106.8677 | 1501117.6088 | 11.2103 | 594617.8429
| UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242029.D | QC | 9 | 2020-02-25 01:41 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 48687.3735 | 5.3002 | 105.021 | 120.024 | 25607.6277 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2111 | 93.1926 | 106.5059 | 709588.8027 | 11.2103 | 558303.7836
| UCGC0213AFS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242030.D | Sample |  | 2020-02-25 02:03 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6334 | 52272.851 | 5.2925 | 1068.2191 | 112.4452 | 283975.4119 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 50.5671 | 99.0893 | 442901.1183 | 11.2103 | 629356.6386
| UCGC0213AFS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242031.D | Sample |  | 2020-02-25 02:24 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6528 | 58919.7768 | 5.308 | 801.1408 | 84.3315 | 239923.9382 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 42.9375 | 84.1386 | 367615.8695 | 11.2103 | 610449.5181
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242032.D | MatrixBlank |  | 2020-02-25 02:45 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.614 | 55026.6449 | 5.0825 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 0.7614 |  | 24847.9653 | 11.2103 | 606801.2007
| UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242033.D | QC | 13 | 2020-02-25 03:07 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 52147.894 | 5.2925 | 649.9639 | 103.9942 | 172189.1334 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 655.167 | 104.8267 | 4974673.0429 | 11.2103 | 568572.0123
| UCGC0213T5hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242034.D | Sample |  | 2020-02-25 03:28 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 56798.111 | 5.3003 | 157.1781 | 79.0996 | 44964.3705 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 165.4822 | 94.3734 | 1414688.4309 | 11.2103 | 633548.0231
| UCGC0213T1hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242035.D | Sample |  | 2020-02-25 03:49 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 51036.8422 | 5.2964 | 262.1253 | 99.2055 | 67688.1658 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 219.6382 | 97.0836 | 1755892.4077 | 11.2103 | 594484.2399
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242036.D | MatrixBlank |  | 2020-02-25 04:11 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.614 | 56151.808 | 5.0786 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 1.2599 |  | 26705.4648 | 11.2103 | 561248.775
| UCGC0213AFS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242037.D | Sample |  | 2020-02-25 04:32 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6373 | 58626.738 | 5.2925 | 980.6115 | 103.2233 | 292329.3904 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 59.591 | 116.7721 | 506962.0295 | 11.2103 | 615379.3913
| ACN | PFAS Group 2 UC | 7010_1_020242038.D | Blank |  | 2020-02-25 04:53 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.8162 | 0 | 5.6736 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.6081 | 0 | 0 | 0 | 11.2104 | 0
| UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242039.D | QC | 1 | 2020-02-25 05:15 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6101 | 48812.6514 | 5.2925 | 2.6007 | 148.6134 | 206.1992 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 2.4259 | 138.6214 | 30331.484 | 11.2103 | 480696.5929
| UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242040.D | QC | 2 | 2020-02-25 05:36 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 51153.7274 | 5.2886 | 3.9598 | 131.9939 | 570.2395 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 4.4588 | 148.6252 | 46657.1962 | 11.2103 | 517542.2318
| UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242041.D | QC | 3 | 2020-02-25 05:57 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 48275.7681 | 5.2886 | 5.0655 | 101.3108 | 810.0765 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 5.8673 | 117.3464 | 55884.5116 | 11.2103 | 513192.2188
| UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242042.D | QC | 4 | 2020-02-25 06:19 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 55478.5799 | 5.2925 | 7.086 | 94.4805 | 1501.9577 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 8.9305 | 119.0733 | 70316.3859 | 11.2103 | 469844.1334
| UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242043.D | QC | 5 | 2020-02-25 06:40 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6295 | 38630.6526 | 5.2808 | 15.1406 | 121.1247 | 2630.8702 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 13.8662 | 110.9298 | 108032.6159 | 11.2103 | 501681.938
| UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242044.D | QC | 6 | 2020-02-25 07:01 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 53645.5089 | 5.2847 | 20.7053 | 103.5266 | 5174.1268 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 21.6539 | 108.2694 | 170841.1938 | 11.2103 | 535594.8155
| UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242045.D | QC | 7 | 2020-02-25 07:23 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6062 | 37816.1574 | 5.273 | 27.8977 | 89.2727 | 5032.9135 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2189 | 31.8927 | 102.0566 | 191360.7939 | 11.2103 | 420363.2797
| UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242046.D | QC | 8 | 2020-02-25 07:44 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6296 | 46257.3871 | 5.277 | 47.1316 | 94.2633 | 10688.5964 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 48.4789 | 96.9578 | 314894.9672 | 11.2104 | 465857.2548
| UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242047.D | QC | 9 | 2020-02-25 08:05 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6334 | 44343.7186 | 5.2847 | 91.3316 | 104.3789 | 20230.7358 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2111 | 92.9825 | 106.2658 | 564515.2468 | 11.2103 | 445139.1458
| UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242048.D | QC | 10 | 2020-02-25 08:27 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 51808.0583 | 5.2808 | 101.9302 | 81.5441 | 26433.2733 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 120.7588 | 96.6071 | 861667.124 | 11.2103 | 526109.7482
| UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242049.D | QC | 11 | 2020-02-25 08:48 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6412 | 44392.9055 | 5.2847 | 207.2031 | 103.6016 | 46456.4468 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 203.2737 | 101.6369 | 1321569.7966 | 11.2103 | 483052.7241
| UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242050.D | QC | 12 | 2020-02-25 09:09 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.649 | 52888.6489 | 5.2769 | 373.9603 | 99.7227 | 100274.5775 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 382.5518 | 102.0138 | 2783968.5303 | 11.2103 | 543573.5052
| UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242051.D | QC | 13 | 2020-02-25 09:31 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6489 | 55103.7877 | 5.2886 | 596.6254 | 95.4601 | 166977.0724 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 636.7335 | 101.8774 | 4365677.4668 | 11.2103 | 513360.6541
| UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242052.D | QC | 14 | 2020-02-25 09:52 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6451 | 52945.9941 | 5.2847 | 845.0492 | 96.577 | 227441.002 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 871.9332 | 99.6495 | 6540834.1463 | 11.2103 | 562216.0944
| UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242053.D | QC | 15 | 2020-02-25 10:13 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6373 | 40411.6953 | 5.273 | 1474.2577 | 117.9406 | 303126.1082 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2112 | 1225.3577 | 98.0286 | 8489605.4621 | 11.2103 | 519648.7958
| SpikeBlank | PFAS Group 2 UC | 7010_1_020242054.D | MatrixBlank |  | 2020-02-25 10:35 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6023 | 46463.7434 | 5.3858 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.219 | 1.4432 |  | 26097.6892 | 11.2103 | 521723.1868
| ACN | PFAS Group 2 UC | 7010_1_020242055.D | Blank |  | 2020-02-25 10:56 | y = 61.128858 * x - 0.009024 | 0.99786 | 5.6684 | 0 | 5.3158 | 0 | 0 | 0 | y = 159.689355 * x + 0.030817 | 0.99966 | 11.2034 | 0 | 0 | 0 | 11.2025 | 0

## Control Chart
| QC Type | Description | Acceptance Criteria | Type
| Internal Standard | Every Sample | 75-125% of spike | Accuracy
[9 empty] | Mean = | 41180.0528 | [9 empty] | Mean = | 31933.81 | [9 empty] | Mean = | 15096.9983 | [9 empty] | Mean = | 457519.468 |  |  |  |  | -1
[6 empty] | MFBET (ISTD) Results |  |  | Stdev = | 4721.0259 | [6 empty] | MFHET (ISTD) Results |  |  | Stdev = | 3861.0681 | [6 empty] | MFOET (ISTD) Results |  |  | Stdev = | 1738.66 | [6 empty] | 4NT13C6 (ISTD) Results |  |  | Stdev = | 50059.0779
Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s
SpikeBlank | PFAS Group 2 UC | 7010_1_02042001.D | MatrixBlank |  | 2020-02-04 16:09 | 5.6528 | 41172.1507 | 7.902 | P | P | P | 41180.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 32044.9747 | 111.1647 | P | P | P | 31933.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1066 | 14570.3396 | 526.6587 | P | P | P | 15096.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 438721.4315 | 18798.0365 | P | P | P | 457519.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042002.D | MatrixBlank |  | 2020-02-04 16:30 | 5.6373 | 40847.531 | 332.5218 | P | P | P | 41181.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2857 | 31968.7845 | 34.9745 | P | P | P | 31934.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1032 | 14480.1653 | 616.833 | P | P | P | 15097.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 463849.0886 | 6329.6206 | P | P | P | 457520.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042003.D | Cal | 1 | 2020-02-04 16:51 | 5.6606 | 42024.0239 | 843.9711 | P | P | P | 41182.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 32827.3281 | 893.5181 | P | P | P | 31935.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1032 | 14691.0928 | 405.9055 | P | P | P | 15098.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 444373.7693 | 13145.6988 | P | P | P | 457521.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042004.D | Cal | 2 | 2020-02-04 17:13 | 5.6023 | 34962.3275 | 6217.7253 | F | P | P | 41183.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2827 | 24972.7331 | 6961.0769 | F | P | P | 31936.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1101 | 10968.6152 | 4128.3831 | F | F | P | 15099.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 362625.8197 | 94893.6483 | F | P | P | 457522.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042005.D | Cal | 3 | 2020-02-04 17:34 | 5.6606 | 41573.1688 | 393.116 | P | P | P | 41184.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 31037.7893 | 896.0206 | P | P | P | 31937.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1067 | 13852.0238 | 1244.9745 | P | P | P | 15100.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 434883.2637 | 22636.2043 | P | P | P | 457523.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042006.D | Cal | 4 | 2020-02-04 17:55 | 5.6879 | 45765.5533 | 4585.5005 | P | P | P | 41185.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 32956.5043 | 1022.6943 | P | P | P | 31938.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.0997 | 15036.032 | 60.9664 | P | P | P | 15101.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 459908.6328 | 2389.1648 | P | P | P | 457524.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042007.D | Cal | 5 | 2020-02-04 18:17 | 5.6878 | 44708.1517 | 3528.0989 | P | P | P | 41186.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3038 | 32560.0815 | 626.2715 | P | P | P | 31939.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.0997 | 15747.7567 | 650.7584 | P | P | P | 15102.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 479036.7769 | 21517.3089 | P | P | P | 457525.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042008.D | Cal | 6 | 2020-02-04 18:38 | 5.6529 | 39345.1842 | 1834.8686 | P | P | P | 41187.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 29718.0841 | 2215.7259 | P | P | P | 31940.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1101 | 13527.271 | 1569.7274 | P | P | P | 15103.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 431616.0593 | 25903.4087 | P | P | P | 457526.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042009.D | Cal | 7 | 2020-02-04 18:59 | 5.6801 | 41774.9961 | 594.9433 | P | P | P | 41188.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 32952.4017 | 1018.5917 | P | P | P | 31941.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1066 | 14617.3179 | 479.6805 | P | P | P | 15104.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 452700.6733 | 4818.7947 | P | P | P | 457527.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042010.D | Cal | 8 | 2020-02-04 19:20 | 5.684 | 40145.8739 | 1034.1789 | P | P | P | 41189.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 33856.1378 | 1922.3278 | P | P | P | 31942.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1032 | 15015.5888 | 81.4096 | P | P | P | 15105.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 445689.3513 | 11830.1167 | P | P | P | 457528.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042011.D | Cal | 9 | 2020-02-04 19:42 | 5.6684 | 39706.8932 | 1473.1596 | P | P | P | 41190.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 31411.0942 | 522.7158 | P | P | P | 31943.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1066 | 14587.0677 | 509.9306 | P | P | P | 15106.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 426840.4477 | 30679.0203 | P | P | P | 457529.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042012.D | Cal | 10 | 2020-02-04 20:03 | 5.5751 | 31360.5911 | 9819.4616 | F | F | P | 41191.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2705 | 24543.8103 | 7389.9996 | F | P | P | 31944.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1132 | 11619.7704 | 3477.228 | F | P | P | 15107.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 347579.0019 | 109940.4661 | F | F | P | 457530.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042013.D | Cal | 11 | 2020-02-04 20:24 | 5.6606 | 40243.2095 | 936.8433 | P | P | P | 41192.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 30910.0715 | 1023.7385 | P | P | P | 31945.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 14271.6128 | 825.3855 | P | P | P | 15108.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 436195.6576 | 21323.8104 | P | P | P | 457531.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042014.D | Cal | 12 | 2020-02-04 20:46 | 5.6917 | 39745.8323 | 1434.2205 | P | P | P | 41193.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 32964.8471 | 1031.0371 | P | P | P | 31946.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1067 | 15495.8861 | 398.8878 | P | P | P | 15109.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 466955.2069 | 9435.7389 | P | P | P | 457532.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042015.D | Cal | 13 | 2020-02-04 21:07 | 5.68 | 43579.4674 | 2399.4146 | P | P | P | 41194.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 32520.1082 | 586.2982 | P | P | P | 31947.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1032 | 15065.9234 | 31.075 | P | P | P | 15110.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 474955.3685 | 17435.9005 | P | P | P | 457533.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042016.D | Cal | 14 | 2020-02-04 21:28 | 5.6762 | 41425.4777 | 245.4249 | P | P | P | 41195.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 31871.6607 | 62.1493 | P | P | P | 31948.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1101 | 14128.672 | 968.3264 | P | P | P | 15111.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 453684.8102 | 3834.6578 | P | P | P | 457534.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042017.D | Cal | 15 | 2020-02-04 21:49 | 5.6606 | 41594.0843 | 414.0315 | P | P | P | 41196.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 32883.5617 | 949.7517 | P | P | P | 31949.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 14196.284 | 900.7144 | P | P | P | 15112.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 418676.8552 | 38842.6128 | P | P | P | 457535.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042018.D | MatrixBlank |  | 2020-02-04 22:11 | 5.6645 | 39337.1317 | 1842.921 | P | P | P | 41197.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 29058.2865 | 2875.5235 | P | P | P | 31950.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1067 | 15077.7524 | 19.2459 | P | P | P | 15113.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 465763.8198 | 8244.3518 | P | P | P | 457536.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | 5.6606 | 42279.1923 | 1099.1395 | P | P | P | 41198.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2948 | 34299.1518 | 2365.3418 | P | P | P | 31951.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1032 | 16387.1308 | 1290.1325 | P | P | P | 15114.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 515901.6354 | 58382.1674 | F | P | P | 457537.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | 5.6529 | 45252.3426 | 4072.2899 | P | P | P | 41199.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 32412.4363 | 478.6263 | P | P | P | 31952.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1132 | 15740.8529 | 643.8546 | P | P | P | 15115.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 487793.8886 | 30274.4206 | P | P | P | 457538.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | 5.6645 | 42273.4175 | 1093.3648 | P | P | P | 41200.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 35240.2399 | 3306.43 | P | P | P | 31953.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1066 | 16448.5129 | 1351.5145 | P | P | P | 15116.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 484818.9148 | 27299.4468 | P | P | P | 457539.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042022.D | MatrixBlank |  | 2020-02-04 23:36 | 5.6451 | 38859.6518 | 2320.4009 | P | P | P | 41201.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 30314.9518 | 1618.8582 | P | P | P | 31954.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 14500.284 | 596.7144 | P | P | P | 15117.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 430906.5473 | 26612.9207 | P | P | P | 457540.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042023.D | QC | 3 | 2020-02-04 23:57 | 5.6762 | 39783.8694 | 1396.1834 | P | P | P | 41202.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 30917.0632 | 1016.7468 | P | P | P | 31955.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1067 | 13960.7816 | 1136.2168 | P | P | P | 15118.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 450232.7156 | 7286.7524 | P | P | P | 457541.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | 5.6451 | 40772.8976 | 407.1551 | P | P | P | 41203.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2948 | 33685.8517 | 1752.0417 | P | P | P | 31956.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1132 | 16249.7142 | 1152.7158 | P | P | P | 15119.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 497897.4822 | 40378.0142 | P | P | P | 457542.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | 5.6684 | 38389.9729 | 2790.0799 | P | P | P | 41204.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 29174.7978 | 2759.0121 | P | P | P | 31957.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 15375.4789 | 278.4806 | P | P | P | 15120.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 429583.7135 | 27935.7545 | P | P | P | 457543.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042026.D | Sample |  | 2020-02-05 01:01 | 5.6762 | 45206.4963 | 4026.4435 | P | P | P | 41205.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 33169.5863 | 1235.7763 | P | P | P | 31958.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1101 | 15535.976 | 438.9777 | P | P | P | 15121.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 428774.9495 | 28744.5185 | P | P | P | 457544.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | 5.6762 | 46610.1305 | 5430.0777 | F | P | P | 41206.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 35166.8058 | 3232.9958 | P | P | P | 31959.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1067 | 16474.9308 | 1377.9325 | P | P | P | 15122.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 497767.6323 | 40248.1643 | P | P | P | 457545.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | 5.6684 | 45385.4767 | 4205.4239 | P | P | P | 41207.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3009 | 35474.482 | 3540.672 | P | P | P | 31960.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1101 | 17112.4072 | 2015.4089 | F | P | P | 15123.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 473692.7868 | 16173.3188 | P | P | P | 457546.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | 5.6256 | 44371.295 | 3191.2422 | P | P | P | 41208.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2826 | 31675.7212 | 258.0888 | P | P | P | 31961.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 15797.7788 | 700.7805 | P | P | P | 15124.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 438351.4479 | 19168.0201 | P | P | P | 457547.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | 5.649 | 44996.0952 | 3816.0424 | P | P | P | 41209.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2948 | 36119.1322 | 4185.3222 | F | P | P | 31962.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1097 | 16177.7404 | 1080.742 | P | P | P | 15125.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 494848.2367 | 37328.7687 | P | P | P | 457548.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042032.D | QC | 9 | 2020-02-05 03:09 | 5.6606 | 36116.4741 | 5063.5786 | F | P | P | 41210.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 29970.6563 | 1963.1537 | P | P | P | 31963.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 14802.8796 | 294.1187 | P | P | P | 15126.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 406575.6376 | 50943.8304 | F | P | P | 457549.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
QCUCGCAACC4_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042033.D | QC | 4 | 2020-02-05 03:30 | 5.6684 | 37408.4757 | 3771.577 | P | P | P | 41211.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 29211.8116 | 2721.9984 | P | P | P | 31964.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 13431.1384 | 1665.86 | P | P | P | 15127.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 393825.2296 | 63694.2384 | F | P | P | 457550.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
QCUCGCAACC8_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042034.D | QC | 8 | 2020-02-05 03:52 | 5.6645 | 38097.3649 | 3082.6878 | P | P | P | 41212.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 27967.0874 | 3966.7226 | F | P | P | 31965.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 13272.138 | 1824.8604 | F | P | P | 15128.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 380542.8038 | 76976.6642 | F | P | P | 457551.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
QCUCGCAACC11_1/30/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042035.D | QC | 11 | 2020-02-05 04:13 | 5.6606 | 37215.3988 | 3964.654 | P | P | P | 41213.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3008 | 28996.3692 | 2937.4408 | P | P | P | 31966.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 13148.2235 | 1948.7749 | F | P | P | 15129.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 390435.4446 | 67084.0234 | F | P | P | 457552.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | 5.6529 | 43830.4305 | 2650.3778 | P | P | P | 41214.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3009 | 32667.0296 | 733.2196 | P | P | P | 31967.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 16527.163 | 1430.1646 | P | P | P | 15130.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 443570.4903 | 13948.9777 | P | P | P | 457553.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | 5.6528 | 46946.472 | 5766.4192 | F | P | P | 41215.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 38576.7512 | 6642.9413 | F | P | P | 31968.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 18071.5743 | 2974.576 | F | P | P | 15131.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 570323.466 | 112803.998 | F | F | P | 457554.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | 5.6451 | 44276.266 | 3096.2132 | P | P | P | 41216.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 34951.2412 | 3017.4312 | P | P | P | 31969.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1201 | 15322.4289 | 225.4306 | P | P | P | 15132.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 456268.9968 | 1250.4712 | P | P | P | 457555.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | 5.5712 | 27629.305 | 13550.7477 | F | F | P | 41217.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2735 | 27013.7501 | 4920.0599 | F | P | P | 31970.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 12225.8416 | 2871.1567 | F | P | P | 15133.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 400149.7862 | 57369.6818 | F | P | P | 457556.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | 5.6529 | 43125.2652 | 1945.2125 | P | P | P | 41218.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3009 | 36772.6699 | 4838.8599 | F | P | P | 31971.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1201 | 17231.3883 | 2134.39 | F | P | P | 15134.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 507581.8712 | 50062.4032 | F | P | P | 457557.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | 5.6723 | 45855.005 | 4674.9522 | P | P | P | 41219.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 35943.9604 | 4010.1504 | F | P | P | 31972.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 16295.4315 | 1198.4331 | P | P | P | 15135.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 497271.4447 | 39751.9767 | P | P | P | 457558.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042042.D | MatrixBlank |  | 2020-02-05 06:42 | 5.6568 | 43295.0213 | 2114.9685 | P | P | P | 41220.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 33984.1072 | 2050.2972 | P | P | P | 31973.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1166 | 15368.0945 | 271.0961 | P | P | P | 15136.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 494808.5795 | 37289.1115 | P | P | P | 457559.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042043.D | QC | 13 | 2020-02-05 07:04 | 5.6723 | 46205.1348 | 5025.0821 | F | P | P | 41221.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 32996.0187 | 1062.2087 | P | P | P | 31974.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 14975.4163 | 121.5821 | P | P | P | 15137.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 457356.7859 | 162.6821 | P | P | P | 457560.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | 5.684 | 48159.5878 | 6979.535 | F | P | P | 41222.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 36298.6945 | 4364.8845 | F | P | P | 31975.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1132 | 17572.9112 | 2475.9128 | F | P | P | 15138.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 489678.7122 | 32159.2442 | P | P | P | 457561.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | 5.6762 | 45158.9158 | 3978.8631 | P | P | P | 41223.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 33682.1857 | 1748.3757 | P | P | P | 31976.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 15982.4355 | 885.4371 | P | P | P | 15139.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 466064.303 | 8544.835 | P | P | P | 457562.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS3A_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042046.D | Sample |  | 2020-02-05 08:08 | 5.6645 | 42126.9768 | 946.9241 | P | P | P | 41224.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 34900.0627 | 2966.2527 | P | P | P | 31977.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 16001.3802 | 904.3819 | P | P | P | 15140.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 460468.8876 | 2949.4196 | P | P | P | 457563.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | 5.5867 | 34358.3476 | 6821.7051 | F | P | P | 41225.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2887 | 27272.3765 | 4661.4335 | F | P | P | 31978.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 13553.8369 | 1543.1614 | P | P | P | 15141.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 373091.7078 | 84427.7602 | F | P | P | 457564.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | 5.6295 | 46634.702 | 5454.6492 | F | P | P | 41226.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2918 | 35634.3775 | 3700.5675 | P | P | P | 31979.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 18208.5429 | 3111.5445 | F | P | P | 15142.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 524339.2574 | 66819.7894 | F | P | P | 457565.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | 5.6256 | 49525.8927 | 8345.8399 | F | P | P | 41227.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2735 | 38036.2844 | 6102.4744 | F | P | P | 31980.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 18880.9947 | 3783.9963 | F | F | P | 15143.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 564202.1179 | 106682.6499 | F | F | P | 457566.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042050.D | MatrixBlank |  | 2020-02-05 09:33 | 5.6645 | 41200.1748 | 20.122 | P | P | P | 41228.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 35888.428 | 3954.618 | F | P | P | 31981.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1132 | 16517.2482 | 1420.2498 | P | P | P | 15144.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 495498.5055 | 37979.0375 | P | P | P | 457567.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042051.D | QC | 6 | 2020-02-05 09:55 | 5.6606 | 43924.8288 | 2744.776 | P | P | P | 41229.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 35124.4078 | 3190.5978 | P | P | P | 31982.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1136 | 16431.8493 | 1334.8509 | P | P | P | 15145.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 486877.1563 | 29357.6883 | P | P | P | 457568.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | 5.6723 | 48219.7539 | 7039.7011 | F | P | P | 41230.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 38121.0607 | 6187.2507 | F | P | P | 31983.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 17456.357 | 2359.3586 | F | P | P | 15146.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 563445.3712 | 105925.9032 | F | F | P | 457569.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | 5.6684 | 48899.4333 | 7719.3806 | F | P | P | 41231.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 38512.1517 | 6578.3417 | F | P | P | 31984.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 16613.0802 | 1516.0819 | P | P | P | 15147.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 539605.0373 | 82085.5693 | F | P | P | 457570.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS1C_GC 2/3_dup | PFAS Group 2 UC | 7010_1_02042054.D | Sample |  | 2020-02-05 10:59 | 5.6723 | 48880.2301 | 7700.1773 | F | P | P | 41232.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 36219.3641 | 4285.5542 | F | P | P | 31985.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 16150.3833 | 1053.3849 | P | P | P | 15148.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 522069.9328 | 64550.4648 | F | P | P | 457571.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | 5.6567 | 45429.6893 | 4249.6365 | P | P | P | 41233.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 33035.4225 | 1101.6125 | P | P | P | 31986.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 15082.165 | 14.8334 | P | P | P | 15149.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 433977.2416 | 23542.2264 | P | P | P | 457572.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | 5.6723 | 47307.1562 | 6127.1034 | F | P | P | 41234.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3039 | 37723.2278 | 5789.4178 | F | P | P | 31987.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1166 | 17563.247 | 2466.2487 | F | P | P | 15150.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 511111.1467 | 53591.6787 | F | P | P | 457573.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | 5.5751 | 35262.4856 | 5917.5672 | F | P | P | 41235.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2705 | 27240.8843 | 4692.9257 | F | P | P | 31988.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 13754.1616 | 1342.8367 | P | P | P | 15151.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 408763.0851 | 48756.3829 | P | P | P | 457574.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | 5.6529 | 44216.1553 | 3036.1026 | P | P | P | 41236.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3009 | 36269.1837 | 4335.3737 | F | P | P | 31989.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 17483.3019 | 2386.3036 | F | P | P | 15152.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 502724.7548 | 45205.2868 | P | P | P | 457575.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | 5.6645 | 43521.4244 | 2341.3717 | P | P | P | 41237.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 35816.3161 | 3882.5061 | F | P | P | 31990.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.117 | 17384.4001 | 2287.4017 | F | P | P | 15153.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 504516.0123 | 46996.5443 | P | P | P | 457576.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042064.D | QC | 1 | 2020-02-05 14:32 | 5.5868 | 27298.1014 | 13881.9513 | F | F | P | 41238.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.2978 | 20213.1454 | 11720.6646 | F | F | F | 31991.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1275 | 11613.6729 | 3483.3254 | F | F | P | 15154.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 337701.7486 | 119817.7194 | F | F | P | 457577.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042065.D | QC | 2 | 2020-02-05 14:53 | 5.6139 | 30145.0196 | 11035.0331 | F | F | P | 41239.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3038 | 21437.2156 | 10496.5944 | F | F | P | 31992.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1274 | 10863.8929 | 4233.1054 | F | F | P | 15155.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 333403.9132 | 124115.5548 | F | F | P | 457578.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042066.D | QC | 3 | 2020-02-05 15:15 | 5.6412 | 34866.664 | 6313.3888 | F | P | P | 41240.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 24943.155 | 6990.655 | F | P | P | 31993.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 12482.98 | 2614.0183 | F | P | P | 15156.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 413428.1999 | 44091.2681 | P | P | P | 457579.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042067.D | QC | 4 | 2020-02-05 15:36 | 5.6528 | 39646.9403 | 1533.1125 | P | P | P | 41241.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3099 | 30275.051 | 1658.759 | P | P | P | 31994.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 15336.2093 | 239.2109 | P | P | P | 15157.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 460867.2116 | 3347.7436 | P | P | P | 457580.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042068.D | QC | 5 | 2020-02-05 15:58 | 5.6645 | 39778.031 | 1402.0218 | P | P | P | 41242.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 32054.8952 | 121.0852 | P | P | P | 31995.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1171 | 15762.5805 | 665.5822 | P | P | P | 15158.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 495756.9008 | 38237.4328 | P | P | P | 457581.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042069.D | QC | 6 | 2020-02-05 16:19 | 5.6606 | 40365.9838 | 814.069 | P | P | P | 41243.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 30067.2479 | 1866.5621 | P | P | P | 31996.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 13388.8649 | 1708.1334 | P | P | P | 15159.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 472128.959 | 14609.491 | P | P | P | 457582.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042070.D | QC | 7 | 2020-02-05 16:40 | 5.6529 | 40269.6793 | 910.3734 | P | P | P | 41244.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 28746.775 | 3187.035 | P | P | P | 31997.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 15624.2761 | 527.2778 | P | P | P | 15160.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 483435.5795 | 25916.1115 | P | P | P | 457583.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042071.D | QC | 8 | 2020-02-05 17:02 | 5.6645 | 39442.9265 | 1737.1263 | P | P | P | 41245.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3099 | 31663.801 | 270.009 | P | P | P | 31998.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 16859.8566 | 1762.8582 | F | P | P | 15161.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 503847.545 | 46328.077 | P | P | P | 457584.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042072.D | QC | 9 | 2020-02-05 17:23 | 5.6529 | 40437.1372 | 742.9155 | P | P | P | 41246.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 26944.2104 | 4989.5996 | F | P | P | 31999.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 15579.0894 | 482.0911 | P | P | P | 15162.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 425958.6316 | 31560.8364 | P | P | P | 457585.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042073.D | QC | 10 | 2020-02-05 17:45 | 5.6178 | 34425.7382 | 6754.3146 | F | P | P | 41247.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3038 | 25493.3563 | 6440.4537 | F | P | P | 32000.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 11297.5873 | 3799.4111 | F | F | P | 15163.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 376600.4022 | 80919.0658 | F | P | P | 457586.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042074.D | QC | 11 | 2020-02-05 18:06 | 5.6723 | 39080.7871 | 2099.2657 | P | P | P | 41248.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.31 | 30385.6208 | 1548.1892 | P | P | P | 32001.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 15498.5133 | 401.515 | P | P | P | 15164.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2103 | 480521.37 | 23001.902 | P | P | P | 457587.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042075.D | QC | 12 | 2020-02-05 18:27 | 5.6645 | 38122.8677 | 3057.185 | P | P | P | 41249.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 31427.1749 | 506.6351 | P | P | P | 32002.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.1205 | 15710.7598 | 613.7615 | P | P | P | 15165.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2103 | 475409.8669 | 17890.3989 | P | P | P | 457588.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042076.D | QC | 13 | 2020-02-05 18:49 | 5.6645 | 35398.619 | 5781.4338 | F | P | P | 41250.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.31 | 27321.942 | 4611.868 | F | P | P | 32003.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 12712.2673 | 2384.7311 | F | P | P | 15166.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2103 | 432268.0107 | 25251.4573 | P | P | P | 457589.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042077.D | QC | 14 | 2020-02-05 19:10 | 5.6645 | 37831.2485 | 3348.8042 | P | P | P | 41251.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 29825.7109 | 2108.0991 | P | P | P | 32004.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 13851.1011 | 1245.8973 | P | P | P | 15167.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2103 | 439804.0144 | 17715.4536 | P | P | P | 457590.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC | 7010_1_02042078.D | QC | 15 | 2020-02-05 19:31 | 5.6723 | 42332.7454 | 1152.6926 | P | P | P | 41252.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3069 | 31678.7751 | 255.0349 | P | P | P | 32005.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 14191.3873 | 905.6111 | P | P | P | 15168.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2103 | 471126.658 | 13607.19 | P | P | P | 457591.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342
SpikeBlank | PFAS Group 2 UC | 7010_1_02042079.D | MatrixBlank |  | 2020-02-05 19:53 | 5.6451 | 41562.5619 | 382.5091 | P | P | P | 41253.0528 | 50622.1045 | 31738.001 | 55343.1303 | 27016.9752 | 6.3038 | 33089.5027 | 1155.6927 | P | P | P | 32006.81 | 39655.9461 | 24211.6739 | 43517.0142 | 20350.6058 | 7.124 | 14386.0657 | 710.9326 | P | P | P | 15169.9983 | 18574.3183 | 11619.6784 | 20312.9782 | 9881.0185 | 11.2181 | 474411.87 | 16892.402 | P | P | P | 457592.468 | 557637.6239 | 357401.3121 | 607696.7018 | 307342.2342

## Control Chart_3125
[9 empty] | Mean = | 1093.7923 | [9 empty] | Mean = | 65995.6421
[7 empty] | MFOET Results |  | Stdev = | 137.6958 | [7 empty] | 4NT13C6 Results |  | Stdev = | 7192.6628
Name | Comment | Info. | Data File | Type | Level | Acq. Date-Time | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032001.D | MatrixBlank |  | 2020-03-03 10:22 | 7.2044 | 1056.4203 | 37.372 | P | P | P | 1093.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 69227.1395 | 3231.4974 | P | P | P | 65995.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032002.D | MatrixBlank |  | 2020-03-03 10:42 | 7.1948 | 1177.6927 | 83.9004 | P | P | P | 1094.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 69136.287 | 3140.6449 | P | P | P | 65996.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032003.D | Cal | 1 | 2020-03-03 11:03 | 7.2044 | 986.3179 | 107.4744 | P | P | P | 1095.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60890.6619 | 5104.9802 | P | P | P | 65997.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032004.D | Cal | 2 | 2020-03-03 11:24 | 7.2044 | 1107.3358 | 13.5436 | P | P | P | 1096.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60618.5669 | 5377.0752 | P | P | P | 65998.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032005.D | Cal | 3 | 2020-03-03 11:44 | 7.2044 | 1066.2183 | 27.5739 | P | P | P | 1097.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 58845.5137 | 7150.1283 | P | P | P | 65999.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032006.D | Cal | 4 | 2020-03-03 12:05 | 7.2044 | 1107.5234 | 13.7312 | P | P | P | 1098.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62414.3875 | 3581.2546 | P | P | P | 66000.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032007.D | Cal | 5 | 2020-03-03 12:26 | 7.2044 | 952.686 | 141.1063 | F | P | P | 1099.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 70190.2357 | 4194.5936 | P | P | P | 66001.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032008.D | Cal | 6 | 2020-03-03 12:46 | 7.2044 | 1148.125 | 54.3327 | P | P | P | 1100.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 61496.6705 | 4498.9716 | P | P | P | 66002.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032009.D | Cal | 7 | 2020-03-03 13:07 | 7.2044 | 1022.3564 | 71.4359 | P | P | P | 1101.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60672.392 | 5323.2501 | P | P | P | 66003.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032010.D | Cal | 8 | 2020-03-03 13:27 | 7.2044 | 1121.8964 | 28.1042 | P | P | P | 1102.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62083.9172 | 3911.7249 | P | P | P | 66004.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032011.D | Cal | 9 | 2020-03-03 13:48 | 7.2044 | 1030.3818 | 63.4104 | P | P | P | 1103.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 58889.069 | 7106.5731 | P | P | P | 66005.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032012.D | Cal | 10 | 2020-03-03 14:08 | 7.2044 | 1018.4304 | 75.3619 | P | P | P | 1104.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 55625.898 | 10369.7441 | F | P | P | 66006.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032013.D | Cal | 11 | 2020-03-03 14:29 | 7.2044 | 976.3613 | 117.431 | P | P | P | 1105.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 58865.0055 | 7130.6366 | P | P | P | 66007.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032014.D | Cal | 12 | 2020-03-03 14:49 | 7.2044 | 1166.9621 | 73.1698 | P | P | P | 1106.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 66548.4652 | 552.8232 | P | P | P | 66008.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032015.D | Cal | 13 | 2020-03-03 15:10 | 7.2044 | 1112.2557 | 18.4634 | P | P | P | 1107.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63809.8812 | 2185.7609 | P | P | P | 66009.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032016.D | Cal | 14 | 2020-03-03 15:31 | 7.2044 | 785.519 | 308.2732 | F | F | P | 1108.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60261.4887 | 5734.1534 | P | P | P | 66010.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032017.D | Cal | 15 | 2020-03-03 15:51 | 7.2044 | 1023.6017 | 70.1906 | P | P | P | 1109.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 61011.6393 | 4984.0028 | P | P | P | 66011.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
QCUCGCAACC4_1/30/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032018.D | QC |  | 2020-03-03 16:12 | 7.2044 | 896.2893 | 197.503 | F | P | P | 1110.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 51558.1134 | 14437.5287 | F | F | P | 66012.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
QCUCGCAACC8_1/30/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032019.D | QC |  | 2020-03-03 16:32 | 7.2044 | 987.1494 | 106.6429 | P | P | P | 1111.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 54538.6494 | 11456.9927 | F | P | P | 66013.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
QCUCGCAACC11_1/30/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032020.D | QC |  | 2020-03-03 16:53 | 7.2044 | 779.0683 | 314.7239 | F | F | P | 1112.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 51533.8445 | 14461.7975 | F | F | P | 66014.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032021.D | MatrixBlank |  | 2020-03-03 17:13 | 7.2044 | 1092.9331 | 0.85912 | P | P | P | 1113.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 69595.3358 | 3599.6938 | P | P | P | 66015.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032022.D | Sample |  | 2020-03-03 17:34 | 7.2044 | 957.7967 | 135.9956 | P | P | P | 1114.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60712.3687 | 5283.2734 | P | P | P | 66016.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032023.D | Sample |  | 2020-03-03 17:54 | 7.2044 | 1152.7691 | 58.9768 | P | P | P | 1115.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64554.5236 | 1441.1185 | P | P | P | 66017.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032024.D | Sample |  | 2020-03-03 18:15 | 7.2044 | 1074.0607 | 19.7316 | P | P | P | 1116.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62965.4822 | 3030.1599 | P | P | P | 66018.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032025.D | Sample |  | 2020-03-03 18:36 | 7.2044 | 1155.4846 | 61.6923 | P | P | P | 1117.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 65493.7533 | 501.8888 | P | P | P | 66019.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032026.D | Sample |  | 2020-03-03 18:56 | 7.2044 | 1016.7513 | 77.041 | P | P | P | 1118.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60035.2194 | 5960.4227 | P | P | P | 66020.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032027.D | Sample |  | 2020-03-03 19:17 | 7.2044 | 1168.5197 | 74.7275 | P | P | P | 1119.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63536.5482 | 2459.0938 | P | P | P | 66021.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032029.D | Sample |  | 2020-03-03 19:58 | 7.2044 | 1048.9794 | 44.8129 | P | P | P | 1120.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62316.6427 | 3678.9994 | P | P | P | 66022.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032030.D | Sample |  | 2020-03-03 20:19 | 7.2044 | 1128.3597 | 34.5674 | P | P | P | 1121.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 67914.8518 | 1919.2097 | P | P | P | 66023.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032031.D | QC | 5 | 2020-03-03 20:39 | 7.2044 | 1174.9714 | 81.1792 | P | P | P | 1122.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 74569.2927 | 8573.6506 | F | P | P | 66024.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032032.D | Sample |  | 2020-03-03 21:00 | 7.2044 | 980.8004 | 112.9919 | P | P | P | 1123.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 66124.7663 | 129.1242 | P | P | P | 66025.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032033.D | Sample |  | 2020-03-03 21:20 | 7.2044 | 1057.9274 | 35.8649 | P | P | P | 1124.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64372.0321 | 1623.61 | P | P | P | 66026.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032034.D | Sample |  | 2020-03-03 21:41 | 7.2044 | 1067.271 | 26.5213 | P | P | P | 1125.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 67728.8416 | 1733.1995 | P | P | P | 66027.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032035.D | MatrixBlank |  | 2020-03-03 22:01 | 7.2044 | 1357.661 | 263.8687 | F | P | P | 1126.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 72751.5903 | 6755.9482 | P | P | P | 66028.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032036.D | Sample |  | 2020-03-03 22:22 | 7.214 | 999.1432 | 94.6491 | P | P | P | 1127.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63180.8596 | 2814.7825 | P | P | P | 66029.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032037.D | Sample |  | 2020-03-03 22:42 | 7.214 | 970.9798 | 122.8125 | P | P | P | 1128.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64105.665 | 1889.9771 | P | P | P | 66030.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS1B_GC 2/3_Dup | PFAS Group 2 UC |  | 7010_2_03032038.D | Sample |  | 2020-03-03 23:03 | 7.214 | 1086.4282 | 7.3641 | P | P | P | 1129.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64503.8845 | 1491.7576 | P | P | P | 66031.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032039.D | Sample |  | 2020-03-03 23:24 | 7.214 | 1198.3503 | 104.558 | P | P | P | 1130.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 66141.9881 | 146.346 | P | P | P | 66032.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032040.D | MatrixBlank |  | 2020-03-03 23:44 | 7.2044 | 1190.4855 | 96.6932 | P | P | P | 1131.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 72595.3435 | 6599.7014 | P | P | P | 66033.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032041.D | Sample |  | 2020-03-04 00:05 | 7.214 | 1193.0884 | 99.2962 | P | P | P | 1132.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63038.4085 | 2957.2336 | P | P | P | 66034.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032042.D | QC | 9 | 2020-03-04 00:25 | 7.214 | 962.8035 | 130.9888 | P | P | P | 1133.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 61424.0237 | 4571.6183 | P | P | P | 66035.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032043.D | Sample |  | 2020-03-04 00:46 | 7.214 | 1067.8575 | 25.9347 | P | P | P | 1134.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 59343.6031 | 6652.039 | P | P | P | 66036.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032045.D | Sample |  | 2020-03-04 01:27 | 7.214 | 948.4752 | 145.3171 | F | P | P | 1135.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 60077.9772 | 5917.6649 | P | P | P | 66037.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032046.D | Sample |  | 2020-03-04 01:48 | 7.2236 | 987.0309 | 106.7613 | P | P | P | 1136.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 61198.2692 | 4797.3729 | P | P | P | 66038.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032047.D | Sample |  | 2020-03-04 02:08 | 7.2044 | 1156.2108 | 62.4186 | P | P | P | 1137.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 68236.6089 | 2240.9668 | P | P | P | 66039.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032048.D | MatrixBlank |  | 2020-03-04 02:29 | 7.2044 | 1292.0009 | 198.2087 | F | P | P | 1138.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 74105.0678 | 8109.4257 | F | P | P | 66040.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032049.D | Sample |  | 2020-03-04 02:49 | 7.2044 | 1172.5662 | 78.7739 | P | P | P | 1139.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 77654.9123 | 11659.2702 | F | P | P | 66041.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032050.D | Sample |  | 2020-03-04 03:10 | 7.214 | 1122.5287 | 28.7364 | P | P | P | 1140.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64975.655 | 1019.9871 | P | P | P | 66042.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032051.D | Sample |  | 2020-03-04 03:30 | 7.2044 | 1037.5133 | 56.279 | P | P | P | 1141.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63750.1201 | 2245.5219 | P | P | P | 66043.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032052.D | Sample |  | 2020-03-04 03:51 | 7.214 | 1227.6331 | 133.8408 | P | P | P | 1142.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 69322.8156 | 3327.1736 | P | P | P | 66044.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032053.D | Sample |  | 2020-03-04 04:11 | 7.214 | 982.7957 | 110.9966 | P | P | P | 1143.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 67306.1309 | 1310.4888 | P | P | P | 66045.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032054.D | QC | 12 | 2020-03-04 04:32 | 7.2044 | 1254.3954 | 160.6032 | F | P | P | 1144.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 72837.0995 | 6841.4574 | P | P | P | 66046.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032055.D | MatrixBlank |  | 2020-03-04 04:52 | 7.2044 | 1145.5744 | 51.7821 | P | P | P | 1145.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 71460.2807 | 5464.6386 | P | P | P | 66047.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032056.D | Sample |  | 2020-03-04 05:13 | 7.214 | 893.1925 | 200.5998 | F | P | P | 1146.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 67109.6281 | 1113.986 | P | P | P | 66048.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032057.D | Sample |  | 2020-03-04 05:34 | 7.214 | 1069.5558 | 24.2365 | P | P | P | 1147.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 64425.9332 | 1569.7089 | P | P | P | 66049.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032058.D | Sample |  | 2020-03-04 05:54 | 7.214 | 950.6351 | 143.1572 | F | P | P | 1148.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63631.4023 | 2364.2398 | P | P | P | 66050.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032059.D | MatrixBlank |  | 2020-03-04 06:15 | 7.2044 | 1145.2775 | 51.4852 | P | P | P | 1149.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 71114.8603 | 5119.2182 | P | P | P | 66051.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC1_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032060.D | QC | 1 | 2020-03-04 06:35 | 7.214 | 1275.6608 | 181.8686 | F | P | P | 1150.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 81763.2966 | 15767.6545 | F | F | P | 66052.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC2_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032061.D | QC | 2 | 2020-03-04 06:56 | 7.214 | 1094.8949 | 1.1026 | P | P | P | 1151.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 65328.5779 | 667.0642 | P | P | P | 66053.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC3_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032062.D | QC | 3 | 2020-03-04 07:16 | 7.2044 | 1213.125 | 119.3327 | P | P | P | 1152.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 74817.3823 | 8821.7402 | F | P | P | 66054.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC4_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032063.D | QC | 4 | 2020-03-04 07:37 | 7.214 | 1334.7875 | 240.9952 | F | P | P | 1153.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 86684.9648 | 20689.3227 | F | F | P | 66055.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC5_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032064.D | QC | 5 | 2020-03-04 07:57 | 7.214 | 1395.5091 | 301.7169 | F | F | P | 1154.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 73487.5397 | 7491.8977 | F | P | P | 66056.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC6_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032065.D | QC | 6 | 2020-03-04 08:18 | 7.2044 | 1570.0219 | 476.2296 | F | F | F | 1155.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 91419.6575 | 25424.0155 | F | F | F | 66057.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC7_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032066.D | QC | 7 | 2020-03-04 08:39 | 7.2044 | 1101.8109 | 8.0186 | P | P | P | 1156.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 66889.4446 | 893.8025 | P | P | P | 66058.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC8_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032067.D | QC | 8 | 2020-03-04 08:59 | 7.2044 | 1052.4966 | 41.2957 | P | P | P | 1157.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 71367.9274 | 5372.2854 | P | P | P | 66059.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC9_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032068.D | QC | 9 | 2020-03-04 09:20 | 7.2044 | 996.2264 | 97.5659 | P | P | P | 1158.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62091.3733 | 3904.2688 | P | P | P | 66060.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC10_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032069.D | QC | 10 | 2020-03-04 09:40 | 7.214 | 960.8637 | 132.9285 | P | P | P | 1159.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 57862.6432 | 8132.9989 | F | P | P | 66061.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC11_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032070.D | QC | 11 | 2020-03-04 10:01 | 7.214 | 1078.6817 | 15.1105 | P | P | P | 1160.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 62841.797 | 3153.8451 | P | P | P | 66062.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC12_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032071.D | QC | 12 | 2020-03-04 10:21 | 7.214 | 1169.3138 | 75.5216 | P | P | P | 1161.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 65503.3161 | 492.326 | P | P | P | 66063.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC13_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032072.D | QC | 13 | 2020-03-04 10:42 | 7.214 | 1063.9294 | 29.8629 | P | P | P | 1162.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 67366.9217 | 1371.2796 | P | P | P | 66064.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC14_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032073.D | QC | 14 | 2020-03-04 11:03 | 7.214 | 1486.7901 | 392.9978 | F | F | P | 1163.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 83707.3707 | 17711.7286 | F | F | P | 66065.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
UCGCAACC15_1/29/20_GC 2/3 | PFAS Group 2 UC |  | 7010_2_03032074.D | QC | 15 | 2020-03-04 11:23 | 7.214 | 1023.8407 | 69.9516 | P | P | P | 1164.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 63586.0781 | 2409.564 | P | P | P | 66066.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536
SpikeBlank | PFAS Group 2 UC |  | 7010_2_03032075.D | MatrixBlank |  | 2020-03-04 11:44 | 7.2044 | 1017.4846 | 76.3077 | P | P | P | 1165.7923 | 1369.1839 | 818.4006 | 1506.8798 | 680.7047 | 11.4968 | 70533.9696 | 4538.3275 | P | P | P | 66067.6421 | 80380.9677 | 51610.3165 | 87573.6305 | 44417.6536

## Control Chart_899
[9 empty] | Mean = | 50266.7455 | [9 empty] | Mean = | 539587.979
[6 empty] | MFBET (ISTD) Results |  |  | Stdev = | 4953.5696 | [6 empty] | 4NT13C6 (ISTD) Results |  |  | Stdev = | 54114.3705
Name | Comment | Data File | Type | Level | Acq. Date-Time | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | RT | Area | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s
SpikeBlank | PFAS Group 2 UC | 7010_1_020242001.D | MatrixBlank |  | 2020-02-24 15:46 | 5.6295 | 48592.6222 | 1674.1233 | P | P | P | 50266.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2181 | 610704.7863 | 71116.8073 | F | P | P | 539587.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242002.D | MatrixBlank |  | 2020-02-24 16:07 | 5.6295 | 45056.2853 | 5210.4601 | F | P | P | 50267.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 547939.7252 | 8351.7461 | P | P | P | 539588.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242003.D | Cal | 1 | 2020-02-24 16:29 | 5.6178 | 45159.1111 | 5107.6343 | F | P | P | 50268.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2181 | 416578.0193 | 123009.9597 | F | F | P | 539589.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242004.D | Cal | 2 | 2020-02-24 16:50 | 5.6334 | 41986.7862 | 8279.9593 | F | P | P | 50269.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2181 | 432368.7952 | 107219.1839 | F | P | P | 539590.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242005.D | Cal | 3 | 2020-02-24 17:11 | 5.6451 | 47553.3329 | 2713.4126 | P | P | P | 50270.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2181 | 508203.9671 | 31384.0119 | P | P | P | 539591.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242006.D | Cal | 4 | 2020-02-24 17:32 | 5.6451 | 53346.989 | 3080.2436 | P | P | P | 50271.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 505560.7045 | 34027.2746 | P | P | P | 539592.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242007.D | Cal | 5 | 2020-02-24 17:54 | 5.6451 | 50078.2096 | 188.5358 | P | P | P | 50272.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 539961.3813 | 373.4023 | P | P | P | 539593.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242008.D | Cal | 6 | 2020-02-24 18:15 | 5.6567 | 43827.3109 | 6439.4345 | F | P | P | 50273.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 509571.151 | 30016.828 | P | P | P | 539594.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242009.D | Cal | 7 | 2020-02-24 18:36 | 5.6295 | 43628.2816 | 6638.4639 | F | P | P | 50274.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 487326.1154 | 52261.8637 | P | P | P | 539595.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242010.D | Cal | 8 | 2020-02-24 18:57 | 5.6334 | 48761.7652 | 1504.9803 | P | P | P | 50275.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2181 | 504469.2157 | 35118.7633 | P | P | P | 539596.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242011.D | Cal | 9 | 2020-02-24 19:19 | 5.6489 | 52843.4966 | 2576.7511 | P | P | P | 50276.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 552775.3856 | 13187.4066 | P | P | P | 539597.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242012.D | Cal | 10 | 2020-02-24 19:40 | 5.649 | 50318.8843 | 52.1389 | P | P | P | 50277.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2104 | 543507.0176 | 3919.0386 | P | P | P | 539598.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242013.D | Cal | 11 | 2020-02-24 20:01 | 5.6528 | 49969.8291 | 296.9163 | P | P | P | 50278.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 523415.4418 | 16172.5372 | P | P | P | 539599.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242014.D | Cal | 12 | 2020-02-24 20:23 | 5.649 | 50683.9632 | 417.2178 | P | P | P | 50279.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 526997.7151 | 12590.2639 | P | P | P | 539600.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242015.D | Cal | 13 | 2020-02-24 20:44 | 5.6451 | 51176.2387 | 909.4932 | P | P | P | 50280.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 546498.4929 | 6910.5138 | P | P | P | 539601.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242016.D | Cal | 14 | 2020-02-24 21:05 | 5.6528 | 51434.5275 | 1167.782 | P | P | P | 50281.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 547146.9709 | 7558.9919 | P | P | P | 539602.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242017.D | Cal | 15 | 2020-02-24 21:26 | 5.6412 | 47441.556 | 2825.1894 | P | P | P | 50282.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 542902.5257 | 3314.5467 | P | P | P | 539603.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242018.D | MatrixBlank |  | 2020-02-24 21:48 | 5.614 | 52230.0349 | 1963.2894 | P | P | P | 50283.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 579489.9307 | 39901.9517 | P | P | P | 539604.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
QCUCGCAACC4_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242019.D | QC | 4 | 2020-02-24 22:09 | 5.6528 | 52147.8654 | 1881.1199 | P | P | P | 50284.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 537799.0956 | 1788.8834 | P | P | P | 539605.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
QCUCGCAACC8_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242020.D | QC | 8 | 2020-02-24 22:30 | 5.6451 | 52724.8583 | 2458.1128 | P | P | P | 50285.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2104 | 557972.9033 | 18384.9243 | P | P | P | 539606.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
QCUCGCAACC11_1/30/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242021.D | QC | 11 | 2020-02-24 22:51 | 5.6178 | 53816.5575 | 3549.812 | P | P | P | 50286.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 586119.3129 | 46531.3338 | P | P | P | 539607.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242022.D | MatrixBlank |  | 2020-02-24 23:12 | 5.6295 | 56865.3096 | 6598.5642 | F | P | P | 50287.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 605881.7141 | 66293.7351 | F | P | P | 539608.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242023.D | QC | 3 | 2020-02-24 23:34 | 5.6373 | 47278.5083 | 2988.2371 | P | P | P | 50288.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 536670.3306 | 2917.6484 | P | P | P | 539609.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T1hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242024.D | Sample |  | 2020-02-24 23:55 | 5.649 | 52552.468 | 2285.7226 | P | P | P | 50289.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 558534.5479 | 18946.5689 | P | P | P | 539610.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T1hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242025.D | Sample |  | 2020-02-25 00:16 | 5.6528 | 52540.6327 | 2273.8872 | P | P | P | 50290.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 616288.7193 | 76700.7403 | F | P | P | 539611.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T5hS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242026.D | Sample |  | 2020-02-25 00:37 | 5.6529 | 57340.2311 | 7073.4857 | F | P | P | 50291.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 641533.1995 | 101945.2204 | F | P | P | 539612.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T5hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242028.D | Sample |  | 2020-02-25 01:20 | 5.649 | 56418.084 | 6151.3385 | F | P | P | 50292.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 594617.8429 | 55029.8639 | F | P | P | 539613.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242029.D | QC | 9 | 2020-02-25 01:41 | 5.6412 | 48687.3735 | 1579.3719 | P | P | P | 50293.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 558303.7836 | 18715.8046 | P | P | P | 539614.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213AFS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242030.D | Sample |  | 2020-02-25 02:03 | 5.6334 | 52272.851 | 2006.1055 | P | P | P | 50294.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 629356.6386 | 89768.6596 | F | P | P | 539615.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213AFS3B_GC 2/19 | PFAS Group 2 UC | 7010_1_020242031.D | Sample |  | 2020-02-25 02:24 | 5.6528 | 58919.7768 | 8653.0314 | F | P | P | 50295.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 610449.5181 | 70861.539 | F | P | P | 539616.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242032.D | MatrixBlank |  | 2020-02-25 02:45 | 5.614 | 55026.6449 | 4759.8995 | P | P | P | 50296.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 606801.2007 | 67213.2217 | F | P | P | 539617.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242033.D | QC | 13 | 2020-02-25 03:07 | 5.6451 | 52147.894 | 1881.1485 | P | P | P | 50297.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 568572.0123 | 28984.0333 | P | P | P | 539618.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T5hS3A_GC 2/19 | PFAS Group 2 UC | 7010_1_020242034.D | Sample |  | 2020-02-25 03:28 | 5.649 | 56798.111 | 6531.3655 | F | P | P | 50298.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 633548.0231 | 93960.0441 | F | P | P | 539619.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213T1hS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242035.D | Sample |  | 2020-02-25 03:49 | 5.649 | 51036.8422 | 770.0968 | P | P | P | 50299.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 594484.2399 | 54896.2609 | F | P | P | 539620.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242036.D | MatrixBlank |  | 2020-02-25 04:11 | 5.614 | 56151.808 | 5885.0625 | F | P | P | 50300.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 561248.775 | 21660.796 | P | P | P | 539621.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGC0213AFS3C_GC 2/19 | PFAS Group 2 UC | 7010_1_020242037.D | Sample |  | 2020-02-25 04:32 | 5.6373 | 58626.738 | 8359.9925 | F | P | P | 50301.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 615379.3913 | 75791.4123 | F | P | P | 539622.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC1_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242039.D | QC | 1 | 2020-02-25 05:15 | 5.6101 | 48812.6514 | 1454.094 | P | P | P | 50302.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 480696.5929 | 58891.3861 | F | P | P | 539623.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC2_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242040.D | QC | 2 | 2020-02-25 05:36 | 5.6412 | 51153.7274 | 886.982 | P | P | P | 50303.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 517542.2318 | 22045.7472 | P | P | P | 539624.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC3_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242041.D | QC | 3 | 2020-02-25 05:57 | 5.6295 | 48275.7681 | 1990.9773 | P | P | P | 50304.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 513192.2188 | 26395.7603 | P | P | P | 539625.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC4_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242042.D | QC | 4 | 2020-02-25 06:19 | 5.6451 | 55478.5799 | 5211.8345 | F | P | P | 50305.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 469844.1334 | 69743.8457 | F | P | P | 539626.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC5_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242043.D | QC | 5 | 2020-02-25 06:40 | 5.6295 | 38630.6526 | 11636.0928 | F | F | P | 50306.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 501681.938 | 37906.0411 | P | P | P | 539627.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC6_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242044.D | QC | 6 | 2020-02-25 07:01 | 5.6412 | 53645.5089 | 3378.7635 | P | P | P | 50307.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 535594.8155 | 3993.1635 | P | P | P | 539628.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC7_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242045.D | QC | 7 | 2020-02-25 07:23 | 5.6062 | 37816.1574 | 12450.5881 | F | F | P | 50308.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 420363.2797 | 119224.6994 | F | F | P | 539629.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC8_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242046.D | QC | 8 | 2020-02-25 07:44 | 5.6296 | 46257.3871 | 4009.3584 | P | P | P | 50309.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2104 | 465857.2548 | 73730.7242 | F | P | P | 539630.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC9_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242047.D | QC | 9 | 2020-02-25 08:05 | 5.6334 | 44343.7186 | 5923.0269 | F | P | P | 50310.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 445139.1458 | 94448.8332 | F | P | P | 539631.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC10_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242048.D | QC | 10 | 2020-02-25 08:27 | 5.6412 | 51808.0583 | 1541.3129 | P | P | P | 50311.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 526109.7482 | 13478.2308 | P | P | P | 539632.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC11_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242049.D | QC | 11 | 2020-02-25 08:48 | 5.6412 | 44392.9055 | 5873.8399 | F | P | P | 50312.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 483052.7241 | 56535.255 | F | P | P | 539633.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC12_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242050.D | QC | 12 | 2020-02-25 09:09 | 5.649 | 52888.6489 | 2621.9035 | P | P | P | 50313.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 543573.5052 | 3985.5261 | P | P | P | 539634.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC13_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242051.D | QC | 13 | 2020-02-25 09:31 | 5.6489 | 55103.7877 | 4837.0423 | P | P | P | 50314.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 513360.6541 | 26227.3249 | P | P | P | 539635.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC14_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242052.D | QC | 14 | 2020-02-25 09:52 | 5.6451 | 52945.9941 | 2679.2487 | P | P | P | 50315.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 562216.0944 | 22628.1154 | P | P | P | 539636.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
UCGCAACC15_2/19/20_GC 2/19 | PFAS Group 2 UC | 7010_1_020242053.D | QC | 15 | 2020-02-25 10:13 | 5.6373 | 40411.6953 | 9855.0502 | F | P | P | 50316.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 519648.7958 | 19939.1832 | P | P | P | 539637.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676
SpikeBlank | PFAS Group 2 UC | 7010_1_020242054.D | MatrixBlank |  | 2020-02-25 10:35 | 5.6023 | 46463.7434 | 3803.002 | P | P | P | 50317.7455 | 60173.8846 | 40359.6063 | 65127.4542 | 35406.0367 | 11.2103 | 521723.1868 | 17864.7922 | P | P | P | 539638.979 | 647816.72 | 431359.2381 | 701931.0905 | 377244.8676

## ValueList_Helper (hidden)
Sample
Blank
Cal
QC
CC
DoubleBlank
MatrixSpike
MatrixSpikeDup
MatrixBlank
TuneCheck
ResponseCheck

## ControlChart_899&900_MFOET
Sample | [7 empty] | MFOET Method | MFOET Results | [10 empty] | Mean = | 1403.485
| Name | Data Path | Comment | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | Stdev = | 345.2846
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420001.D | MatrixBlank |  | 2020-02-04 14:17 | 12000 | 7.1852 | 1678.5451 | Difference | 1s | 2s | 3s | 1403.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420002.D | MatrixBlank |  | 2020-02-04 14:37 | 12000 | 7.1852 | 1540.1071 | 136.6221 | P | P | P | 1404.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC1_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420003.D | Cal | 1 | 2020-02-04 14:58 | 12000 | 7.1852 | 1475.6344 | 72.1494 | P | P | P | 1405.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC2_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420004.D | Cal | 2 | 2020-02-04 15:19 | 12000 | 7.1852 | 1386.7954 | 16.6896 | P | P | P | 1406.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC3_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420005.D | Cal | 3 | 2020-02-04 15:39 | 12000 | 7.1852 | 1400.8425 | 2.6425 | P | P | P | 1407.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420006.D | Cal | 4 | 2020-02-04 16:00 | 12000 | 7.1852 | 1382.2084 | 21.2766 | P | P | P | 1408.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC5_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420007.D | Cal | 5 | 2020-02-04 16:20 | 12000 | 7.1852 | 1355.2381 | 48.2469 | P | P | P | 1409.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC6_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420008.D | Cal | 6 | 2020-02-04 16:41 | 12000 | 7.1852 | 1095.8063 | 307.6787 | P | P | P | 1410.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC7_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420009.D | Cal | 7 | 2020-02-04 17:01 | 12000 | 7.1852 | 1453.9224 | 50.4374 | P | P | P | 1411.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420010.D | Cal | 8 | 2020-02-04 17:22 | 12000 | 7.1852 | 1352.7667 | 50.7183 | P | P | P | 1412.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420011.D | Cal | 9 | 2020-02-04 17:42 | 12000 | 7.1852 | 1440.3343 | 36.8493 | P | P | P | 1413.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC10_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420012.D | Cal | 10 | 2020-02-04 18:03 | 12000 | 7.1852 | 1572.8193 | 169.3343 | P | P | P | 1414.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420013.D | Cal | 11 | 2020-02-04 18:23 | 12000 | 7.1852 | 1693.0594 | 289.5744 | P | P | P | 1415.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC12_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420014.D | Cal | 12 | 2020-02-04 18:44 | 12000 | 7.1948 | 1675.9603 | 272.4753 | P | P | P | 1416.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420015.D | Cal | 13 | 2020-02-04 19:05 | 12000 | 7.1948 | 1303.5317 | 99.9533 | P | P | P | 1417.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC14_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420016.D | Cal | 14 | 2020-02-04 19:25 | 12000 | 7.1948 | 1468.8075 | 65.3225 | P | P | P | 1418.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC15_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420017.D | Cal | 15 | 2020-02-04 19:46 | 12000 | 7.1852 | 1953.1702 | 549.6852 | F | P | P | 1419.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420018.D | MatrixBlank |  | 2020-02-04 20:06 | 12000 | 7.1852 | 1702.1964 | 298.7114 | P | P | P | 1420.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| QCUCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420019.D | QC | 4 | 2020-02-04 20:27 | 12000 | 7.1852 | 1462.6929 | 59.2079 | P | P | P | 1421.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| QCUCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420020.D | QC | 8 | 2020-02-04 20:47 | 12000 | 7.1852 | 1489.8128 | 86.3278 | P | P | P | 1422.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| QCUCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420021.D | QC | 11 | 2020-02-04 21:08 | 12000 | 7.1948 | 1436.0345 | 32.5495 | P | P | P | 1423.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420022.D | MatrixBlank |  | 2020-02-04 21:28 | 12000 | 7.1852 | 1892.198 | 488.713 | F | P | P | 1424.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T1hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420023.D | Sample |  | 2020-02-04 21:49 | 12000 | 7.214 | 0 | 1403.485 | F | F | F | 1425.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T1hS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420024.D | Sample |  | 2020-02-04 22:10 | 12000 | 7.1948 | 1337.4599 | 66.0251 | P | P | P | 1426.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T1hS3B_GC 2/3_dup | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420025.D | Sample |  | 2020-02-04 22:30 | 12000 | 7.1948 | 1520.5123 | 117.0273 | P | P | P | 1427.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T5hS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420026.D | Sample |  | 2020-02-04 22:51 | 12000 | 7.1948 | 1341.0442 | 62.4408 | P | P | P | 1428.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420028.D | QC | 9 | 2020-02-04 23:32 | 12000 | 7.1948 | 1504.6128 | 101.1278 | P | P | P | 1429.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T5hS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420029.D | Sample |  | 2020-02-04 23:52 | 12000 | 7.1948 | 1843.1399 | 439.6549 | F | P | P | 1430.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3AFS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420030.D | Sample |  | 2020-02-05 00:13 | 12000 | 7.1948 | 1526.157 | 122.672 | P | P | P | 1431.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3AFS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420031.D | Sample |  | 2020-02-05 00:34 | 12000 | 7.1948 | 1269.7729 | 133.7121 | P | P | P | 1432.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420032.D | MatrixBlank |  | 2020-02-05 00:54 | 12000 | 7.1948 | 1457.733 | 54.248 | P | P | P | 1433.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420033.D | QC | 13 | 2020-02-05 01:15 | 12000 | 7.1948 | 1388.96 | 14.525 | P | P | P | 1434.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T5hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420034.D | Sample |  | 2020-02-05 01:35 | 12000 | 7.1948 | 1667.3832 | 263.8982 | P | P | P | 1435.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T1hS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420035.D | Sample |  | 2020-02-05 01:56 | 12000 | 7.1948 | 1408.1346 | 4.6496 | P | P | P | 1436.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3AFS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420036.D | Sample |  | 2020-02-05 02:16 | 12000 | 7.1948 | 1866.3611 | 462.8761 | F | P | P | 1437.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420037.D | MatrixBlank |  | 2020-02-05 02:37 | 12000 | 7.1948 | 1509.2634 | 105.7784 | P | P | P | 1438.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC1_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420038.D | QC | 1 | 2020-02-05 02:57 | 12000 | 7.1948 | 1325.0316 | 78.4534 | P | P | P | 1439.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC2_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420039.D | QC | 2 | 2020-02-05 03:18 | 12000 | 7.1948 | 1111.9583 | 291.5267 | P | P | P | 1440.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC3_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420040.D | QC | 3 | 2020-02-05 03:38 | 12000 | 7.1948 | 1135.115 | 268.37 | P | P | P | 1441.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420041.D | QC | 4 | 2020-02-05 03:59 | 12000 | 7.1948 | 1310.4551 | 93.0299 | P | P | P | 1442.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC5_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420042.D | QC | 5 | 2020-02-05 04:20 | 12000 | 7.1948 | 1384.4092 | 19.0758 | P | P | P | 1443.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC6_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420043.D | QC | 6 | 2020-02-05 04:40 | 12000 | 7.1948 | 1087.8318 | 315.6532 | P | P | P | 1444.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC7_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420044.D | QC | 7 | 2020-02-05 05:01 | 12000 | 7.1948 | 1323.8509 | 79.6341 | P | P | P | 1445.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420045.D | QC | 8 | 2020-02-05 05:21 | 12000 | 7.1948 | 1266.946 | 136.539 | P | P | P | 1446.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420046.D | QC | 9 | 2020-02-05 05:42 | 12000 | 7.1948 | 1492.5478 | 89.0628 | P | P | P | 1447.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC10_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420047.D | QC | 10 | 2020-02-05 06:02 | 12000 | 7.1948 | 1650.1718 | 246.6868 | P | P | P | 1448.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420048.D | QC | 11 | 2020-02-05 06:23 | 12000 | 7.1948 | 1466.8161 | 63.3311 | P | P | P | 1449.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC12_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420049.D | QC | 12 | 2020-02-05 06:44 | 12000 | 7.1948 | 1429.3358 | 25.8508 | P | P | P | 1450.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420050.D | QC | 13 | 2020-02-05 07:04 | 12000 | 7.1948 | 1397.2523 | 6.2327 | P | P | P | 1451.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC14_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420051.D | QC | 14 | 2020-02-05 07:25 | 12000 | 7.1948 | 1524.4085 | 120.9235 | P | P | P | 1452.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3CC15_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420052.D | QC | 15 | 2020-02-05 07:45 | 12000 | 7.3481 | 0 | 1403.485 | F | F | F | 1453.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420053.D | MatrixBlank |  | 2020-02-05 08:06 | 12000 | 7.1948 | 1669.7925 | 266.3075 | P | P | P | 1454.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313
| UCGC3T1hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420055.D | Sample |  | 2020-02-05 08:47 | 12000 | 7.1948 | 955.764 | 447.721 | F | P | P | 1455.485 | 2094.0542 | 712.9158 | 2439.3387 | 367.6313

## ControlChart_899&900_4NT13C6
Sample | [7 empty] | 4NT13C6 Method | 4NT13C6 Results | [10 empty] | Mean = | 56181.0063
| Name | Data Path | Comment | Data File | Type | Level | Acq. Date-Time | Exp. Conc. | RT | Resp. | Difference | 1s | 2s | 3s | Mean | 2s | -2s | 3s | -3s | Stdev = | 13705.9559
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420001.D | MatrixBlank |  | 2020-02-04 14:17 | 12000 | 11.4954 | 61305.5363 | 5124.5301 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420002.D | MatrixBlank |  | 2020-02-04 14:37 | 12000 | 11.4955 | 61022.7272 | 4841.721 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC1_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420003.D | Cal | 1 | 2020-02-04 14:58 | 12000 | 11.4954 | 54868.3616 | 1312.6446 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC2_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420004.D | Cal | 2 | 2020-02-04 15:19 | 12000 | 11.4955 | 51994.6133 | 4186.393 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC3_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420005.D | Cal | 3 | 2020-02-04 15:39 | 12000 | 11.4954 | 53724.3593 | 2456.6469 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420006.D | Cal | 4 | 2020-02-04 16:00 | 12000 | 11.4955 | 51892.681 | 4288.3252 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC5_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420007.D | Cal | 5 | 2020-02-04 16:20 | 12000 | 11.4954 | 54760.6056 | 1420.4007 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC6_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420008.D | Cal | 6 | 2020-02-04 16:41 | 12000 | 11.4955 | 51389.3208 | 4791.6855 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC7_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420009.D | Cal | 7 | 2020-02-04 17:01 | 12000 | 11.4954 | 55999.1071 | 181.8991 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420010.D | Cal | 8 | 2020-02-04 17:22 | 12000 | 11.4955 | 52774.4391 | 3406.5671 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420011.D | Cal | 9 | 2020-02-04 17:42 | 12000 | 11.4954 | 52424.0015 | 3757.0047 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC10_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420012.D | Cal | 10 | 2020-02-04 18:03 | 12000 | 11.4955 | 67086.5711 | 10905.5648 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420013.D | Cal | 11 | 2020-02-04 18:23 | 12000 | 11.4954 | 65928.9361 | 9747.9298 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC12_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420014.D | Cal | 12 | 2020-02-04 18:44 | 12000 | 11.4955 | 58733.6975 | 2552.6913 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420015.D | Cal | 13 | 2020-02-04 19:05 | 12000 | 11.4954 | 59257.3607 | 3076.3544 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC14_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420016.D | Cal | 14 | 2020-02-04 19:25 | 12000 | 11.4955 | 56407.1932 | 226.187 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC15_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420017.D | Cal | 15 | 2020-02-04 19:46 | 12000 | 11.4954 | 96074.1978 | 39893.1915 | F | F | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420018.D | MatrixBlank |  | 2020-02-04 20:06 | 12000 | 11.4955 | 61166.3809 | 4985.3747 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| QCUCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420019.D | QC | 4 | 2020-02-04 20:27 | 12000 | 11.4954 | 64723.7441 | 8542.7378 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| QCUCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420020.D | QC | 8 | 2020-02-04 20:47 | 12000 | 11.4955 | 65285.6514 | 9104.6451 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| QCUCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420021.D | QC | 11 | 2020-02-04 21:08 | 12000 | 11.4954 | 61245.5627 | 5064.5564 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420022.D | MatrixBlank |  | 2020-02-04 21:28 | 12000 | 11.4955 | 64416.1022 | 8235.096 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T1hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420023.D | Sample |  | 2020-02-04 21:49 | 12000 | 11.5014 | 0 | 56181.0063 | F | F | F | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T1hS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420024.D | Sample |  | 2020-02-04 22:10 | 12000 | 11.4955 | 64386.3302 | 8205.3239 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T1hS3B_GC 2/3_dup | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420025.D | Sample |  | 2020-02-04 22:30 | 12000 | 11.4954 | 60885.2332 | 4704.227 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T5hS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420026.D | Sample |  | 2020-02-04 22:51 | 12000 | 11.4955 | 61259.6653 | 5078.659 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420028.D | QC | 9 | 2020-02-04 23:32 | 12000 | 11.4955 | 56723.4922 | 542.4859 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T5hS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420029.D | Sample |  | 2020-02-04 23:52 | 12000 | 11.4954 | 74046.2217 | 17865.2155 | F | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3AFS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420030.D | Sample |  | 2020-02-05 00:13 | 12000 | 11.4955 | 59391.3479 | 3210.3416 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3AFS3B_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420031.D | Sample |  | 2020-02-05 00:34 | 12000 | 11.4954 | 55640.4423 | 540.564 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420032.D | MatrixBlank |  | 2020-02-05 00:54 | 12000 | 11.4955 | 57640.7532 | 1459.747 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420033.D | QC | 13 | 2020-02-05 01:15 | 12000 | 11.4954 | 51469.1702 | 4711.8361 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T5hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420034.D | Sample |  | 2020-02-05 01:35 | 12000 | 11.4955 | 64139.5691 | 7958.5628 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T1hS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420035.D | Sample |  | 2020-02-05 01:56 | 12000 | 11.4954 | 59747.7672 | 3566.7609 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3AFS3C_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420036.D | Sample |  | 2020-02-05 02:16 | 12000 | 11.4955 | 69473.5789 | 13292.5727 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420037.D | MatrixBlank |  | 2020-02-05 02:37 | 12000 | 11.4954 | 57444.1318 | 1263.1256 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC1_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420038.D | QC | 1 | 2020-02-05 02:57 | 12000 | 11.4955 | 56752.0246 | 571.0183 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC2_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420039.D | QC | 2 | 2020-02-05 03:18 | 12000 | 11.4954 | 49535.2038 | 6645.8024 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC3_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420040.D | QC | 3 | 2020-02-05 03:38 | 12000 | 11.4955 | 53101.5278 | 3079.4784 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC4_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420041.D | QC | 4 | 2020-02-05 03:59 | 12000 | 11.4954 | 52937.9004 | 3243.1059 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC5_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420042.D | QC | 5 | 2020-02-05 04:20 | 12000 | 11.4955 | 52147.6965 | 4033.3097 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC6_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420043.D | QC | 6 | 2020-02-05 04:40 | 12000 | 11.4954 | 51778.2372 | 4402.769 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC7_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420044.D | QC | 7 | 2020-02-05 05:01 | 12000 | 11.4955 | 54433.2288 | 1747.7775 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC8_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420045.D | QC | 8 | 2020-02-05 05:21 | 12000 | 11.4954 | 48426.2058 | 7754.8004 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC9_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420046.D | QC | 9 | 2020-02-05 05:42 | 12000 | 11.4955 | 57364.2828 | 1183.2765 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC10_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420047.D | QC | 10 | 2020-02-05 06:02 | 12000 | 11.4954 | 60926.6464 | 4745.6402 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC11_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420048.D | QC | 11 | 2020-02-05 06:23 | 12000 | 11.4955 | 60268.444 | 4087.4378 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC12_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420049.D | QC | 12 | 2020-02-05 06:44 | 12000 | 11.4954 | 58074.7273 | 1893.7211 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC13_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420050.D | QC | 13 | 2020-02-05 07:04 | 12000 | 11.4955 | 55538.0057 | 643.0006 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC14_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420051.D | QC | 14 | 2020-02-05 07:25 | 12000 | 11.4954 | 55900.2573 | 280.749 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3CC15_1/17/20_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420052.D | QC | 15 | 2020-02-05 07:45 | 12000 | 11.4895 | 0 | 56181.0063 | F | F | F | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| SpikeBlank | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420053.D | MatrixBlank |  | 2020-02-05 08:06 | 12000 | 11.4954 | 56349.6778 | 168.6716 | P | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384
| UCGC3T1hS3A_GC 2/3 | E:\D554A Drive\7010_2_020420_rerun | PFAS UC | 7010_2_020420055.D | Sample |  | 2020-02-05 08:47 | 12000 | 11.4954 | 39330.4113 | 16850.5949 | F | P | P | 56508.1028 | 83592.9181 | 28769.0944 | 97298.8741 | 15063.1384

## 900 Cal
CC15 and CC13 Dropped for 900, MFOET used as IS

## 900_MDL_CC1
MDL Calculation
899
Sample | [5 empty] | 900 Method |  | 900 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_MDL_1 | 7010_2_020520000.D | Cal | 1 | 2020-02-05 12:56 | y = 4712.329312 * x | 1.1102e-16 | 8.9484 | 1.769 | 101.0834 | 1034.1381
| CC1_MDL_1 | 7010_2_020520001.D | Cal | 1 | 2020-02-05 13:17 | y = 4712.329312 * x | 1.1102e-16 | 8.9578 | 1.9843 | 113.3886 | 1085.5786
| CC1_MDL_1 | 7010_2_020520002.D | Cal | 1 | 2020-02-05 13:37 | y = 4712.329312 * x | 1.1102e-16 | 8.9578 | 1.8319 | 104.6822 | 981.4253
| CC1_MDL_1 | 7010_2_020520003.D | Cal | 1 | 2020-02-05 13:58 | y = 4712.329312 * x | 1.1102e-16 | 8.9626 | 1.8882 | 107.8968 | 924.146
| CC1_MDL_1 | 7010_2_020520004.D | Cal | 1 | 2020-02-05 14:18 | y = 4712.329312 * x | 1.1102e-16 | 8.9626 | 1.6984 | 97.0498 | 916.8857
| CC1_MDL_1 | 7010_2_020520005.D | Cal | 1 | 2020-02-05 14:39 | y = 4712.329312 * x | 1.1102e-16 | 8.9626 | 1.4735 | 84.199 | 827.4449
| CC1_MDL_1 | 7010_2_020520006.D | Cal | 1 | 2020-02-05 15:00 | y = 4712.329312 * x | 1.1102e-16 | 8.9626 | 1.6048 | 91.7002 | 795.7976
[8 empty] | sd | 0.17
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.53
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 899 Cal
CC8 Dropped for 899, MFBET used as IS

## 899_MDL_CC1
MDL Calculation
899
Sample | [5 empty] | 899 Method |  | 899 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_MDL_1 | 7010_2_020520000.D | Cal | 1 | 2020-02-05 12:56 | y = 161.316904 * x | 0 | 5.4599 | 1.5778 | 90.1609 | 129.5148
| CC1_MDL_1 | 7010_2_020520001.D | Cal | 1 | 2020-02-05 13:17 | y = 161.316904 * x | 0 | 5.4675 | 1.5397 | 87.9802 | 123.4866
| CC1_MDL_1 | 7010_2_020520002.D | Cal | 1 | 2020-02-05 13:37 | y = 161.316904 * x | 0 | 5.4704 | 2.3282 | 133.0412 | 181.7682
| CC1_MDL_1 | 7010_2_020520003.D | Cal | 1 | 2020-02-05 13:58 | y = 161.316904 * x | 0 | 5.4823 | 1.4753 | 84.3051 | 119.2935
| CC1_MDL_1 | 7010_2_020520004.D | Cal | 1 | 2020-02-05 14:18 | y = 161.316904 * x | 0 | 5.4704 | 2.3675 | 135.288 | 190.9427
| CC1_MDL_1 | 7010_2_020520005.D | Cal | 1 | 2020-02-05 14:39 | y = 161.316904 * x | 0 | 5.4675 | 1.452 | 82.9722 | 118.7497
| CC1_MDL_1 | 7010_2_020520006.D | Cal | 1 | 2020-02-05 15:00 | y = 161.316904 * x | 0 | 5.463 | 1.5094 | 86.2524 | 120.4894
[8 empty] | sd | 0.41
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.29
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## MDL Data
Sample | [6 empty] | 907 Method |  | 907 Results |  |  |  | MFBET (ISTD) Results |  | 915 Method |  | 915 Results |  |  |  | MFBET (ISTD) Results |  | 965 Method |  | 965 Results |  |  |  | MFHET (ISTD) Results |  | 476 Method |  | 476 Results |  |  |  | MFHET (ISTD) Results |  | 267 Method |  | 267 Results |  |  |  | MFOET (ISTD) Results |  | 3125 Method |  | 3125 Results |  |  |  | MFOET (ISTD) Results |  | 906 Method |  | 906 Results |  |  |  | MFOET (ISTD) Results |  | 273 Method |  | 273 Results |  |  |  | MFOET (ISTD) Results |  | 913 Method |  | 913 Results |  |  |  | MFOET (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 988150.074874 * x | 0 | 5.0263 | 1.7859 | 102.0501 | 3845228.387 | 5.5634 | 26147.4196 | y = 694.318369 * x | 0 | 5.5763 | 3.3224 | 189.852 | 5026.4286 | 5.5634 | 26147.4196 | y = 1521.247013 * x | 0 | 6.0906 | 2.4826 | 141.8644 | 6230.8186 | 6.2857 | 19797.7168 | y = 497.198143 * x | 0 | 6.9108 | 2.0271 | 115.8369 | 1662.8321 | 6.2857 | 19797.7168 | y = 501.004701 * x | 0 | 7.8147 | 1.4966 | 85.5199 | 707.711 | 7.1309 | 11326.354 | y = 89.256838 * x | -2.2204e-16 | 8.227 | 1.2506 | 71.462 | 105.357 | 7.1309 | 11326.354 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.7711 | 101.2037 | 11323.7657 | 7.1309 | 11326.354 | y = 483.247405 * x | -2.2204e-16 | 9.6529 | 2.3566 | 134.6642 | 1074.901 | 7.1309 | 11326.354 | y = 1414.900673 * x | 0 | 11.1147 | 0.074074 | 4.2328 | 98.9237 | 7.1309 | 11326.354 | y = 236.473244 * x | 0 | 11.2189 | 1.8147 | 103.6949 | 12971.6064 | 11.2181 | 362741.8499
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 988150.074874 * x | 0 | 5.2013 | 2.0713 | 118.3609 | 6659495.332 | 5.6567 | 39043.9022 | y = 694.318369 * x | 0 | 5.5803 | 1.089 | 62.2295 | 2460.1674 | 5.6567 | 39043.9022 | y = 1521.247013 * x | 0 | 6.1027 | 1.2913 | 73.7865 | 5121.0024 | 6.3069 | 31283.9418 | y = 497.198143 * x | 0 | 6.9143 | 1.7574 | 100.4215 | 2277.8986 | 6.3069 | 31283.9418 | y = 501.004701 * x | 0 | 7.8112 | 2.0644 | 117.9652 | 1257.2732 | 7.1275 | 14587.3778 | y = 89.256838 * x | -2.2204e-16 | 8.2296 | 2.2239 | 127.0781 | 241.2939 | 7.1275 | 14587.3778 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.8856 | 107.746 | 15526.8247 | 7.1275 | 14587.3778 | y = 483.247405 * x | -2.2204e-16 | 9.6606 | 1.2497 | 71.4112 | 734.1247 | 7.1275 | 14587.3778 | y = 1414.900673 * x | 0 | 11.1848 | 0.36164 | 20.6654 | 622.0186 | 7.1275 | 14587.3778 | y = 236.473244 * x | 0 | 11.219 | 1.7473 | 99.8444 | 15260.9596 | 11.2103 | 443219.8778
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 988150.074874 * x | 0 | 5.0341 | 1.4398 | 82.2757 | 3435373.546 | 5.5712 | 28974.9505 | y = 694.318369 * x | 0 | 5.5763 | 2.7253 | 155.7336 | 4568.9921 | 5.5712 | 28974.9505 | y = 1521.247013 * x | 0 | 6.0815 | 2.3148 | 132.2722 | 6840.6851 | 6.2857 | 23311.7215 | y = 497.198143 * x | 0 | 6.9108 | 1.5381 | 87.8905 | 1485.6022 | 6.2857 | 23311.7215 | y = 501.004701 * x | 0 | 7.8008 | 1.8958 | 108.3335 | 962.8596 | 7.1274 | 12164.7024 | y = 89.256838 * x | -2.2204e-16 | 8.2244 | 1.0195 | 58.2548 | 92.2426 | 7.1274 | 12164.7024 | y = 6774.033597 * x | -2.2204e-16 | 9.3664 | 1.6045 | 91.6853 | 11018.0672 | 7.1274 | 12164.7024 | y = 483.247405 * x | -2.2204e-16 | 9.658 | 2.5114 | 143.5097 | 1230.2941 | 7.1274 | 12164.7024 | y = 1414.900673 * x | 0 | 11.7918 | 0.34379 | 19.645 | 493.1025 | 7.1274 | 12164.7024 | y = 236.473244 * x | 0 | 11.2189 | 1.7029 | 97.3067 | 11936.1848 | 11.2181 | 355700.0325
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 988150.074874 * x | 0 | 5.178 | 2.053 | 117.3133 | 6353532.805 | 5.649 | 37582.7449 | y = 694.318369 * x | 0 | 5.5686 | 0.90549 | 51.7421 | 1969.0088 | 5.649 | 37582.7449 | y = 1521.247013 * x | 0 | 6.1027 | 1.2621 | 72.1204 | 5038.318 | 6.31 | 31489.8732 | y = 497.198143 * x | 0 | 6.9312 | 1.4921 | 85.263 | 1946.7844 | 6.31 | 31489.8732 | y = 501.004701 * x | 0 | 7.8004 | 1.5734 | 89.9061 | 1013.7104 | 7.127 | 15432.1432 | y = 89.256838 * x | -2.2204e-16 | 8.227 | 1.1643 | 66.5294 | 133.6406 | 7.127 | 15432.1432 | y = 6774.033597 * x | -2.2204e-16 | 9.3639 | 1.7067 | 97.5269 | 14868.0822 | 7.127 | 15432.1432 | y = 483.247405 * x | -2.2204e-16 | 9.6606 | 1.1223 | 64.1299 | 697.4504 | 7.127 | 15432.1432 | y = 1414.900673 * x | 0 | 11.6673 | 3.9001 | 222.865 | 7096.6091 | 7.127 | 15432.1432 | y = 236.473244 * x | 0 | 11.219 | 1.6588 | 94.786 | 15215.8079 | 11.2181 | 465491.652
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 988150.074874 * x | 0 | 5.2052 | 1.8054 | 103.1683 | 5745940.067 | 5.6567 | 38648.7208 | y = 694.318369 * x | 0 | 5.5803 | 0.88975 | 50.8427 | 1989.6624 | 5.6567 | 38648.7208 | y = 1521.247013 * x | 0 | 6.1057 | 1.3049 | 74.5667 | 5091.7414 | 6.3069 | 30779.7289 | y = 497.198143 * x | 0 | 6.9143 | 1.8511 | 105.7783 | 2360.7369 | 6.3069 | 30779.7289 | y = 501.004701 * x | 0 | 7.8043 | 1.6847 | 96.271 | 1134.6655 | 7.124 | 16131.4661 | y = 89.256838 * x | -2.2204e-16 | 8.2244 | 1.5402 | 88.0087 | 184.7983 | 7.124 | 16131.4661 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.6221 | 92.6903 | 14771.0869 | 7.124 | 16131.4661 | y = 483.247405 * x | -2.2204e-16 | 9.6605 | 1.4009 | 80.0501 | 910.0431 | 7.124 | 16131.4661 | y = 1414.900673 * x | 0 | 11.107 | 0.10164 | 5.8078 | 193.3149 | 7.124 | 16131.4661 | y = 236.473244 * x | 0 | 11.2189 | 1.7833 | 101.9025 | 16803.7841 | 11.2181 | 478171.4541
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 988150.074874 * x | 0 | 5.0457 | 0.81443 | 46.5387 | 2250133.696 | 5.5829 | 33551.6578 | y = 694.318369 * x | 0 | 5.5842 | 2.3072 | 131.8398 | 4478.9484 | 5.5829 | 33551.6578 | y = 1521.247013 * x | 0 | 6.0936 | 2.181 | 124.6275 | 6656.7635 | 6.2918 | 24076.4553 | y = 497.198143 * x | 0 | 6.9143 | 1.7592 | 100.5281 | 1754.9565 | 6.2918 | 24076.4553 | y = 501.004701 * x | 0 | 7.8008 | 1.5398 | 87.9904 | 773.0778 | 7.124 | 12025.1176 | y = 89.256838 * x | -2.2204e-16 | 8.2244 | 2.4908 | 142.3316 | 222.7866 | 7.124 | 12025.1176 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.7758 | 101.4766 | 12054.7812 | 7.124 | 12025.1176 | y = 483.247405 * x | -2.2204e-16 | 9.6606 | 2.2165 | 126.6552 | 1073.3435 | 7.124 | 12025.1176 | y = 1414.900673 * x | 0 | 10.6867 | 0.41186 | 23.5349 | 583.9622 | 7.124 | 12025.1176 | y = 236.473244 * x | 0 | 11.219 | 1.7068 | 97.534 | 12419.7493 | 11.2181 | 369247.9243
| UCGCAACC1_MDL | PFAS Group 2 UC | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 988150.074874 * x | 0 | 5.1818 | 2.2801 | 130.2929 | 6994613.453 | 5.649 | 37253.1591 | y = 694.318369 * x | 0 | 5.5686 | 1.0108 | 57.7603 | 2178.7547 | 5.649 | 37253.1591 | y = 1521.247013 * x | 0 | 6.0996 | 1.4133 | 80.7622 | 5160.4849 | 6.3069 | 28802.2184 | y = 497.198143 * x | 0 | 6.9177 | 1.8249 | 104.2817 | 2177.8121 | 6.3069 | 28802.2184 | y = 501.004701 * x | 0 | 7.8008 | 1.9952 | 114.0139 | 1251.9501 | 7.124 | 15029.0127 | y = 89.256838 * x | -2.2204e-16 | 8.227 | 2.5609 | 146.3353 | 286.2714 | 7.124 | 15029.0127 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.8842 | 107.6711 | 15985.7701 | 7.124 | 15029.0127 | y = 483.247405 * x | -2.2204e-16 | 9.6631 | 1.3926 | 79.5798 | 842.8678 | 7.124 | 15029.0127 | y = 1414.900673 * x | 0 | 11.7996 | 7.0569 | 403.2491 | 12505.0839 | 7.124 | 15029.0127 | y = 236.473244 * x | 0 | 11.2189 | 1.8363 | 104.9315 | 15748.6125 | 11.2181 | 435208.6169
[9 empty] | sd | 0.492 | [6 empty] | sd | 1.014 | [6 empty] | sd | 0.548 | [6 empty] | sd | 0.184 | [6 empty] | sd | 0.232 | [6 empty] | sd | 0.658 | [6 empty] | sd | 0.113 | [6 empty] | sd | 0.586 | [6 empty] | sd | 2.708 | [6 empty] | sd | 0.065
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.55 | [6 empty] | MDL (nM) | 3.19 | [6 empty] | MDL (nM) | 1.72 | [6 empty] | MDL (nM) | 0.58 | [6 empty] | MDL (nM) | 0.73 | [6 empty] | MDL (nM) | 2.07 | [6 empty] | MDL (nM) | 0.36 | [6 empty] | MDL (nM) | 1.84 | [6 empty] | MDL (nM) | 8.51 | [6 empty] | MDL (nM) | 0.2
Sample | [6 empty] | 907 Method |  | 907 Results |  |  |  | MFBET (ISTD) Results |  | 915 Method |  | 915 Results |  |  |  | MFBET (ISTD) Results |  | 965 Method |  | 965 Results |  |  |  | MFHET (ISTD) Results |  | 476 Method |  | 476 Results |  |  |  | MFHET (ISTD) Results |  | 267 Method |  | 267 Results |  |  |  | MFOET (ISTD) Results |  | 3125 Method |  | 3125 Results |  |  |  | MFOET (ISTD) Results |  | 906 Method |  | 906 Results |  |  |  | MFOET (ISTD) Results |  | 273 Method |  | 273 Results |  |  |  | MFOET (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results |  | 913 Method |  | 913 Results |  |  |  | MFOET (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042088.D | Cal | 4 | 2020-02-05 23:04 | y = 211762.969083 * x | 0 | 5.2674 | 6.8789 | 91.719 | 4526056.163 | 5.6723 | 37284.6909 | y = 516.242687 * x | 0 | 5.6036 | 8.0254 | 107.0055 | 12872.7301 | 5.6723 | 37284.6909 | y = 1254.136319 * x | 1.1102e-16 | 6.1118 | 6.8639 | 91.5185 | 20475.5033 | 6.31 | 28543.092 | y = 518.960488 * x | -2.2204e-16 | 6.9143 | 7.1998 | 95.9978 | 8887.4418 | 6.31 | 28543.092 | y = 509.437850 * x | 0 | 7.8043 | 8.5209 | 113.6115 | 5438.1285 | 7.124 | 15033.3532 | y = 32.587088 * x | 0 | 8.2295 | 6.752 | 90.0268 | 275.6471 | 7.124 | 15033.3532 | y = 8772.072887 * x | 0 | 9.3613 | 7.0524 | 94.0317 | 77501.8903 | 7.124 | 15033.3532 | y = 462.277108 * x | 0 | 9.6606 | 6.6416 | 88.5551 | 3846.3778 | 7.124 | 15033.3532 | y = 166.327112 * x | 0 | 11.2112 | 7.2801 | 97.0674 | 49530.5011 | 11.2103 | 490858.2797 | y = 4059.078801 * x | 0 | 11.9864 | 7.3922 | 98.5624 | 37590.2047 | 7.124 | 15033.3532
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042089.D | Cal | 4 | 2020-02-05 23:26 | y = 211762.969083 * x | 0 | 5.2246 | 6.3617 | 84.8233 | 4164686.468 | 5.6606 | 37096.8575 | y = 516.242687 * x | 0 | 5.5802 | 7.4937 | 99.9165 | 11959.3643 | 5.6606 | 37096.8575 | y = 1254.136319 * x | 1.1102e-16 | 6.1087 | 6.6057 | 88.0756 | 18955.7915 | 6.316 | 27457.5212 | y = 518.960488 * x | -2.2204e-16 | 6.9142 | 6.8226 | 90.9678 | 8101.4572 | 6.316 | 27457.5212 | y = 509.437850 * x | 0 | 7.8077 | 7.7315 | 103.0865 | 4350.1829 | 7.1309 | 13253.6199 | y = 32.587088 * x | 0 | 8.227 | 3.0479 | 40.639 | 109.6992 | 7.1309 | 13253.6199 | y = 8772.072887 * x | 0 | 9.3638 | 7.8754 | 105.0051 | 76300.4907 | 7.1309 | 13253.6199 | y = 462.277108 * x | 0 | 9.6605 | 7.0715 | 94.286 | 3610.4736 | 7.1309 | 13253.6199 | y = 166.327112 * x | 0 | 11.2189 | 7.4736 | 99.6478 | 46222.2082 | 11.2103 | 446210.4675 | y = 4059.078801 * x | 0 | 11.8385 | 7.8431 | 104.5746 | 35161.5789 | 7.1309 | 13253.6199
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042090.D | Cal | 4 | 2020-02-05 23:47 | y = 211762.969083 * x | 0 | 5.2674 | 4.8904 | 65.205 | 2900184.798 | 5.6723 | 33605.8457 | y = 516.242687 * x | 0 | 5.6036 | 7.4276 | 99.0344 | 10738.2806 | 5.6723 | 33605.8457 | y = 1254.136319 * x | 1.1102e-16 | 6.1148 | 6.8359 | 91.1454 | 19725.8111 | 6.3099 | 27610.566 | y = 518.960488 * x | -2.2204e-16 | 6.9143 | 7.0628 | 94.1705 | 8433.4327 | 6.3099 | 27610.566 | y = 509.437850 * x | 0 | 7.8043 | 7.5975 | 101.2994 | 4589.5149 | 7.1275 | 14229.4662 | y = 32.587088 * x | 0 | 8.2296 | 9.6213 | 128.2837 | 371.7798 | 7.1275 | 14229.4662 | y = 8772.072887 * x | 0 | 9.3613 | 7.3053 | 97.4035 | 75988.0329 | 7.1275 | 14229.4662 | y = 462.277108 * x | 0 | 9.658 | 6.9434 | 92.5786 | 3806.114 | 7.1275 | 14229.4662 | y = 166.327112 * x | 0 | 11.219 | 7.3588 | 98.1174 | 50243.1191 | 11.2103 | 492592.0107 | y = 4059.078801 * x | 0 | 11.823 | 7.1764 | 95.6852 | 34541.4803 | 7.1275 | 14229.4662
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042091.D | Cal | 4 | 2020-02-06 00:08 | y = 211762.969083 * x | 0 | 5.2674 | 16.7685 | 223.5797 | 11093457.06 | 5.6723 | 37489.0622 | y = 516.242687 * x | 0 | 5.6036 | 8.0304 | 107.0722 | 12951.353 | 5.6723 | 37489.0622 | y = 1254.136319 * x | 1.1102e-16 | 6.1148 | 7.2643 | 96.8573 | 19562.9476 | 6.3099 | 25767.7889 | y = 518.960488 * x | -2.2204e-16 | 6.9142 | 7.7185 | 102.913 | 8601.2526 | 6.3099 | 25767.7889 | y = 509.437850 * x | 0 | 7.8008 | 8.0483 | 107.311 | 4164.7043 | 7.124 | 12189.0189 | y = 32.587088 * x | 0 | 8.2218 | 6.2923 | 83.8977 | 208.2784 | 7.124 | 12189.0189 | y = 8772.072887 * x | 0 | 9.3613 | 8.4437 | 112.5827 | 75235.4857 | 7.124 | 12189.0189 | y = 462.277108 * x | 0 | 9.658 | 8.2486 | 109.9815 | 3873.2084 | 7.124 | 12189.0189 | y = 166.327112 * x | 0 | 11.2189 | 7.5174 | 100.2314 | 49258.7629 | 11.2103 | 472755.5882 | y = 4059.078801 * x | 0 | 11.7529 | 8.8571 | 118.095 | 36518.0519 | 7.124 | 12189.0189
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042092.D | Cal | 4 | 2020-02-06 00:29 | y = 211762.969083 * x | 0 | 5.2052 | 8.0371 | 107.1615 | 4610337.278 | 5.6529 | 32506.0374 | y = 516.242687 * x | 0 | 5.5725 | 8.0078 | 106.7709 | 11198.2641 | 5.6529 | 32506.0374 | y = 1254.136319 * x | 1.1102e-16 | 6.1027 | 7.9299 | 105.7326 | 19116.6251 | 6.313 | 23066.278 | y = 518.960488 * x | -2.2204e-16 | 6.9143 | 8.4099 | 112.1316 | 8389.1883 | 6.313 | 23066.278 | y = 509.437850 * x | 0 | 7.8043 | 7.0006 | 93.3407 | 4201.0451 | 7.1309 | 14135.6228 | y = 32.587088 * x | 0 | 8.2295 | 8.336 | 111.1465 | 319.99 | 7.1309 | 14135.6228 | y = 8772.072887 * x | 0 | 9.3638 | 6.816 | 90.8804 | 70431.574 | 7.1309 | 14135.6228 | y = 462.277108 * x | 0 | 9.6606 | 6.7506 | 90.0085 | 3676.0462 | 7.1309 | 14135.6228 | y = 166.327112 * x | 0 | 11.219 | 7.7672 | 103.5625 | 46959.2409 | 11.2181 | 436189.8583 | y = 4059.078801 * x | 0 | 11.7685 | 7.3068 | 97.4235 | 34937.0479 | 7.1309 | 14135.6228
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042093.D | Cal | 4 | 2020-02-06 00:51 | y = 211762.969083 * x | 0 | 5.0535 | 5.6359 | 75.1447 | 2751029.092 | 5.5828 | 27660.913 | y = 516.242687 * x | 0 | 5.4908 | 6.3855 | 85.1403 | 7598.6329 | 5.5828 | 27660.913 | y = 1254.136319 * x | 1.1102e-16 | 6.0814 | 6.7561 | 90.0811 | 17751.4984 | 6.2917 | 25140.6434 | y = 518.960488 * x | -2.2204e-16 | 6.9177 | 6.0583 | 80.7769 | 6586.8556 | 6.2917 | 25140.6434 | y = 509.437850 * x | 0 | 7.8043 | 7.1764 | 95.6853 | 3586.7777 | 7.1309 | 11773.0186 | y = 32.587088 * x | 0 | 8.2218 | 11.7755 | 157.0072 | 376.4722 | 7.1309 | 11773.0186 | y = 8772.072887 * x | 0 | 9.3664 | 7.1811 | 95.7484 | 61801.8507 | 7.1309 | 11773.0186 | y = 462.277108 * x | 0 | 9.6605 | 7.9627 | 106.1691 | 3611.3397 | 7.1309 | 11773.0186 | y = 166.327112 * x | 0 | 11.2189 | 7.9256 | 105.6744 | 42775.2006 | 11.2181 | 389385.1031 | y = 4059.078801 * x | 0 | 11.6906 | 7.1223 | 94.9635 | 28363.0026 | 7.1309 | 11773.0186
| UCGCAACC4_MDL | PFAS Group 2 UC | 7010_1_02042094.D | Cal | 4 | 2020-02-06 01:12 | y = 211762.969083 * x | 0 | 5.0496 | 3.9275 | 52.3668 | 1702762.055 | 5.5945 | 24567.8769 | y = 516.242687 * x | 0 | 5.4986 | 7.1295 | 95.0603 | 7535.3056 | 5.5945 | 24567.8769 | y = 1254.136319 * x | 1.1102e-16 | 6.0906 | 10.2442 | 136.5894 | 18621.0934 | 6.3039 | 17392.5545 | y = 518.960488 * x | -2.2204e-16 | 6.9247 | 9.2282 | 123.0424 | 6941.1643 | 6.3039 | 17392.5545 | y = 509.437850 * x | 0 | 7.8043 | 6.4249 | 85.6655 | 2982.4089 | 7.1309 | 10934.274 | y = 32.587088 * x | 0 | 8.2244 | 6.6749 | 88.999 | 198.1987 | 7.1309 | 10934.274 | y = 8772.072887 * x | 0 | 9.3639 | 7.8261 | 104.3482 | 62554.3251 | 7.1309 | 10934.274 | y = 462.277108 * x | 0 | 9.6606 | 8.8816 | 118.4211 | 3741.1184 | 7.1309 | 10934.274 | y = 166.327112 * x | 0 | 11.219 | 7.1774 | 95.6991 | 37125.8324 | 11.2181 | 373186.093 | y = 4059.078801 * x | 0 | 11.7607 | 6.8022 | 90.6957 | 25158.4758 | 7.1309 | 10934.274

## 915_Cal
IS = MFBET

## 915 MDL
MDL Calculation
915 (CC4 Used)
Sample | [5 empty] | 915 Method |  | 915 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC4_MDL | 7010_1_02042088.D | Cal | 4 | 2020-02-05 23:04 | y = 516.242687 * x | 0 | 5.6036 | 8.0254 | 107.0055 | 12872.7301
| UCGCAACC4_MDL | 7010_1_02042089.D | Cal | 4 | 2020-02-05 23:26 | y = 516.242687 * x | 0 | 5.5802 | 7.4937 | 99.9165 | 11959.3643
| UCGCAACC4_MDL | 7010_1_02042090.D | Cal | 4 | 2020-02-05 23:47 | y = 516.242687 * x | 0 | 5.6036 | 7.4276 | 99.0344 | 10738.2806
| UCGCAACC4_MDL | 7010_1_02042091.D | Cal | 4 | 2020-02-06 00:08 | y = 516.242687 * x | 0 | 5.6036 | 8.0304 | 107.0722 | 12951.353
| UCGCAACC4_MDL | 7010_1_02042092.D | Cal | 4 | 2020-02-06 00:29 | y = 516.242687 * x | 0 | 5.5725 | 8.0078 | 106.7709 | 11198.2641
| UCGCAACC4_MDL | 7010_1_02042093.D | Cal | 4 | 2020-02-06 00:51 | y = 516.242687 * x | 0 | 5.4908 | 6.3855 | 85.1403 | 7598.6329
| UCGCAACC4_MDL | 7010_1_02042094.D | Cal | 4 | 2020-02-06 01:12 | y = 516.242687 * x | 0 | 5.4986 | 7.1295 | 95.0603 | 7535.3056
[8 empty] | sd | 0.61
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.92
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 965 Cal
IS = MFHET

## 965 MDL
MDL Calculation
965 (CC1 Used)
Sample | [5 empty] | 965 Method |  | 965 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_MDL | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 1521.247013 * x | 0 | 6.0906 | 2.4826 | 141.8644 | 6230.8186
| UCGCAACC1_MDL | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 1521.247013 * x | 0 | 6.1027 | 1.2913 | 73.7865 | 5121.0024
| UCGCAACC1_MDL | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 1521.247013 * x | 0 | 6.0815 | 2.3148 | 132.2722 | 6840.6851
| UCGCAACC1_MDL | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 1521.247013 * x | 0 | 6.1027 | 1.2621 | 72.1204 | 5038.318
| UCGCAACC1_MDL | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 1521.247013 * x | 0 | 6.1057 | 1.3049 | 74.5667 | 5091.7414
| UCGCAACC1_MDL | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 1521.247013 * x | 0 | 6.0936 | 2.181 | 124.6275 | 6656.7635
| UCGCAACC1_MDL | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 1521.247013 * x | 0 | 6.0996 | 1.4133 | 80.7622 | 5160.4849
[8 empty] | sd | 0.55
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.73
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 476 Cal
IS = MFHET

## 476 MDL
MDL Calculation
476 (CC1 Used)
Sample | [5 empty] | 476 Method |  | 476 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_MDL | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 497.198143 * x | 0 | 6.9108 | 2.0271 | 115.8369 | 1662.8321
| UCGCAACC1_MDL | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 497.198143 * x | 0 | 6.9143 | 1.7574 | 100.4215 | 2277.8986
| UCGCAACC1_MDL | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 497.198143 * x | 0 | 6.9108 | 1.5381 | 87.8905 | 1485.6022
| UCGCAACC1_MDL | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 497.198143 * x | 0 | 6.9312 | 1.4921 | 85.263 | 1946.7844
| UCGCAACC1_MDL | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 497.198143 * x | 0 | 6.9143 | 1.8511 | 105.7783 | 2360.7369
| UCGCAACC1_MDL | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 497.198143 * x | 0 | 6.9143 | 1.7592 | 100.5281 | 1754.9565
| UCGCAACC1_MDL | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 497.198143 * x | 0 | 6.9177 | 1.8249 | 104.2817 | 2177.8121
[8 empty] | sd | 0.18
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.57
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 267 Cal
IS = MFOET

## 267 MDL
MDL Calculation
267 (CC1 Used)
Sample | [5 empty] | 267 Method |  | 267 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_MDL | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 501.004701 * x | 0 | 7.8147 | 1.4966 | 85.5199 | 707.711
| UCGCAACC1_MDL | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 501.004701 * x | 0 | 7.8112 | 2.0644 | 117.9652 | 1257.2732
| UCGCAACC1_MDL | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 501.004701 * x | 0 | 7.8008 | 1.8958 | 108.3335 | 962.8596
| UCGCAACC1_MDL | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 501.004701 * x | 0 | 7.8004 | 1.5734 | 89.9061 | 1013.7104
| UCGCAACC1_MDL | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 501.004701 * x | 0 | 7.8043 | 1.6847 | 96.271 | 1134.6655
| UCGCAACC1_MDL | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 501.004701 * x | 0 | 7.8008 | 1.5398 | 87.9904 | 773.0778
| UCGCAACC1_MDL | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 501.004701 * x | 0 | 7.8008 | 1.9952 | 114.0139 | 1251.9501
[8 empty] | sd | 0.23
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.72
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 906 Cal
IS = MFOET

## 906 MDL
MDL Calculation
906 (CC1 Used)
Sample | [5 empty] | 906 Method |  | 906 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_MDL | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.7711 | 101.2037 | 11323.7657
| UCGCAACC1_MDL | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.8856 | 107.746 | 15526.8247
| UCGCAACC1_MDL | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 6774.033597 * x | -2.2204e-16 | 9.3664 | 1.6045 | 91.6853 | 11018.0672
| UCGCAACC1_MDL | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 6774.033597 * x | -2.2204e-16 | 9.3639 | 1.7067 | 97.5269 | 14868.0822
| UCGCAACC1_MDL | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.6221 | 92.6903 | 14771.0869
| UCGCAACC1_MDL | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.7758 | 101.4766 | 12054.7812
| UCGCAACC1_MDL | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 6774.033597 * x | -2.2204e-16 | 9.3638 | 1.8842 | 107.6711 | 15985.7701
[8 empty] | sd | 0.11
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.35
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 273 Cal
IS = MFOET

## 273 MDL
MDL Calculation
273 (CC4 Used)
Sample | [5 empty] | 273 Method |  | 273 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC4_MDL | 7010_1_02042088.D | Cal | 4 | 2020-02-05 23:04 | y = 462.277108 * x | 0 | 9.6606 | 6.6416 | 88.5551 | 3846.3778
| UCGCAACC4_MDL | 7010_1_02042089.D | Cal | 4 | 2020-02-05 23:26 | y = 462.277108 * x | 0 | 9.6605 | 7.0715 | 94.286 | 3610.4736
| UCGCAACC4_MDL | 7010_1_02042090.D | Cal | 4 | 2020-02-05 23:47 | y = 462.277108 * x | 0 | 9.658 | 6.9434 | 92.5786 | 3806.114
| UCGCAACC4_MDL | 7010_1_02042091.D | Cal | 4 | 2020-02-06 00:08 | y = 462.277108 * x | 0 | 9.658 | 8.2486 | 109.9815 | 3873.2084
| UCGCAACC4_MDL | 7010_1_02042092.D | Cal | 4 | 2020-02-06 00:29 | y = 462.277108 * x | 0 | 9.6606 | 6.7506 | 90.0085 | 3676.0462
| UCGCAACC4_MDL | 7010_1_02042093.D | Cal | 4 | 2020-02-06 00:51 | y = 462.277108 * x | 0 | 9.6605 | 7.9627 | 106.1691 | 3611.3397
| UCGCAACC4_MDL | 7010_1_02042094.D | Cal | 4 | 2020-02-06 01:12 | y = 462.277108 * x | 0 | 9.6606 | 8.8816 | 118.4211 | 3741.1184
[8 empty] | sd | 0.86
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 2.7
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 913 Cal
IS = MFOET

## 913 MDL
MDL Calculation
913 (CC4 Used)
Sample | [5 empty] | 913 Method |  | 913 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC4_MDL | 7010_1_02042088.D | Cal | 4 | 2020-02-05 23:04 | y = 4059.078801 * x | 0 | 11.9864 | 7.3922 | 98.5624 | 37590.2047
| UCGCAACC4_MDL | 7010_1_02042089.D | Cal | 4 | 2020-02-05 23:26 | y = 4059.078801 * x | 0 | 11.8385 | 7.8431 | 104.5746 | 35161.5789
| UCGCAACC4_MDL | 7010_1_02042090.D | Cal | 4 | 2020-02-05 23:47 | y = 4059.078801 * x | 0 | 11.823 | 7.1764 | 95.6852 | 34541.4803
| UCGCAACC4_MDL | 7010_1_02042091.D | Cal | 4 | 2020-02-06 00:08 | y = 4059.078801 * x | 0 | 11.7529 | 8.8571 | 118.095 | 36518.0519
| UCGCAACC4_MDL | 7010_1_02042092.D | Cal | 4 | 2020-02-06 00:29 | y = 4059.078801 * x | 0 | 11.7685 | 7.3068 | 97.4235 | 34937.0479
| UCGCAACC4_MDL | 7010_1_02042093.D | Cal | 4 | 2020-02-06 00:51 | y = 4059.078801 * x | 0 | 11.6906 | 7.1223 | 94.9635 | 28363.0026
| UCGCAACC4_MDL | 7010_1_02042094.D | Cal | 4 | 2020-02-06 01:12 | y = 4059.078801 * x | 0 | 11.7607 | 6.8022 | 90.6957 | 25158.4758
[8 empty] | sd | 0.68
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 2.14
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## Cal_Curve 3125
All points used, Quadratic

## 3125_MDL
MDL Calculation
3125
Sample | [5 empty] | 900 Method |  | 900 Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032060.D | Cal | 1 | 2020-03-04 06:35 | y = 1729.487335 * x | 0 | 10.3573 | 1.9018 | 108.6746 | 399.2569
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032077.D | Cal | 1 | 2020-03-04 12:25 | y = 1729.487335 * x | 0 | 10.3572 | 1.8354 | 104.8808 | 364.6062
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032078.D | Cal | 1 | 2020-03-04 12:45 | y = 1729.487335 * x | 0 | 10.3573 | 1.4127 | 80.7277 | 299.8208
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032079.D | Cal | 1 | 2020-03-04 13:06 | y = 1729.487335 * x | 0 | 10.3572 | 1.9095 | 109.1167 | 384.4152
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032080.D | Cal | 1 | 2020-03-04 13:27 | y = 1729.487335 * x | 0 | 10.3657 | 1.7957 | 102.6095 | 343.5416
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032081.D | Cal | 1 | 2020-03-04 13:47 | y = 1729.487335 * x | 0 | 10.3572 | 1.6813 | 96.0771 | 344.9902
| UCGCAACC1_1/29/20_GC 2/3 | 7010_2_03032082.D | Cal | 1 | 2020-03-04 14:08 | y = 1729.487335 * x | 0 | 10.3573 | 1.7135 | 97.9135 | 353.9471
[8 empty] | sd | 0.17
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.53
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 4NT Cal
IS =4NT13C6

## 4NT MDL
MDL Calculation
4NT (CC1 Used)
Sample | [5 empty] | 4NT Method |  | 4NT Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| UCGCAACC1_MDL | 7010_1_02042087.D | Cal | 1 | 2020-02-05 22:43 | y = 236.473244 * x | 0 | 11.2189 | 1.8147 | 103.6949 | 12971.6064
| UCGCAACC1_MDL | 7010_1_02042086.D | Cal | 1 | 2020-02-05 22:22 | y = 236.473244 * x | 0 | 11.219 | 1.7473 | 99.8444 | 15260.9596
| UCGCAACC1_MDL | 7010_1_02042085.D | Cal | 1 | 2020-02-05 22:00 | y = 236.473244 * x | 0 | 11.2189 | 1.7029 | 97.3067 | 11936.1848
| UCGCAACC1_MDL | 7010_1_02042084.D | Cal | 1 | 2020-02-05 21:39 | y = 236.473244 * x | 0 | 11.219 | 1.6588 | 94.786 | 15215.8079
| UCGCAACC1_MDL | 7010_1_02042083.D | Cal | 1 | 2020-02-05 21:18 | y = 236.473244 * x | 0 | 11.2189 | 1.7833 | 101.9025 | 16803.7841
| UCGCAACC1_MDL | 7010_1_02042082.D | Cal | 1 | 2020-02-05 20:56 | y = 236.473244 * x | 0 | 11.219 | 1.7068 | 97.534 | 12419.7493
| UCGCAACC1_MDL | 7010_1_02042081.D | Cal | 1 | 2020-02-05 20:35 | y = 236.473244 * x | 0 | 11.2189 | 1.8363 | 104.9315 | 15748.6125
[8 empty] | sd | 0.06
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.19
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## 3125Cal Curve_4NT
All points used, Linear

## 4NT Cal_899&900
All points used for 4NT, 4NT13C6 used as IS

## 4NT_MDL_CC1_899&900
MDL Calculation
4NT
Sample | [5 empty] | 4NT Method |  | 4NT Results
| Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_MDL_1 | 7010_2_020520000.D | Cal | 1 | 2020-02-05 12:56 | y = 226.065662 * x | 0 | 11.4848 | 1.4999 | 85.7111 | 1699.288
| CC1_MDL_1 | 7010_2_020520001.D | Cal | 1 | 2020-02-05 13:17 | y = 226.065662 * x | 0 | 11.4906 | 1.4301 | 81.7178 | 1540.6242
| CC1_MDL_1 | 7010_2_020520002.D | Cal | 1 | 2020-02-05 13:37 | y = 226.065662 * x | 0 | 11.4967 | 2.2096 | 126.2612 | 2330.2362
| CC1_MDL_1 | 7010_2_020520003.D | Cal | 1 | 2020-02-05 13:58 | y = 226.065662 * x | 0 | 11.4906 | 1.6354 | 93.4537 | 1604.6204
| CC1_MDL_1 | 7010_2_020520004.D | Cal | 1 | 2020-02-05 14:18 | y = 226.065662 * x | 0 | 11.4967 | 1.5918 | 90.9591 | 1713.1025
| CC1_MDL_1 | 7010_2_020520005.D | Cal | 1 | 2020-02-05 14:39 | y = 226.065662 * x | 0 | 11.4966 | 1.6527 | 94.4426 | 1743.6507
| CC1_MDL_1 | 7010_2_020520006.D | Cal | 1 | 2020-02-05 15:00 | y = 226.065662 * x | 0 | 11.4907 | 2.2305 | 127.4546 | 2241.1202
[8 empty] | sd | 0.33
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  | 3.143 |  |  |  |  | MDL (nM) | 1.04
References:
40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [7 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf

## Analysis
Sample | [7 empty] | 907 Method |  | 907 Results | [10 empty] | MFBET (ISTD) Results |  |  | 3125 Method |  | 3125 Results |  |  |  | MFOET (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj |  |  |  |  | Area | RT | Area |  | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
|  | UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1585 | 0 |  |  | 752.9841 | 3011.9364 |  |  |  |  | 4656747.6265 | 5.6451 | 44276.266 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2322 | 0 |  | 1004.3118 | 7.1201 | 15322.4289 | y = 153.621749 * x | 0.99954 | 11.219 | 55.6284 |  | 324929.2471 | 11.2181 | 456268.9968
|  | UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.0496 | 0 |  |  | 722.2826 | 2889.1305 |  |  |  |  | 3576428.7517 | 5.5712 | 27629.305 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2347 | 0 |  | 1452.8442 | 7.117 | 12225.8416 | y = 153.621749 * x | 0.99954 | 11.2189 | 53.1138 |  | 272083.1737 | 11.2181 | 400149.7862
|  | UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.038 | 0 | 378.7819 |  | 797.4247 | 3189.6988 |  |  |  |  | 3543110.4513 | 5.5867 | 34358.3476 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2373 | 0 | 0 | 0 | 7.124 | 13553.8369 | y = 153.621749 * x | 0.99954 | 11.219 | 54.2685 |  | 259199.7576 | 11.2181 | 373091.7078
|  | UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1818 | 516.6606 |  |  | 1764.1073 | 7056.429 |  |  |  |  | 9328606.7185 | 5.6606 | 42279.1923 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2218 | 122.0789 |  | 40349.7739 | 7.1032 | 16387.1308 | y = 153.621749 * x | 0.99954 | 11.2189 | 236.1243 |  | 1559477.053 | 11.2181 | 515901.6354
|  | UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1663 | 0 |  |  | 1757.6622 | 7030.6489 |  |  |  |  | 7387053.6599 | 5.6529 | 45252.3426 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 120.5983 |  | 38338.1948 | 7.1132 | 15740.8529 | y = 153.621749 * x | 0.99954 | 11.219 | 239.3649 |  | 1494748.7556 | 11.2181 | 487793.8886
|  | UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2246 | 0 | 186.6874 |  | 2078.8533 | 8315.4133 |  |  |  |  | 7878227.3542 | 5.6684 | 48899.4333 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 28.9607 |  | 13010.678 | 7.117 | 16613.0802 | y = 153.621749 * x | 0.99954 | 11.2189 | 229.1937 |  | 1583252.3854 | 11.2181 | 539605.0373
|  | UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1352 | 0 |  |  | 1687.8795 | 6751.5179 |  |  |  |  | 2200812.7226 | 5.6256 | 49525.8927 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 2.3472 |  | 5725.8095 | 7.1205 | 18880.9947 | y = 153.621749 * x | 0.99954 | 11.2189 | 200.2697 |  | 1446510.0873 | 11.2181 | 564202.1179
|  | UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2286 | 0 |  |  | 1828.5526 | 7314.2106 |  |  |  |  | 6311525.0992 | 5.6684 | 45385.4767 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 13.6111 |  | 8665.2176 | 7.1101 | 17112.4072 | y = 153.621749 * x | 0.99954 | 11.219 | 197.845 |  | 1199756.8571 | 11.2181 | 473692.7868
|  | UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1196 | 0 | 169.3189 |  | 1563.1346 | 6252.5384 |  |  |  |  | 1466209.919 | 5.6256 | 44371.295 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 18.8334 |  | 9487.2043 | 7.117 | 15797.7788 | y = 153.621749 * x | 0.99954 | 11.2189 | 192.3357 |  | 1079328.812 | 11.2181 | 438351.4479
|  | UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2168 | 0 | [8 empty] | 5012794.55 | 5.6645 | 43521.4244 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2372 | 0 | 0 | 0 | 7.117 | 17384.4001 | y = 153.621749 * x | 0.99954 | 11.2189 | 62.9781 |  | 406758.0232 | 11.2181 | 504516.0123
|  | UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2246 | 0 | [8 empty] | 9015447.5624 | 5.6723 | 47307.1562 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 | 0 | 0 | 7.1166 | 17563.247 | y = 153.621749 * x | 0.99954 | 11.219 | 53.4542 |  | 349758.6536 | 11.2181 | 511111.1467
|  | UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.0613 | 0 | [8 empty] | 4050244.2467 | 5.5751 | 35262.4856 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.245 | 0 | 0 | 0 | 7.1205 | 13754.1616 | y = 153.621749 * x | 0.99954 | 11.2189 | 58.7453 |  | 307408.5122 | 11.2181 | 408763.0851
|  | UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1935 | 0 | [8 empty] | 7785594.8256 | 5.6567 | 45429.6893 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 0 | 7.1205 | 15082.165 | y = 153.621749 * x | 0.99954 | 11.2189 | 228.357 |  | 1268681.9773 | 11.2181 | 433977.2416
|  | UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2013 | 195.9941 | [8 empty] | 8831162.8367 | 5.6645 | 42273.4175 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 476.9612 | 7.1066 | 16448.5129 | y = 153.621749 * x | 0.99954 | 11.2189 | 216.1466 |  | 1341527.1913 | 11.2181 | 484818.9148
|  | UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1858 | 0 | [8 empty] | 6660197.6794 | 5.6451 | 40772.8976 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 280.6394 | 7.1132 | 16249.7142 | y = 153.621749 * x | 0.99954 | 11.219 | 211.3997 |  | 1347459.2993 | 11.2181 | 497897.4822
|  | UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1663 | 327.1044 | [8 empty] | 10032684.8646 | 5.6528 | 46946.472 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.227 | 0 |  | 159.9197 | 7.1136 | 18071.5743 | y = 153.621749 * x | 0.99954 | 11.2189 | 182.8837 |  | 1335265.836 | 11.2181 | 570323.466
|  | UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1663 | 0 | [8 empty] | 8864967.8544 | 5.6529 | 44216.1553 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 288.742 | 7.1171 | 17483.3019 | y = 153.621749 * x | 0.99954 | 11.219 | 188.2681 |  | 1211653.6671 | 11.2181 | 502724.7548
|  | UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2752 | 0 | [8 empty] | 9269022.7122 | 5.684 | 48159.5878 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 232.3474 | 7.1132 | 17572.9112 | y = 153.621749 * x | 0.99954 | 11.219 | 166.8947 |  | 1046225.116 | 11.2181 | 489678.7122
|  | UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1702 | 0 | [8 empty] | 7905384.7634 | 5.6529 | 43125.2652 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2373 | 0 |  | 397.1058 | 7.1201 | 17231.3883 | y = 153.621749 * x | 0.99954 | 11.219 | 50.3812 |  | 327375.3083 | 11.2181 | 507581.8712
|  | UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2441 | 2104.6568 | [8 empty] | 12782881.4178 | 5.6723 | 45855.005 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 |  | 320.1417 | 7.117 | 16295.4315 | y = 153.621749 * x | 0.99954 | 11.2189 | 63.4059 |  | 403640.3802 | 11.2181 | 497271.4447
|  | UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2441 | 0 | [8 empty] | 4818235.5452 | 5.6723 | 48219.7539 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2347 | 0 |  | 426.6992 | 7.1171 | 17456.357 | y = 153.621749 * x | 0.99954 | 11.219 | 52.1492 |  | 376158.2799 | 11.2181 | 563445.3712
|  | UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2363 | 0 | [8 empty] | 6192370.4485 | 5.6684 | 38389.9729 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 |  | 380.6254 | 7.1136 | 15375.4789 | y = 153.621749 * x | 0.99954 | 11.2189 | 227.5675 |  | 1251495.903 | 11.2181 | 429583.7135
|  | UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2285 | 0 | [8 empty] | 8925258.0522 | 5.6762 | 46610.1305 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2295 | 0 |  | 306.5811 | 7.1067 | 16474.9308 | y = 153.621749 * x | 0.99954 | 11.2189 | 225.7573 |  | 1438599.7058 | 11.2181 | 497767.6323
|  | UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1002 | 159.7704 | [8 empty] | 9680428.8318 | 5.6295 | 46634.702 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2321 | 0 | 0 | 0 | 7.1205 | 18208.5429 | y = 153.621749 * x | 0.99954 | 11.219 | 211.6066 |  | 1420407.5714 | 11.2181 | 524339.2574
|  | UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.2518 | 2357.6608 | [8 empty] | 13007030.8878 | 5.6762 | 45158.9158 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2398 | 0 | 0 | 0 | 7.1136 | 15982.4355 | y = 153.621749 * x | 0.99954 | 11.2189 | 187.6978 |  | 1119892.7533 | 11.2181 | 466064.303
|  | UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1547 | 0 | [8 empty] | 6588931.0593 | 5.649 | 44996.0952 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 311.0805 | 7.1097 | 16177.7404 | y = 153.621749 * x | 0.99954 | 11.219 | 182.2168 |  | 1154335.3982 | 11.2181 | 494848.2367
|  | UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = 439.228605 * x + 201.731965 | 0.036747 | 5.1858 | 0 | [8 empty] | 6698905.0428 | 5.6529 | 43830.4305 |  | y = 216.386359 * x + 0.260933 | 0.98762 | 8.2296 | 0 |  | 177.5834 | 7.1171 | 16527.163 | y = 153.621749 * x | 0.99954 | 11.219 | 183.9699 |  | 1044674.4162 | 11.2181 | 443570.4903
Compound | 915
Sample | [7 empty] | 915 Method |  | 915 Results | [10 empty] | MFBET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5725 | 398.3919 |  |  | 796.7838 | 3187.135 | 0.48014 |  |  |  | 572966.5654 | 5.6529 | 43125.2652
|  | UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 414.2254 |  |  | 828.4509 | 3313.8036 | 0.54217 |  |  |  | 633501.796 | 5.6723 | 45855.005
|  | UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 388.5011 | 400.3728 |  | 777.0023 | 3108.0091 | 0.50922 | 0.51051 | 0.031035 |  | 624713.6456 | 5.6723 | 48219.7539
|  | UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5919 | 205.0918 |  |  | 2050.9178 | 8203.6713 |  |  |  |  | 262299.4346 | 5.6684 | 38389.9729
|  | UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5997 | 192.5821 |  |  | 1925.8213 | 7703.2854 |  |  |  |  | 299018.5217 | 5.6762 | 46610.1305
|  | UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5492 | 181.4107 | 193.0282 |  | 1814.1072 | 7256.4288 |  | 7721.1285 | 473.8733 |  | 281804.2735 | 5.6295 | 46634.702
|  | UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.6036 | 165.9491 |  |  | 1659.4912 | 6637.9647 |  |  |  |  | 249607.3666 | 5.6762 | 45158.9158
|  | UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5686 | 152.8038 |  |  | 1528.0379 | 6112.1517 |  |  |  |  | 228990.1946 | 5.649 | 44996.0952
|  | UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = 26.056578 * x ^ 2 + 399.326815 * x | 0.99953 | 5.5764 | 152.5874 | 157.1134 |  | 1525.874 | 6103.4961 |  | 6284.5375 | 306.1075 | -18.61% | 222741.8559 | 5.6529 | 43830.4305
Sample | [7 empty] | 965 Method |  | 965 Results | [10 empty] | MFHET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 175.3527 |  |  | 350.7054 | 1402.8217 | 0.17909 |  |  |  | 581711.0964 | 6.3009 | 36772.6699
|  | UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 204.4494 |  |  | 408.8987 | 1635.5949 | 0.22368 |  |  |  | 660811.0812 | 6.3069 | 35943.9604
|  | UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 187.033 | 188.945 |  | 374.0659 | 1496.2636 | 0.20712 | 0.2033 | 0.022537 |  | 642376.6085 | 6.3069 | 38121.0607
|  | UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1118 | 228.1491 |  |  | 2281.4905 | 9125.9621 |  |  |  |  | 596959.8831 | 6.3039 | 29174.7978
|  | UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 206.0158 |  |  | 2060.1575 | 8240.63 |  |  |  |  | 651363.2661 | 6.3008 | 35166.8058
|  | UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.0936 | 194.392 | 209.5189 |  | 1943.9201 | 7775.6803 |  | 8380.7575 | 685.9606 |  | 623589.8648 | 6.2918 | 35634.3775
|  | UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1148 | 195.8233 |  |  | 1958.2328 | 7832.9313 |  |  |  |  | 593672.5283 | 6.3069 | 33682.1857
|  | UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 182.8059 |  |  | 1828.0586 | 7312.2343 |  |  |  |  | 595165.7019 | 6.2948 | 36119.1322
|  | UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = -1441.098499 * x ^ 2 + 1103.615825 * x | 0.99777 | 6.1027 | 180.6059 | 186.4117 |  | 1806.059 | 7224.236 |  | 7456.4672 | 328.983 | -11.03% | 531934.4626 | 6.3009 | 32667.0296
Sample | [7 empty] | 476 Method |  | 476 Results | [10 empty] | MFHET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 671.1463 |  |  | 1342.2926 | 5369.1706 | 0.50742 |  |  |  | 814200.3783 | 6.3039 | 34951.2412
|  | UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9108 | 653.8039 |  |  | 1307.6077 | 5230.4309 | 0.52473 |  |  |  | 613019.7317 | 6.2735 | 27013.7501
|  | UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 741.6943 | 688.8815 |  | 1483.3887 | 5933.5547 | 0.60308 | 0.54507 | 0.050973 |  | 702155.1796 | 6.2887 | 27272.3765
|  | UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 270.4607 |  |  | 2704.6071 | 10818.4284 |  |  |  |  | 321598.6557 | 6.2948 | 34299.1518
|  | UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9208 | 275.029 |  |  | 2750.2897 | 11001.1589 |  |  |  |  | 309051.8833 | 6.2978 | 32412.4363
|  | UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9177 | 266.9594 | 270.8164 |  | 2669.5938 | 10678.3754 |  | 10832.6542 | 161.8613 |  | 356416.7453 | 6.3069 | 38512.1517
|  | UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 264.5346 |  |  | 2645.3463 | 10581.3851 |  |  |  |  | 348808.9109 | 6.2735 | 38036.2844
|  | UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9212 | 249.1974 |  |  | 2491.9741 | 9967.8963 |  |  |  |  | 306415.8203 | 6.3009 | 35474.482
|  | UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | y = 416.856543 * x - 0.018990 | 0.99928 | 6.9246 | 245.97 | 253.234 |  | 2459.7003 | 9838.8011 |  | 10129.3608 | 396.7504 | -6.492% | 270052.2341 | 6.2826 | 31675.7212
Sample | [7 empty] | 267 Method |  | 267 Results | [10 empty] | MFOET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | y = 484.688584 * x | 0.99801 | 7.7939 | 37.924 |  |  | 75.848 | 303.3919 | 0.02772 |  |  |  | 26629.0199 | 7.117 | 17384.4001
|  | UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | y = 484.688584 * x | 0.99801 | 7.7935 | 45.4157 |  |  | 90.8313 | 363.3254 | 0.034755 |  |  |  | 32217.5211 | 7.1166 | 17563.247
|  | UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | y = 484.688584 * x | 0.99801 | 7.7939 | 50.8567 | 44.7321 |  | 101.7134 | 406.8537 | 0.044771 | 0.035748 | 0.0085689 |  | 28252.9584 | 7.1205 | 13754.1616
|  | UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | y = 484.688584 * x | 0.99801 | 7.7939 | 298.8573 |  |  | 2988.5733 | 11954.2931 |  |  |  |  | 182057.7361 | 7.1205 | 15082.165
|  | UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | y = 484.688584 * x | 0.99801 | 7.7835 | 276.4869 |  |  | 2764.8687 | 11059.4748 |  |  |  |  | 183688.8077 | 7.1066 | 16448.5129
|  | UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | y = 484.688584 * x | 0.99801 | 7.7866 | 277.1838 | 284.176 |  | 2771.8383 | 11087.353 |  | 11367.0403 | 508.7668 |  | 181926.1609 | 7.1132 | 16249.7142
|  | UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | y = 484.688584 * x | 0.99801 | 7.7904 | 273.6249 |  |  | 2736.249 | 10944.9958 |  |  |  |  | 199725.3272 | 7.1136 | 18071.5743
|  | UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | y = 484.688584 * x | 0.99801 | 7.7939 | 261.3487 |  |  | 2613.4865 | 10453.9462 |  |  |  |  | 184554.768 | 7.1171 | 17483.3019
|  | UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | y = 484.688584 * x | 0.99801 | 7.79 | 227.1863 | 254.0533 |  | 2271.8627 | 9087.4506 |  | 10162.1309 | 962.5413 | -10.6% | 161252.8244 | 7.1132 | 17572.9112
Sample | [7 empty] | 906 Method |  | 906 Results | [10 empty] | MFOET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 56.6688 |  |  | 113.3376 | 453.3503 | 0.05153 |  |  |  | 642952.0649 | 7.117 | 17384.4001
|  | UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 50.5734 |  |  | 101.1468 | 404.5873 | 0.045727 |  |  |  | 582772.9937 | 7.1166 | 17563.247
|  | UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 58.6135 | 55.2852 |  | 117.2269 | 468.9078 | 0.063792 | 0.053683 | 0.0092232 |  | 525378.1677 | 7.1205 | 13754.1616
|  | UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 268.5056 |  |  | 2685.0565 | 10740.2258 |  |  |  |  | 2551207.2326 | 7.1205 | 15082.165
|  | UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 220.7353 |  |  | 2207.3527 | 8829.4108 |  |  |  |  | 2292083.3033 | 7.1066 | 16448.5129
|  | UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 228.7215 | 239.3208 |  | 2287.2151 | 9148.8602 |  | 9572.8323 | 1023.532 |  | 2345349.6761 | 7.1132 | 16249.7142
|  | UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3664 | 219.9468 |  |  | 2199.4681 | 8797.8725 |  |  |  |  | 2509365.1856 | 7.1136 | 18071.5743
|  | UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3638 | 221.197 |  |  | 2211.9699 | 8847.8794 |  |  |  |  | 2441316.5818 | 7.1171 | 17483.3019
|  | UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | y = 7487.053151 * x + 1.627577 | 0.99933 | 9.3639 | 183.763 | 208.3023 |  | 1837.63 | 7350.5199 |  | 8332.0906 | 850.4328 | -12.96% | 2043398.9891 | 7.1132 | 17572.9112
Sample | [7 empty] | 273 Method |  | 273 Results | [10 empty] | MFOET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 691.5618 |  |  | 1383.1237 | 5532.4948 | 0.78791 |  |  |  | 305308.0013 | 7.1201 | 15322.4289
|  | UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 684.3315 |  |  | 1368.663 | 5474.6522 | 0.78705 |  |  |  | 241140.2062 | 7.117 | 12225.8416
|  | UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 715.5762 | 697.1565 |  | 1431.1524 | 5724.6096 | 0.85187 | 0.80894 | 0.037181 |  | 279136.4205 | 7.124 | 13553.8369
|  | UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 251.357 |  |  | 2513.5699 | 10054.2794 |  |  |  |  | 121087.9193 | 7.1032 | 16387.1308
|  | UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 236.3069 |  |  | 2363.0686 | 9452.2742 |  |  |  |  | 109422.5473 | 7.1132 | 15740.8529
|  | UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 240.9389 | 242.8676 |  | 2409.3893 | 9637.5571 |  | 9714.7036 | 308.3282 |  | 117724.9479 | 7.117 | 16613.0802
|  | UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6631 | 175.5438 |  |  | 1755.4375 | 7021.7502 |  |  |  |  | 97769.3431 | 7.1205 | 18880.9947
|  | UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6657 | 173.8982 |  |  | 1738.9824 | 6955.9295 |  |  |  |  | 87787.144 | 7.1101 | 17112.4072
|  | UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | y = -191.319123 * x ^ 2 + 356.774638 * x | 0.99981 | 9.6656 | 168.0008 | 172.4809 |  | 1680.0077 | 6720.0306 |  | 6899.2368 | 158.6481 | -28.98% | 78315.4156 | 7.117 | 15797.7788
Sample | [7 empty] | 913 Method |  | 913 Results | [10 empty] | MFOET (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = 4960.639652 * x | 0.99845 | 11.3794 | 367.0031 |  |  | 734.0062 | 2936.0246 | 0.34914 |  |  |  | 2614245.7392 | 7.1201 | 17231.3883
|  | UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = 4960.639652 * x | 0.99845 | 11.3794 | 407.3641 |  |  | 814.7281 | 3258.9126 | 0.33867 |  |  |  | 2744132.1562 | 7.117 | 16295.4315
|  | UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = 4960.639652 * x | 0.99845 | 11.4261 | 422.5855 | 398.9842 |  | 845.171 | 3380.6841 | 0.42096 | 0.36959 | 0.044797 |  | 3049472.0421 | 7.1171 | 17456.357
|  | UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = 4960.639652 * x | 0.99845 | 11.2159 | 238.0729 |  |  | 2380.7293 | 9522.9171 |  |  |  |  | 1513195.7042 | 7.1136 | 15375.4789
|  | UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = 4960.639652 * x | 0.99845 | 11.3171 | 242.0251 |  |  | 2420.2509 | 9681.0035 |  |  |  |  | 1648315.7882 | 7.1067 | 16474.9308
|  | UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = 4960.639652 * x | 0.99845 | 11.3716 | 197.2418 | 225.7799 |  | 1972.4183 | 7889.6732 |  | 9031.1979 | 991.7443 |  | 1484672.4604 | 7.1205 | 18208.5429
|  | UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = 4960.639652 * x | 0.99845 | 11.3482 | 210.2326 |  |  | 2102.326 | 8409.304 |  |  |  |  | 1388991.0846 | 7.1136 | 15982.4355
|  | UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = 4960.639652 * x | 0.99845 | 11.3949 | 240.5697 |  |  | 2405.6968 | 9622.7874 |  |  |  |  | 1608848.6635 | 7.1097 | 16177.7404
|  | UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = 4960.639652 * x | 0.99845 | 11.3171 | 200.7712 | 217.1912 |  | 2007.7124 | 8030.8496 |  | 8687.647 | 831.6685 | -3.804% | 1371690.8616 | 7.1171 | 16527.163
Sample | [7 empty] | 4NT Method |  | 4NT Results | [9 empty] | 4NT13C6 (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UCGCAAAFS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042038.D | Sample |  | 2020-02-05 05:17 | y = 153.621749 * x | 0.99954 | 11.219 | 55.6284 |  |  | 111.2568 | 445.027 | 0.055553 |  |  |  | 324929.2471 | 11.2181 | 456268.9968
|  | UCGCAAAFS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042039.D | Sample |  | 2020-02-05 05:39 | y = 153.621749 * x | 0.99954 | 11.2189 | 53.1138 |  |  | 106.2277 | 424.9107 | 0.053692 |  |  |  | 272083.1737 | 11.2181 | 400149.7862
|  | UCGCAAAFS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042047.D | Sample |  | 2020-02-05 08:29 | y = 153.621749 * x | 0.99954 | 11.219 | 54.2685 |  |  | 108.5369 | 434.1477 | 0.056431 | 0.055226 | 0.0013984 |  | 259199.7576 | 11.2181 | 373091.7078
|  | UCGCAAT1hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042019.D | Sample |  | 2020-02-04 22:32 | y = 153.621749 * x | 0.99954 | 11.2189 | 236.1243 |  |  | 2361.2427 | 9444.9707 |  |  |  |  | 1559477.053 | 11.2181 | 515901.6354
|  | UCGCAAT1hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042020.D | Sample |  | 2020-02-04 22:53 | y = 153.621749 * x | 0.99954 | 11.219 | 239.3649 |  |  | 2393.6486 | 9574.5942 |  |  |  |  | 1494748.7556 | 11.2181 | 487793.8886
|  | UCGCAAT1hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042053.D | Sample |  | 2020-02-05 10:37 | y = 153.621749 * x | 0.99954 | 11.2189 | 229.1937 |  |  | 2291.9371 | 9167.7485 |  | 9395.7711 | 207.8373 |  | 1583252.3854 | 11.2181 | 539605.0373
|  | UCGCAAT5hS1A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042049.D | Sample |  | 2020-02-05 09:12 | y = 153.621749 * x | 0.99954 | 11.2189 | 200.2697 |  |  | 2002.6972 | 8010.7889 |  |  |  |  | 1446510.0873 | 11.2181 | 564202.1179
|  | UCGCAAT5hS1B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042028.D | Sample |  | 2020-02-05 01:44 | y = 153.621749 * x | 0.99954 | 11.219 | 197.845 |  |  | 1978.4497 | 7913.7988 |  |  |  |  | 1199756.8571 | 11.2181 | 473692.7868
|  | UCGCAAT5hS1C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042029.D | Sample |  | 2020-02-05 02:05 | y = 153.621749 * x | 0.99954 | 11.2189 | 192.3357 |  |  | 1923.3568 | 7693.4274 |  | 7872.6717 | 162.6289 | -16.21% | 1079328.812 | 11.2181 | 438351.4479
|  | UCGCAAAFS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042059.D | Sample |  | 2020-02-05 12:45 | y = 153.621749 * x | 0.99954 | 11.2189 | 62.9781 |  |  | 125.9562 | 503.825 | 0.068872 |  |  |  | 406758.0232 | 11.2181 | 504516.0123
|  | UCGCAAAFS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042056.D | Sample |  | 2020-02-05 11:41 | y = 153.621749 * x | 0.99954 | 11.219 | 53.4542 |  |  | 106.9084 | 427.6334 | 0.056785 |  |  |  | 349758.6536 | 11.2181 | 511111.1467
|  | UCGCAAAFS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042057.D | Sample |  | 2020-02-05 12:03 | y = 153.621749 * x | 0.99954 | 11.2189 | 58.7453 |  |  | 117.4905 | 469.962 | 0.070398 | 0.065352 | 0.0074581 |  | 307408.5122 | 11.2181 | 408763.0851
|  | UCGCAAT1hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042055.D | Sample |  | 2020-02-05 11:20 | y = 153.621749 * x | 0.99954 | 11.2189 | 228.357 |  |  | 2283.5701 | 9134.2806 |  |  |  |  | 1268681.9773 | 11.2181 | 433977.2416
|  | UCGCAAT1hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042021.D | Sample |  | 2020-02-04 23:15 | y = 153.621749 * x | 0.99954 | 11.2189 | 216.1466 |  |  | 2161.4663 | 8645.8651 |  |  |  |  | 1341527.1913 | 11.2181 | 484818.9148
|  | UCGCAAT1hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042024.D | Sample |  | 2020-02-05 00:19 | y = 153.621749 * x | 0.99954 | 11.219 | 211.3997 |  |  | 2113.9965 | 8455.9861 |  | 8745.3773 | 349.9255 |  | 1347459.2993 | 11.2181 | 497897.4822
|  | UCGCAAT5hS2A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042037.D | Sample |  | 2020-02-05 04:56 | y = 153.621749 * x | 0.99954 | 11.2189 | 182.8837 |  |  | 1828.8372 | 7315.3486 |  |  |  |  | 1335265.836 | 11.2181 | 570323.466
|  | UCGCAAT5hS2B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042058.D | Sample |  | 2020-02-05 12:24 | y = 153.621749 * x | 0.99954 | 11.219 | 188.2681 |  |  | 1882.6811 | 7530.7245 |  |  |  |  | 1211653.6671 | 11.2181 | 502724.7548
|  | UCGCAAT5hS2C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042044.D | Sample |  | 2020-02-05 07:25 | y = 153.621749 * x | 0.99954 | 11.219 | 166.8947 |  |  | 1668.9467 | 6675.7867 |  | 7173.9533 | 444.6619 | -17.97% | 1046225.116 | 11.2181 | 489678.7122
|  | UCGCAAAFS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042040.D | Sample |  | 2020-02-05 06:00 | y = 153.621749 * x | 0.99954 | 11.219 | 50.3812 |  |  | 100.7624 | 403.0495 | 0.053683 |  |  |  | 327375.3083 | 11.2181 | 507581.8712
|  | UCGCAAAFS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042041.D | Sample |  | 2020-02-05 06:21 | y = 153.621749 * x | 0.99954 | 11.2189 | 63.4059 |  |  | 126.8118 | 507.2471 | 0.069594 |  |  |  | 403640.3802 | 11.2181 | 497271.4447
|  | UCGCAAAFS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042052.D | Sample |  | 2020-02-05 10:16 | y = 153.621749 * x | 0.99954 | 11.219 | 52.1492 |  |  | 104.2983 | 417.1933 | 0.056693 | 0.05999 | 0.0084522 |  | 376158.2799 | 11.2181 | 563445.3712
|  | UCGCAAT1hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042025.D | Sample |  | 2020-02-05 00:40 | y = 153.621749 * x | 0.99954 | 11.2189 | 227.5675 |  |  | 2275.6746 | 9102.6985 |  |  |  |  | 1251495.903 | 11.2181 | 429583.7135
|  | UCGCAAT1hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042027.D | Sample |  | 2020-02-05 01:23 | y = 153.621749 * x | 0.99954 | 11.2189 | 225.7573 |  |  | 2257.5733 | 9030.2931 |  |  |  |  | 1438599.7058 | 11.2181 | 497767.6323
|  | UCGCAAT1hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042048.D | Sample |  | 2020-02-05 08:51 | y = 153.621749 * x | 0.99954 | 11.219 | 211.6066 |  |  | 2116.0657 | 8464.2628 |  | 8865.7514 | 349.579 |  | 1420407.5714 | 11.2181 | 524339.2574
|  | UCGCAAT5hS3A_GC 2/3 | PFAS Group 2 UC | 7010_1_02042045.D | Sample |  | 2020-02-05 07:47 | y = 153.621749 * x | 0.99954 | 11.2189 | 187.6978 |  |  | 1876.9778 | 7507.9112 |  |  |  |  | 1119892.7533 | 11.2181 | 466064.303
|  | UCGCAAT5hS3B_GC 2/3 | PFAS Group 2 UC | 7010_1_02042030.D | Sample |  | 2020-02-05 02:27 | y = 153.621749 * x | 0.99954 | 11.219 | 182.2168 |  |  | 1822.1685 | 7288.6739 |  |  |  |  | 1154335.3982 | 11.2181 | 494848.2367
|  | UCGCAAT5hS3C_GC 2/3 | PFAS Group 2 UC | 7010_1_02042036.D | Sample |  | 2020-02-05 04:35 | y = 153.621749 * x | 0.99954 | 11.219 | 183.9699 |  |  | 1839.6993 | 7358.7974 |  | 7385.1275 | 111.9652 | -16.7% | 1044674.4162 | 11.2181 | 443570.4903
[5 empty] | 3125 Method |  | 3125 Results | [10 empty] | 4NT Method |  | 4NT Results
Name | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Area | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | CF Formula | CF R2 | RT | Area | Final Conc. | Accuracy
UCGCAAAFS1A_GC 2/3 | 7010_2_03032036.D | Sample |  | 2020-03-03 22:22 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2460.2076 | 16.98 |  |  | 33.96 | 135.84 | 0.025665 |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 43377.5179 | 53.0758
UCGCAAAFS1B_GC 2/3 | 7010_2_03032037.D | Sample |  | 2020-03-03 22:42 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 3018.0244 | 16.8063 |  |  | 33.6127 | 134.4506 | 0.02449 |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41931.0929 | 50.6106
UCGCAAAFS1C_GC 2/3 | 7010_2_03032046.D | Sample |  | 2020-03-04 01:48 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 2284.9813 | 15.7352 | 16.5072 |  | 31.4705 | 125.8818 | 0.025619 | 0.025258 | 0.00066538 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 41590.239 | 52.5845
UCGCAAT1hS1A_GC 2/3 | 7010_2_03032022.D | Sample |  | 2020-03-03 17:34 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 31973.0127 | 204.482 |  |  | 2044.8201 | 8179.2804 |  |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 183985.2055 | 235.8448
UCGCAAT1hS1B_GC 2/3 | 7010_2_03032023.D | Sample |  | 2020-03-03 17:54 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 33892.3029 | 194.5951 |  |  | 1945.9507 | 7783.803 |  |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 194178.0654 | 234.8444
UCGCAAT1hS1C_GC 2/3 | 7010_2_03032051.D | Sample |  | 2020-03-04 03:30 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 32317.108 | 183.3286 | 194.1352 |  | 1833.2863 | 7333.145 |  | 7765.4095 | 423.3675 |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 186786.148 | 228.8408
UCGCAAT5hS1A_GC 2/3 | 7010_2_03032049.D | Sample |  | 2020-03-04 02:49 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 25828.324 | 132.3226 |  |  | 1323.2257 | 5292.903 |  |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 198589.9057 | 199.9504
UCGCAAT5hS1B_GC 2/3 | 7010_2_03032029.D | Sample |  | 2020-03-03 19:58 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3572 | 20779.909 | 137.2516 |  |  | 1372.5162 | 5490.0647 |  |  |  |  | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 156886.9825 | 196.2927
UCGCAAT5hS1C_GC 2/3 | 7010_2_03032030.D | Sample |  | 2020-03-03 20:19 | y = 2478.481268 * x ^ 2 + 1485.332714 * x + 0.017548 | 0.99042 | 10.3573 | 22140.7158 | 122.8421 | 130.8054 |  | 1228.4206 | 4913.6823 |  | 5232.2167 | 292.9442 | -32.62% | y = 153.340196 * x + 0.007796 | 0.99968 | 11.498 | 164350.6487 | 189.4232