## Cover Sheet
Analytical Cover Sheet
QAPP: | In Vitro TK Assessments of PFAS (D-CED-0031343)
Analysis Set ID: | 2019_PFAS_PPB_UC_AMIDES_A_ALK | *Ultracentrifugation assay to assess plasma protein binding with PFAS acids
Analytes: | DTXSID2060965, DTXSID60400587, DTXSID70366226, DTXSID80310730, DTXSID60195123 | PFAS
Dates Prepared: | 8/13/2019, 10/11/2019, 12/6/2019 | PPB_CC
Assay Preparer: | ALK
Dates Prepared: | 8/9/2019, 10/9/2019, 12/6/2019 | PPB_UC
Dates of MS Run, resp.: | 10/29/2019, 11/1/2019, 12/10/2019
Assay Preparer: | ALK
Report Preparer: | ALK
Sample Type: | Standards, QCs, Unknown Analyte Samples | CC1-CC15
Internal Standard (conc) | Perfluoro-1-[13C8]octanesulfonamide [lot M8FOSA1017] (3 pg/uL); 13C6-4-Nitrotoluene [lot SDFK-011] (3 pg/uL) | Amides; Ref
Total # Samples (not including blanks): | 171
Actual # Samples (not including replicates): | 141
MDL Calculation |  |  |  | References:
|  |  |  | 40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
eLOQ Definition | Lowest curve concentration at which precision checks can be quantitated against curve ±30%
QC Samples and Acceptance Criteria:
QC Type | Description | Acceptance Criteria | Type
QC Sample Type | Frequency | Acceptance Criteria | QC Metric | Corrective Action
(Analytical, TK Assay, or Both) |  |  | Assessed | if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | r2≥0.98; | Linearity | Rerun samples and curve
7-pt curve min.
Curve Check (Analytical) | 2-3 samples per run | 75-125 % | Accuracy | Flag in comments
Instrument Blanks (Analytical) | 1 every 6-10 injections | <½ LOD of instrument method | Specificity | Flag in comments, assess and resolve
Technical Replicate assessment | Every assay sample | 75-125% | Precision | Flag in comments
Reference Chemical(s) (Assay) | 1 per assay batch | RSD = ±20% of historical or published values | Reproducibility | Flag in comments
Analyte-IS Matching
Analyte | Name | Sample ID | IS | Reference Compound | Avg. MW
DTXSID2060965 | Heptafluorobutyramide | 908 | M8FOSA | 4-nitrotoluene | 213.055
DTXSID60400587 | Nonafluoropentanamide | 909 |  |  | 263.063
DTXSID70366226 | Perfluoropentanamide | 916 |  |  | 245.072
DTXSID80310730 | Octafluoroadipamide | 923 |  |  | 288.097
DTXSID60195123 | Perfluorooctanamide | 3117 |  |  | 413.086
Sample ID Key
CC | Calibration Curve
QC | Quality Check
UC_S#letterAF | Ultracentrifugation Assay_Mix Replicate Aqueous Fraction
UC_S#letterT# | Ultracentrifugation Assay_Mix Replicate Stability Timepoint
Review of Data Timeline
Person | Action
ALK | Assays conducted 8/9, 10/9, 12/6/2019
MSC | Analytical data generated 10/29, 11/1, 12/10/2019
ALK | QC Report completed 9/16/2020
Additional Notes
Concentrations listed are at the instrument, unless otherwise specified

## Executive Summary
Executive Summary
|  | Quantitative Limits |  |  | Plasma Protein Binding
DTXSID | Name | Sample ID | LOD (nM) | eLOQ (nM) | 4-nitrotoluene |  |  |  |  | Perfluoropentanamide
DTXSID2060965 | Heptafluorobutyramide | 908 | 0.46 | 1.75 | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
DTXSID60400587 | Nonafluoropentanamide | 909 | 0.59 | 1.75 | AF | 0.36556 | 2019-10-29 | 0.060815 | 0.065365 | AF | 2.8683 | 2019-10-29 | 0.60148 | 0.55693
DTXSID70366226 | Perfluoropentanamide | 916 | 0.47 | 3; 1.75 on 11/1 CC |  |  | 2019-11-01 | 0.069349 |  |  |  | 2019-11-01 | 0.61889
DTXSID80310730 | Octafluoroadipamide | 923 | 7.91 | 5; 12.5 on 12/10 CC |  |  | 2019-12-10 | 0.06593 |  |  |  | 2019-12-10 | 0.45042
DTXSID60195123 | Perfluorooctanamide | 3117 | 0.6 | 1.75; 3 on 1st 12/10 CC | T1h | 6.0753 |  |  |  | T1h | 5.7197
DTXSID5023792 | 4-nitrotoluene | 4NT | 0.36 | 12.5nM 10/29 & 11/1; 1.75 12/10
[5 empty] | T5h | 5.6294 |  |  |  | T5h | 5.292
[5 empty] | Heptafluorobutyramide |  |  |  |  | Octafluoroadipamide
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 5.5423 | 2019-10-29 | 0.60092 | 0.60297 | AF | 3.7827 | 2019-10-29 | 0.66804 | 0.81178
[7 empty] | 2019-11-01 | 0.60503 |  |  |  | 2019-11-01 | 0.89716
[7 empty] | 2019-12-10 | N/A |  |  |  | 2019-12-10 | 0.87014
[5 empty] | T1h | 10.0715 |  |  |  | T1h | 5.0507
[5 empty] | T5h | 9.2441 |  |  |  | T5h | 4.4282
[5 empty] | Nonafluoropentanamide |  |  |  |  | Perfluorooctanamide
[5 empty] | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu | Sample Text | Avg. Measured Conc. (μM) | GC Date | Fraction Unbound (Fu) | Avg Fu
[5 empty] | AF | 1.8515 | 2019-10-29 | 0.19933 | 0.20311 | AF | 0.17572 | 2019-10-29 | 0.026637 | 0.028193
[7 empty] | 2019-11-01 | 0.20689 |  |  |  | 2019-11-01 | N/A
[7 empty] | 2019-12-10 | N/A |  |  |  | 2019-12-10 | 0.02975
[5 empty] | T1h | 9.9855 |  |  |  | T1h | 7.9147
[5 empty] | T5h | 9.3934 |  |  |  | T5h | 6.406

## FractionUnbound
Ultracentrifugation Plasma protein binding assay - Experimental data and Fu calculations
|  |  |  | Mean | SD | CV | Mean |  | Time 1 hr |  |  | Time 5 hr |  |  | Aqueous Fraction |  |  | Fraction Unbound (fu) |  |  | Stability (T5hr/ T1hr)
DTXSID | Chemical | AbbrevSampleID | CAS# | Fu | Fu | Fu | Stability | UC Assay Date | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | t test
DTXSID2060965 | Heptafluorobutyramide | 908 | 662-50-0 | 0.60297 | 0.072721 | 12.0604 | 0.9244 | Run-specific Avg | 9336.0975 | 10806.9467 | NA | 8971.0714 | 9517.1723 | NA | 5392.0092 | 5692.5358 | NA | 0.60092 | 0.60503 | NA | 0.9609 | 0.88065 | NA | 0.11528
[8 empty] | Run-specific CV | [9 empty] | 0.095644 | 0.16449 | NA
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC | 10040.2707 | 8718.1158 | 9249.9061 | 9020.544 | 8615.3735 | 9277.2968 | 6018.7286 | 4913.6275 | 5243.6716 | 0.66722 | 0.57033 | 0.56522 | 0.89844 | 0.98822 | 1.003
[8 empty] | 10/9 UC, 11/1 GC | 11241.4456 | 9422.5055 | 11756.889 | 8437.6948 | 9272.6665 | 10841.1555 | 6072.1841 | 5144.944 | 5860.4793 | 0.71965 | 0.55485 | 0.54058 | 0.75059 | 0.9841 | 0.92211
[9 empty] | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06
DTXSID60400587 | Nonafluoropentanamide | 909 | 13485-61-5 | 0.19849 | 0.028876 | 14.548 | 0.94696 | Run-specific Avg | 9452.1187 | 10518.8719 | NA | 9187.7737 | 9599.1221 | NA | 1830.4887 | 1872.5425 | NA | 0.19933 | 0.20689 | NA | 0.97203 | 0.91256 | NA | 0.24042
[8 empty] | Run-specific CV | [9 empty] | 0.13049 | 0.18971 | NA
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC | 10113.7416 | 8734.5576 | 9508.057 | 9001.9577 | 8863.6707 | 9697.6928 | 2056.7322 | 1581.8386 | 1852.8951 | 0.22848 | 0.17846 | 0.19107 | 0.89007 | 1.0148 | 1.0199
[8 empty] | 10/9 UC, 11/1 GC | 10851.6486 | 9171.4901 | 11533.4771 | 8398.349 | 9600.6943 | 10798.323 | 2022.2565 | 1660.709 | 1934.6621 | 0.24079 | 0.17298 | 0.17916 | 0.77392 | 1.0468 | 0.93626
[9 empty] | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06
DTXSID70366226 | Perfluoropentanamide | 916 | 355-81-7 | 0.54171 | 0.096514 | 17.8164 | 0.93156 | Run-specific Avg | 4161.8645 | 5529.712 | 7467.4971 | 3829.7574 | 5273.5873 | 6772.7556 | 2279.212 | 3295.2978 | 3030.2552 | 0.60148 | 0.61889 | 0.45042 | 0.9202 | 0.95368 | 0.90696 | 0.095937
[8 empty] | Run-specific CV | [9 empty] | 0.63925 | 0.60408 | 0.12808
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC | 4095.9058 | 4172.1055 | 4217.5821 | 3665.5834 | 3923.1826 | 3900.5063 | 2291.4746 | 2266.9494 | 6458.3434 | 0.62513 | 0.57783 | 1.6558 | 0.89494 | 0.94034 | 0.92482
[8 empty] | 10/9 UC, 11/1 GC | 4947.5419 | 5944.7974 | 5696.7965 | 5187.4522 | 5483.1346 | 5150.1753 | 3444.0068 | 3146.5887 | 8284.0441 | 0.66391 | 0.57387 | 1.6085 | 1.0485 | 0.92234 | 0.90405
[8 empty] | 12/6 UC, 12/10 GC | 7056.429 | 7030.6489 | 8315.4133 | 6751.5179 | 7314.2106 | 6252.5384 | 3011.9364 | 2889.1305 | 3189.6988 | 0.44611 | 0.395 | 0.51014 | 0.95679 | 1.0403 | 0.75192
[9 empty] | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06
DTXSID80310730 | Octafluoroadipamide | 923 | 355-66-8 | 0.82012 | 0.13096 | 15.9689 | 0.94379 | Run-specific Avg | 3276.4618 | 4690.7442 | 7184.8348 | 3070.5632 | 5450.7678 | 4763.3721 | 2151.3564 | 5073.8077 | 4122.8203 | 0.66804 | 0.89716 | 0.87014 | 0.93716 | 1.162 | 0.66298 | 0.29999
[8 empty] | Run-specific CV | [9 empty] | 0.84349 | 0.61331 | 0.14343
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC | 2552.6823 | 3416.1861 | 3860.5171 | 3300.1272 | 3131.7374 | 2779.8249 | 2321.7176 | 1980.9952 | 7140.1805 | 0.70352 | 0.63255 | 2.5686 | 1.2928 | 0.91674 | 0.72007
[8 empty] | 10/9 UC, 11/1 GC | 4082.0435 | 5048.8845 | 4941.3045 | 5843.8616 | 5448.2446 | 5060.1972 | 5490.9299 | 4656.6856 | 11991.8322 | 0.93961 | 0.85471 | 2.3698 | 1.4316 | 1.0791 | 1.0241
[8 empty] | 12/6 UC, 12/10 GC | 7544.5257 | 7932.1355 | 6077.8433 | 4734.0692 | 4420.6245 | 5135.4227 | 4614.8609 | 3994.3871 | 3759.213 | 0.97482 | 0.90358 | 0.73202 | 0.62748 | 0.55731 | 0.84494
[9 empty] | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06
DTXSID60195123 | Perfluorooctanamide | 3117 | 423-54-1 | 0.028193 | 0.0052592 | 18.6541 | 0.58411 | Run-specific Avg | 9601.585 | NC | 6227.899 | 7784.0508 | NC | 5027.9308 | 204.0906 | NA | 147.3501 | 0.026637 | NA | 0.02975 | 0.8107 | NA | 0.80732 | 0.10004
[8 empty] | Run-specific CV | [9 empty] | 0.18998 | 0.035397 | 0.20245
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC | 8325.333 | 11802.896 | 8676.5259 | 8907.8824 | 7238.8387 | 7205.4314 | 185.7934 | 208.3703 | 218.108 | 0.020857 | 0.028785 | 0.03027 | 1.07 | 0.61331 | 0.83045
[8 empty] | 10/9 UC, 11/1 GC | 10280.3749 | 6716.5239 | 6829.2544 | 667.9436 | 730.3315 | 769.6188 | 1631.5613 | 1668.4301 | 1784.8827 | 2.4427 | 2.2845 | 2.3192 | 0.064973 | 0.10874 | 0.11269
[8 empty] | 12/6 UC, 12/10 GC | 6849.6302 | 5468.3077 | 6365.7592 | 5187.8988 | 5618.9627 | 4276.931 | 127.7074 | 158.7489 | 155.594 | 0.024616 | 0.028252 | 0.03638 | 0.7574 | 1.0276 | 0.67187
[9 empty] | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06 | 2020-08-09 | 2020-10-09 | 2020-12-06
DTXSID5023792 | 4-Nitrotoluene | 4NT | 99-99-0 | 0.064441 | 0.0081777 | 12.6903 | 0.82455 | Run-specific Avg | 6858.5913 | 3289.1457 | 8078.1456 | 5650.8226 | 2809.7885 | 8427.7273 | 356.1409 | 184.481 | 556.0639 | 0.060815 | 0.069349 | 0.06593 | 0.82904 | 0.85743 | 1.2648 | 0.04867
[8 empty] | Run-specific CV | [9 empty] | 0.13825 | 0.52903 | 0.050468
[9 empty] | A | B | C | A | B | C | A | B | C | A | B | C | A | B | C
[8 empty] | 8/9 UC, 10/29 GC- 1 | 8167.2063 | 7664.9002 | 8197.8987 | 6351.6352 | 6455.8455 | 6439.005 | 444.3562 | 341.7792 | 397.9333 | 0.069959 | 0.052941 | 0.0618 | 0.7777 | 0.84226 | 0.78545
[8 empty] | 8/9 UC, 10/29 GC- 2 | 4810.0489 | 4986.4352 | 5060.5875 | 4320.1018 | 4238.9201 | 4212.3725 | 234.1026 | 258.7356 | 561.4516 | 0.054189 | 0.061038 | 0.13329 | 0.89814 | 0.85009 | 0.83239
[8 empty] | 8/9 UC, 10/29 GC- 3 | 7466.0414 | 7610.1409 | 7764.0624 | 6273.6769 | 6362.4219 | 6203.4246 | 397.9876 | 410.5423 | 363.6903 | 0.063438 | 0.064526 | 0.058627 | 0.84029 | 0.83605 | 0.79899
[8 empty] | 10/9 UC, 11/1 GC-1 | 1954.2695 | 1818.1095 | 1901.0281 | 1624.4261 | 1540.9086 | 1723.7609 | 101.2771 | 112.298 | 105.9022 | 0.062346 | 0.072878 | 0.061437 | 0.83122 | 0.84753 | 0.90675
[8 empty] | 10/9 UC, 11/1 GC-2 | 4431.5057 | 4797.523 | 4832.4381 | 4068.5126 | 3945.2845 | 3955.8385 | 356.4699 | 246.4577 | 637.237 | 0.087617 | 0.062469 | 0.16109 | 0.91809 | 0.82236 | 0.8186
[8 empty] | 10/9 UC, 11/1 GC-3 | 8042.1088 | 7354.2937 | 7366.9209 | 2355.4909 | 2275.1757 | 2306.964 | 1245.2229 | 1310.018 | 1449.4279 | 0.52865 | 0.57579 | 0.62828 | 0.29289 | 0.30937 | 0.31315
[8 empty] | 12/6 UC, 12/10 GC | 3955.8385 | 10330.806 | 9947.7923 | 8430.1037 | 8299.8226 | 8553.2557 | 555.4359 | 506.0104 | 606.7453 | 0.065887 | 0.060966 | 0.070937 | 2.1311 | 0.80341 | 0.85981
Highlighted cells indicate omitted outliers. These variable values were identified based on the reference standard, 4NT, and are likely due to contamination of the AF with protein.

## SampleIDs
2019-10-29 | [8 empty] | Total # Samples: | 45
Sample | [8 empty] | Actual # Samples: | 42
| Name | Comment | Data File | Type | Level | Conc. pre-Crash (nM) | Conc. after Crash (nM) | Acq. Date-Time
| MTBE |  | 102919000.D | Blank |  |  |  | 2019-10-29 11:04
| UCG1CC1_8/13/2019_SE 10/15 |  | 102919001.D | Cal | 1 | 7 | 1.75 | 2019-10-29 11:34
| UCG1CC2_8/13/2019_SE 10/15 |  | 102919002.D | Cal | 2 | 12 | 3 | 2019-10-29 12:04
| UCG1CC3_8/13/2019_SE 10/15 |  | 102919003.D | Cal | 3 | 20 | 5 | 2019-10-29 12:34
| UCG1CC4_8/13/2019_SE 10/15 |  | 102919004.D | Cal | 4 | 30 | 7.5 | 2019-10-29 13:04
| UCG1CC5_8/13/2019_SE 10/15 |  | 102919005.D | Cal | 5 | 50 | 12.5 | 2019-10-29 13:34
| UCG1CC6_8/13/2019_SE 10/15 |  | 102919006.D | Cal | 6 | 80 | 20 | 2019-10-29 14:05
| UCG1CC7_8/13/2019_SE 10/15 |  | 102919007.D | Cal | 7 | 125 | 31.25 | 2019-10-29 14:35
| UCG1CC8_8/13/2019_SE 10/15 |  | 102919008.D | Cal | 8 | 200 | 50 | 2019-10-29 15:05
| UCG1CC9_8/13/2019_SE 10/15 |  | 102919009.D | Cal | 9 | 350 | 87.5 | 2019-10-29 15:35
| UCG1CC10_8/13/2019_SE 10/15 |  | 102919010.D | Cal | 10 | 500 | 125 | 2019-10-29 16:05
| UCG1CC11_8/13/2019_SE 10/15 |  | 102919011.D | Cal | 11 | 800 | 200 | 2019-10-29 16:35
| UCG1CC12_8/13/2019_SE 10/15 |  | 102919012.D | Cal | 12 | 1500 | 375 | 2019-10-29 17:05
| UCG1CC13_8/13/2019_SE 10/15 |  | 102919013.D | Cal | 13 | 2500 | 625 | 2019-10-29 17:35
| UCG1CC14_8/13/2019_SE 10/15 |  | 102919014.D | Cal | 14 | 3500 | 875 | 2019-10-29 18:05
| UCG1CC15_8/13/2019_SE 10/15 |  | 102919015.D | Cal | 15 | 5000 | 1250 | 2019-10-29 18:35
| UCG1T1hS1A_SE 10/11 |  | 102919016.D | Sample |  |  |  | 2019-10-29 19:05
| UCG1T1hS1B_SE 10/11 |  | 102919017.D | Sample |  |  |  | 2019-10-29 19:35
| UCG1T1hS2B_SE 10/11 |  | 102919018.D | Sample |  |  |  | 2019-10-29 20:05
| UCG1T1hS2C_SE 10/11 |  | 102919019.D | Sample |  |  |  | 2019-10-29 20:35
| UCG1T1hS3A_SE 10/11 |  | 102919020.D | Sample |  |  |  | 2019-10-29 21:06
| UCG1T1hS3B_SE 10/11 |  | 102919021.D | Sample |  |  |  | 2019-10-29 21:36
| MTBE |  | 102919022.D | Blank |  |  |  | 2019-10-29 22:06
| UCG1T5hS1B_SE 10/11 |  | 102919023.D | Sample |  |  |  | 2019-10-29 22:36
| UCG1T5hS1C_SE 10/11 |  | 102919024.D | Sample |  |  |  | 2019-10-29 23:06
| UCG1T5hS3B_SE 10/11 |  | 102919025.D | Sample |  |  |  | 2019-10-29 23:36
| UCG1T5hS3C_SE 10/11 |  | 102919026.D | Sample |  |  |  | 2019-10-30 00:06
| UCG1T5hS2A_SE 10/11 |  | 102919027.D | Sample |  |  |  | 2019-10-30 00:36
| UCG1CC8_8/13/2019_SE 10/15 |  | 102919028.D | QC | 8 | 200 | 50 | 2019-10-30 01:07
| UCG1AFS1A_SE 10/11 |  | 102919029.D | Sample |  |  |  | 2019-10-30 01:37
| UCG1AFS1B_SE 10/11 |  | 102919030.D | Sample |  |  |  | 2019-10-30 02:07
| UCG1AFS3A_SE 10/11 |  | 102919031.D | Sample |  |  |  | 2019-10-30 02:37
| UCG1AFS3B_SE 10/11 |  | 102919032.D | Sample |  |  |  | 2019-10-30 03:07
| UCG1T5hS2C_SE 10/11 |  | 102919033.D | Sample |  |  |  | 2019-10-30 03:37
| MTBE |  | 102919034.D | Blank |  |  |  | 2019-10-30 04:07
| UCG1T5hS3A_SE 10/11 |  | 102919035.D | Sample |  |  |  | 2019-10-30 04:37
| UCG1AFS1C_SE 10/11 |  | 102919036.D | Sample |  |  |  | 2019-10-30 05:08
| UCG1T1hS3C_SE 10/11 |  | 102919037.D | Sample |  |  |  | 2019-10-30 05:38
| UCG1T5hS1A_SE 10/11 |  | 102919038.D | Sample |  |  |  | 2019-10-30 06:08
| UCG1AFS3C_SE 10/11 |  | 102919039.D | Sample |  |  |  | 2019-10-30 06:38
| UCG1CC4_8/13/2019_SE 10/15 |  | 102919040.D | QC | 4 | 30 | 7.5 | 2019-10-30 07:08
| UCG1T1hS1C_SE 10/11 |  | 102919041.D | Sample |  |  |  | 2019-10-30 07:38
| UCG1T1hS2A_SE 10/11 |  | 102919042.D | Sample |  |  |  | 2019-10-30 08:08
| UCG1AFS2B_SE 10/11 |  | 102919043.D | Sample |  |  |  | 2019-10-30 08:38
| UCG1AFS2C_SE 10/11 |  | 102919044.D | Sample |  |  |  | 2019-10-30 09:08
| UCG1T5hS2B_SE 10/11 |  | 102919045.D | Sample |  |  |  | 2019-10-30 09:38
| UCG1AFS2A_SE 10/11 |  | 102919046.D | Sample |  |  |  | 2019-10-30 10:08
| UCG1CC13_8/13/2019_SE 10/15 |  | 102919047.D | Cal |  | 2500 | 625 | 2019-10-30 10:38
| MTBE |  | 102919048.D | Blank |  |  |  | 2019-10-30 11:09
2019-11-01 | [8 empty] | Total # Samples: | 79
Sample | [8 empty] | Actual # Samples: | 69
| Name | Comment | Data File | Type | Level | Conc. pre-Crash (nM) | Conc. after Crash (nM) | Acq. Date-Time
| MTBE |  | 110119000.D | Blank |  |  |  | 2019-11-01 08:57
| G1+CC Spiked MM_Nov.1 |  | 110119001.D | MatrixBlank |  |  |  | 2019-11-01 09:27
| G1+CC Spiked MM_Nov.1 |  | 110119002.D | MatrixBlank |  |  |  | 2019-11-01 10:22
| G1+CC CC15_Nov.1 |  | 110119003.D | Cal | 15 | 5000 | 1250 | 2019-11-01 10:52
| G1+CC CC14_Nov.1 |  | 110119004.D | Cal | 14 | 3500 | 875 | 2019-11-01 11:22
| G1+CC CC13_Nov.1 |  | 110119005.D | Cal | 13 | 2500 | 625 | 2019-11-01 11:52
| G1+CC CC12_Nov.1 |  | 110119006.D | Cal | 12 | 1500 | 375 | 2019-11-01 12:22
| G1+CC CC11_Nov.1 |  | 110119007.D | Cal | 11 | 800 | 200 | 2019-11-01 12:52
| G1+CC CC10_Nov.1 |  | 110119008.D | Cal | 10 | 500 | 125 | 2019-11-01 13:22
| G1+CC CC9_Nov.1 |  | 110119009.D | Cal | 9 | 350 | 87.5 | 2019-11-01 13:52
| G1+CC CC8_Nov.1 |  | 110119010.D | Cal | 8 | 200 | 50 | 2019-11-01 14:22
| G1+CC CC7_Nov.1 |  | 110119011.D | Cal | 7 | 125 | 31.25 | 2019-11-01 14:52
| G1+CC CC6_Nov.1 |  | 110119012.D | Cal | 6 | 80 | 20 | 2019-11-01 15:22
| G1+CC CC5_Nov.1 |  | 110119013.D | Cal | 5 | 50 | 12.5 | 2019-11-01 15:52
| G1+CC CC4_Nov.1 |  | 110119014.D | Cal | 4 | 30 | 7.5 | 2019-11-01 16:22
| G1+CC CC3_Nov.1 |  | 110119015.D | Cal | 3 | 20 | 5 | 2019-11-01 16:52
| G1+CC CC2_Nov.1 |  | 110119016.D | Cal | 2 | 12 | 3 | 2019-11-01 17:22
| G1+CC CC1_Nov.1 |  | 110119017.D | Cal | 1 | 7 | 1.75 | 2019-11-01 17:52
| UCG1 T1 S1A1009_SE 11/1 |  | 110119018.D | Sample |  |  |  | 2019-11-01 18:22
| UCG1 T1 S1B1009_SE 11/1 |  | 110119019.D | Sample |  |  |  | 2019-11-01 18:52
| UCG1 T1 S1C1009_SE 11/1 |  | 110119020.D | Sample |  |  |  | 2019-11-01 19:22
| UCG1 T1 S1C1009_SE 11/1b |  | 110119021.D | Sample |  |  |  | 2019-11-01 19:52
| MTBE |  | 110119022.D | Blank |  |  |  | 2019-11-01 20:23
| UCG1 T1 S1C1009_SE 11/1c |  | 110119023.D | Sample |  |  |  | 2019-11-01 20:53
| UCG1 AF S1A1009_SE 11/1 |  | 110119024.D | Sample |  |  |  | 2019-11-01 21:23
| UCG1 AF S1B1009_SE 11/1 |  | 110119025.D | Sample |  |  |  | 2019-11-01 21:53
| UCG1 AF S1C1009_SE 11/1 |  | 110119026.D | Sample |  |  |  | 2019-11-01 22:23
| UCG1 T5 S1A1009_SE 11/1 |  | 110119027.D | Sample |  |  |  | 2019-11-01 22:53
| UCG1 T5 S1B1009_SE 11/1 |  | 110119028.D | Sample |  |  |  | 2019-11-01 23:23
| G1+CC CC10_Nov.1 |  | 110119029.D | QC | 10 | 500 | 125 | 2019-11-01 23:53
| QCGCCC11_1015 |  | 110119030.D | QC | 11 | 800 | 200 | 2019-11-02 00:23
| UCG1 T5 S1C1009_SE 11/1 |  | 110119031.D | Sample |  |  |  | 2019-11-02 00:54
| UCG1 AF S2A1009_SE 11/1 |  | 110119032.D | Sample |  |  |  | 2019-11-02 01:24
| UCG1 AF S2B1009_SE 11/1 |  | 110119033.D | Sample |  |  |  | 2019-11-02 01:54
| UCG1 AF S2C1009_SE 11/1 |  | 110119034.D | Sample |  |  |  | 2019-11-02 02:24
| UCG1 T5 S2A1009_SE 11/1 |  | 110119035.D | Sample |  |  |  | 2019-11-02 02:54
| MTBE |  | 110119036.D | Blank |  |  |  | 2019-11-02 03:24
| UCG1 T5 S2B1009_SE 11/1 |  | 110119037.D | Sample |  |  |  | 2019-11-02 03:54
| UCG1 T5 S2B1009_SE 11/1b |  | 110119038.D | Sample |  |  |  | 2019-11-02 04:24
| UCG1 T5 S2C1009_SE 11/1 |  | 110119039.D | Sample |  |  |  | 2019-11-02 04:54
| UCG1 T1 S2A1009_SE 11/1 |  | 110119040.D | Sample |  |  |  | 2019-11-02 05:24
| UCG1 T1 S2B1009_SE 11/1 |  | 110119041.D | Sample |  |  |  | 2019-11-02 05:54
| G1+CC CC6_Nov.1 |  | 110119042.D | QC | 6 | 80 | 20 | 2019-11-02 06:24
| UCG1 T1 S2C1009_SE 11/1 |  | 110119043.D | Sample |  |  |  | 2019-11-02 06:54
| UCG1 T1 S3A1009_SE 11/1 |  | 110119044.D | Sample |  |  |  | 2019-11-02 07:24
| UCG1 T1 S3B1009_SE 11/1 |  | 110119045.D | Sample |  |  |  | 2019-11-02 07:54
| UCG1 T1 S3C1009_SE 11/1 |  | 110119046.D | Sample |  |  |  | 2019-11-02 08:24
| UCG1 AF S3A1009_SE 11/1 |  | 110119047.D | Sample |  |  |  | 2019-11-02 08:54
| QCGCCC8_1015 |  | 110119048.D | QC | 8 | 200 | 50 | 2019-11-02 09:25
| MTBE |  | 110119049.D | Blank |  |  |  | 2019-11-02 09:55
| UCG1 AF S3B1009_SE 11/1 |  | 110119050.D | Sample |  |  |  | 2019-11-02 10:25
| UCG1 AF S3C1009_SE 11/1 |  | 110119051.D | Sample |  |  |  | 2019-11-02 10:55
| UCG1 T5 S3A1009_SE 11/1 |  | 110119052.D | Sample |  |  |  | 2019-11-02 11:25
| UCG1 T5 S3B1009_SE 11/1 |  | 110119053.D | Sample |  |  |  | 2019-11-02 11:55
| UCG1 T5 S3B1009_SE 11/1b |  | 110119054.D | Sample |  |  |  | 2019-11-02 12:25
| G1+CC CC11_Nov.1 |  | 110119055.D | QC | 11 | 800 | 200 | 2019-11-02 12:55
| UCG1 T5 S3C1009_SE 11/1 |  | 110119056.D | Sample |  |  |  | 2019-11-02 13:25
| UCG1 AF S1A809_Cr10/11_SE 11/1 |  | 110119057.D | Sample |  |  |  | 2019-11-02 13:55
| UCG1 AF S1B809_Cr10/11_SE 11/1 |  | 110119058.D | Sample |  |  |  | 2019-11-02 14:25
| UCG1 AF S1C809_Cr10/11_SE 11/1 |  | 110119059.D | Sample |  |  |  | 2019-11-02 14:55
| UCG1 T5 S1A809_Cr10/11_SE 11/1 |  | 110119060.D | Sample |  |  |  | 2019-11-02 15:25
| MTBE |  | 110119061.D | Blank |  |  |  | 2019-11-02 15:55
| UCG1 T5 S1B809_Cr10/11_SE 11/1 |  | 110119062.D | Sample |  |  |  | 2019-11-02 16:26
| UCG1 T5 S1C809_Cr10/11_SE 11/1 |  | 110119063.D | Sample |  |  |  | 2019-11-02 16:56
| UCG1 T1 S1A809_Cr10/11_SE 11/1 |  | 110119064.D | Sample |  |  |  | 2019-11-02 17:26
| UCG1 T1 S2A809_Cr10/11_SE 11/1 |  | 110119065.D | Sample |  |  |  | 2019-11-02 17:56
| G1+CC CC5_Nov.1 |  | 110119066.D | QC | 5 | 50 | 12.5 | 2019-11-02 18:26
| UCG1 T5 S2A809_Cr10/11_SE 11/1 |  | 110119067.D | Sample |  |  |  | 2019-11-02 18:56
| UCG1 T5 S2B809_Cr10/11_SE 11/1 |  | 110119068.D | Sample |  |  |  | 2019-11-02 19:26
| MTBE |  | 110119069.D | Blank |  |  |  | 2019-11-02 19:56
| QCGCCC4_1015 |  | 110119070.D | QC | 4 | 30 | 7.5 | 2019-11-02 20:26
| UCG1 T5 S2C809_Cr10/11_SE 11/1 |  | 110119071.D | Sample |  |  |  | 2019-11-02 20:56
| UCG1 AF S2A809_Cr10/11_SE 11/1 |  | 110119072.D | Sample |  |  |  | 2019-11-02 21:26
| UCG1 AF S2B809_Cr10/11_SE 11/1 |  | 110119073.D | Sample |  |  |  | 2019-11-02 21:56
| UCG1 AF S2C809_Cr10/11_SE 11/1 |  | 110119074.D | Sample |  |  |  | 2019-11-02 22:27
| UCG1 T1 S3A809_Cr10/11_SE 11/1 |  | 110119075.D | Sample |  |  |  | 2019-11-02 22:57
| G1+CC CC9_Nov.1 |  | 110119076.D | QC | 9 | 350 | 87.5 | 2019-11-02 23:27
| UCG1 T5 S3A809_Cr10/11_SE 11/1 |  | 110119077.D | Sample |  |  |  | 2019-11-02 23:57
| UCG1 T5 S3B809_Cr10/11_SE 11/1 |  | 110119078.D | Sample |  |  |  | 2019-11-03 00:27
| UCG1 T5 S3C809_Cr10/11_SE 11/1 |  | 110119079.D | Sample |  |  |  | 2019-11-03 00:57
| MTBE |  | 110119080.D | Blank |  |  |  | 2019-11-03 01:27
| UCG1 AF S3A809_Cr10/11_SE 11/1 |  | 110119081.D | Sample |  |  |  | 2019-11-03 01:57
| UCG1 AF S3B809_Cr10/11_SE 11/1 |  | 110119082.D | Sample |  |  |  | 2019-11-03 02:02
| UCG1 AF S3C809_Cr10/11_SE 11/1 |  | 110119083.D | Sample |  |  |  | 2019-11-03 02:32
| UCG1 T5 S1B1009_SE 11/1b |  | 110119084.D | Sample |  |  |  | 2019-11-03 03:02
| UCG1 AF S2C1009_SE 11/1b |  | 110119085.D | Sample |  |  |  | 2019-11-03 03:32
| UCG1 T1 S3A1009_SE 11/1b |  | 110119086.D | Sample |  |  |  | 2019-11-03 04:02
| G1+CC CC4_Nov.1 |  | 110119087.D | QC | 4 | 30 | 7.5 | 2019-11-03 04:32
| MTBE |  | 110119088.D | Blank |  |  |  | 2019-11-03 05:02
2019-12-10 | [8 empty] | Total # Samples: | 47
Sample | [8 empty] | Actual # Samples: | 30
| Name | Comment | Data File | Type | Level | Conc. pre-Crash (nM) | Conc. after Crash (nM) | Acq. Date-Time
| Spike Blank |  | 121019000.D | MatrixBlank |  |  |  | 2019-12-10 11:49
| CC1 | 923 dropped | 121019001.D | Cal | 1 | 7 | 1.75 | 2019-12-10 12:19
| CC2 | 924 dropped | 121019002.D | Cal | 2 | 12 | 3 | 2019-12-10 12:49
| CC3 | 925 dropped | 121019003.D | Cal | 3 | 20 | 5 | 2019-12-10 13:19
| CC4 | 926 dropped | 121019004.D | Cal | 4 | 30 | 7.5 | 2019-12-10 13:49
| CC5 |  | 121019005.D | Cal | 5 | 50 | 12.5 | 2019-12-10 14:20
| CC6 |  | 121019006.D | Cal | 6 | 80 | 20 | 2019-12-10 14:50
| CC7 |  | 121019007.D | Cal | 7 | 125 | 31.25 | 2019-12-10 15:20
| CC8 |  | 121019008.D | Cal | 8 | 200 | 50 | 2019-12-10 15:50
| CC9 | Dropped | 121019009.D | Cal | 9 | 350 | 87.5 | 2019-12-10 16:20
| CC10 |  | 121019010.D | Cal | 10 | 500 | 125 | 2019-12-10 16:50
| CC11 |  | 121019011.D | Cal | 11 | 800 | 200 | 2019-12-10 17:20
| CC12 | Dropped | 121019012.D | Cal | 12 | 1500 | 375 | 2019-12-10 17:50
| CC13 | Dropped | 121019013.D | Cal | 13 | 2500 | 625 | 2019-12-10 18:20
| CC14 |  | 121019014.D | Cal | 14 | 3500 | 875 | 2019-12-10 18:50
| CC15 |  | 121019015.D | Cal | 15 | 5000 | 1250 | 2019-12-10 19:20
| QCGC4 |  | 121019016.D | QC | 4 | 30 | 7.5 | 2019-12-10 19:51
| QCGC8 |  | 121019017.D | QC | 8 | 200 | 50 | 2019-12-10 20:21
| QCGC11 |  | 121019018.D | QC | 11 | 800 | 200 | 2019-12-10 20:51
| UC1206_S3bT1 |  | 121019019.D | Sample |  |  |  | 2019-12-10 21:21
| UC1206_S3cAF |  | 121019020.D | Sample |  |  |  | 2019-12-10 21:51
| Spike Blank |  | 121019021.D | MatrixBlank |  |  |  | 2019-12-10 22:21
| UC1206_S3aAF |  | 121019023.D | Sample |  |  |  | 2019-12-10 23:21
| CC7 |  | 121019025.D | QC | 7 | 125 | 31.25 | 2019-12-11 00:21
| UC1206_S3aT1 |  | 121019026.D | Sample |  |  |  | 2019-12-11 00:51
| UC1206_S3aT1_dup |  | 121019027.D | Sample |  |  |  | 2019-12-11 01:21
| Spike Blank |  | 121019028.D | MatrixBlank |  |  |  | 2019-12-11 01:52
| UC1206_S3bAF |  | 121019031.D | Sample |  |  |  | 2019-12-11 03:22
| Spike Blank |  | 121019033.D | MatrixBlank |  |  |  | 2019-12-11 04:22
| CC9 | Cal dropped | 121019036.D | QC | 9 | 350 | 87.5 | 2019-12-11 05:52
| Spike Blank |  | 121019040.D | MatrixBlank |  |  |  | 2019-12-11 07:52
| UC1206_S3aT5 |  | 121019043.D | Sample |  |  |  | 2019-12-11 08:22
| Spike Blank |  | 121019045.D | MatrixBlank |  |  |  | 2019-12-11 08:52
| UC1206_S3bT5 |  | 121019048.D | Sample |  |  |  | 2019-12-11 09:22
| UC1206_S3bT5_dup |  | 121019049.D | Sample |  |  |  | 2019-12-11 09:52
| UC1206_S3cT1 |  | 121019050.D | Sample |  |  |  | 2019-12-11 10:22
| Spike Blank |  | 121019051.D | MatrixBlank |  |  |  | 2019-12-11 10:52
| UC1206_S3cT5 |  | 121019052.D | Sample |  |  |  | 2019-12-11 11:22
| Spike Blank |  | 121019056.D | MatrixBlank |  |  |  | 2019-12-11 11:52
| CC1 |  | 121019060.D | CC | 1 | 7 | 1.75 | 2019-12-11 12:22
| CC2 |  | 121019061.D | CC | 2 | 12 | 3 | 2019-12-11 12:52
| CC3 |  | 121019062.D | CC | 3 | 20 | 5 | 2019-12-11 13:22
| CC4 |  | 121019063.D | CC | 4 | 30 | 7.5 | 2019-12-11 13:52
| CC5 |  | 121019064.D | CC | 5 | 50 | 12.5 | 2019-12-11 14:22
| CC6 |  | 121019065.D | CC | 6 | 80 | 20 | 2019-12-11 14:52
| CC7 |  | 121019066.D | CC | 7 | 125 | 31.25 | 2019-12-11 15:22
| CC8 |  | 121019067.D | CC | 8 | 200 | 50 | 2019-12-11 15:52
| CC9 |  | 121019068.D | CC | 9 | 350 | 87.5 | 2019-12-11 16:22
| CC10 |  | 121019069.D | CC | 10 | 500 | 125 | 2019-12-11 16:52
| CC11 |  | 121019070.D | CC | 11 | 800 | 200 | 2019-12-11 17:22
| CC12 |  | 121019071.D | CC | 12 | 1500 | 375 | 2019-12-11 17:52
| CC13 |  | 121019072.D | CC | 13 | 2500 | 625 | 2019-12-11 18:22
| CC14 |  | 121019073.D | CC | 14 | 3500 | 875 | 2019-12-11 18:52
| CC15 |  | 121019074.D | CC | 15 | 5000 | 1250 | 2019-12-11 19:23
| QCGC4 |  | 121019075.D | QC | 4 | 30 | 7.5 | 2019-12-12 10:02

## LOD
MDL
References:
[10 empty] | 40 CFR Part 136 | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
[10 empty] | EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016) | [10 empty] | https://www.epa.gov/sites/production/files/2016-12/documents/mdl-procedure_rev2_12-13-2016.pdf
908
Sample | [6 empty] | 908 Method |  | 908 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc.(nM) | Accuracy | Area
| CC1_101119SE | CC1 for MDL | 011419001.D | Cal | 1 | 2020-01-14 10:59 | y = 0.467949 * x | 1.1102e-16 | 9.0356 | 1.7954 | 102.5954 | 3951.2418
| CC1_101119SE | CC1 for MDL | 011419002.D | Cal | 1 | 2020-01-14 11:29 | y = 0.467949 * x | 1.1102e-16 | 9.03 | 1.9579 | 111.8793 | 4501.6133
| CC1_101119SE | CC1 for MDL | 011419003.D | Cal | 1 | 2020-01-14 11:59 | y = 0.467949 * x | 1.1102e-16 | 9.0355 | 1.7127 | 97.8675 | 3941.6991
| CC1_101119SE | CC1 for MDL | 011419004.D | Cal | 1 | 2020-01-14 12:29 | y = 0.467949 * x | 1.1102e-16 | 9.03 | 1.6841 | 96.2355 | 3919.907
| CC1_101119SE | CC1 for MDL | 011419005.D | Cal | 1 | 2020-01-14 12:59 | y = 0.467949 * x | 1.1102e-16 | 9.03 | 1.4924 | 85.2777 | 3737.5161
| CC1_101119SE | CC1 for MDL | 011419006.D | Cal | 1 | 2020-01-14 13:29 | y = 0.467949 * x | 1.1102e-16 | 9.03 | 1.8626 | 106.4352 | 4266.5565
| CC1_101119SE | CC1 for MDL | 011419007.D | Cal | 1 | 2020-01-14 13:59 | y = 0.467949 * x | 1.1102e-16 | 9.0299 | 1.7449 | 99.7094 | 4199.2264
[9 empty] | sd | 0.147
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.46
909
Sample | [6 empty] | 909 Method |  | 909 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_101119SE | CC1 for MDL | 011419001.D | Cal | 1 | 2020-01-14 10:59 | y = 0.480273 * x | 0 | 9.1174 | 1.9773 | 112.9912 | 4466.2091
| CC1_101119SE | CC1 for MDL | 011419002.D | Cal | 1 | 2020-01-14 11:29 | y = 0.480273 * x | 0 | 9.1172 | 1.7774 | 101.5641 | 4194.187
| CC1_101119SE | CC1 for MDL | 011419003.D | Cal | 1 | 2020-01-14 11:59 | y = 0.480273 * x | 0 | 9.1172 | 1.6061 | 91.7769 | 3793.7351
| CC1_101119SE | CC1 for MDL | 011419004.D | Cal | 1 | 2020-01-14 12:29 | y = 0.480273 * x | 0 | 9.1172 | 1.8137 | 103.6414 | 4332.7431
| CC1_101119SE | CC1 for MDL | 011419005.D | Cal | 1 | 2020-01-14 12:59 | y = 0.480273 * x | 0 | 9.1118 | 1.4887 | 85.0662 | 3826.4281
| CC1_101119SE | CC1 for MDL | 011419006.D | Cal | 1 | 2020-01-14 13:29 | y = 0.480273 * x | 0 | 9.1118 | 1.9682 | 112.4668 | 4627.0657
| CC1_101119SE | CC1 for MDL | 011419007.D | Cal | 1 | 2020-01-14 13:59 | y = 0.480273 * x | 0 | 9.1118 | 1.6186 | 92.4934 | 3997.9089
[9 empty] | sd | 0.187
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.59
916
Sample | [6 empty] | 916 Method |  | 916 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_101119SE | CC1 for MDL | 011419001.D | Cal | 1 | 2020-01-14 10:59 | y = 0.391972 * x | 2.2204e-16 | 11.7258 | 1.8282 | 104.4703 | 3370.188
| CC1_101119SE | CC1 for MDL | 011419002.D | Cal | 1 | 2020-01-14 11:29 | y = 0.391972 * x | 2.2204e-16 | 11.7173 | 2.0162 | 115.2117 | 3883.0316
| CC1_101119SE | CC1 for MDL | 011419003.D | Cal | 1 | 2020-01-14 11:59 | y = 0.391972 * x | 2.2204e-16 | 11.7172 | 1.6494 | 94.2528 | 3179.7633
| CC1_101119SE | CC1 for MDL | 011419004.D | Cal | 1 | 2020-01-14 12:29 | y = 0.391972 * x | 2.2204e-16 | 11.7172 | 1.7585 | 100.4863 | 3428.4916
| CC1_101119SE | CC1 for MDL | 011419005.D | Cal | 1 | 2020-01-14 12:59 | y = 0.391972 * x | 2.2204e-16 | 11.7172 | 1.6164 | 92.3629 | 3390.7897
| CC1_101119SE | CC1 for MDL | 011419006.D | Cal | 1 | 2020-01-14 13:29 | y = 0.391972 * x | 2.2204e-16 | 11.7172 | 1.5906 | 90.8906 | 3051.8749
| CC1_101119SE | CC1 for MDL | 011419007.D | Cal | 1 | 2020-01-14 13:59 | y = 0.391972 * x | 2.2204e-16 | 11.7172 | 1.7907 | 102.3253 | 3609.7042
[9 empty] | sd | 0.148
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.47
923
Sample | [6 empty] | 923 Method |  | 923 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC5_101119SE | CC1 for MDL | 011419008.D | Cal | 5 | 2020-01-14 14:29 | y = 0.090450 * x | 0 | 22.8935 | 12.3206 | 98.5651 | 6453.7736
| CC5_101119SE | CC1 for MDL | 011419009.D | Cal | 5 | 2020-01-14 14:59 | y = 0.090450 * x | 0 | 22.8936 | 9.7463 | 77.9703 | 4898.4533
| CC5_101119SE | CC1 for MDL | 011419010.D | Cal | 5 | 2020-01-14 15:29 | y = 0.090450 * x | 0 | 22.902 | 11.7833 | 94.2667 | 5964.4241
| CC5_101119SE | CC1 for MDL | 011419011.D | Cal | 5 | 2020-01-14 16:00 | y = 0.090450 * x | 0 | 22.9105 | 14.7536 | 118.0291 | 7039.0475
| CC5_101119SE | CC1 for MDL | 011419012.D | Cal | 5 | 2020-01-14 16:30 | y = 0.090450 * x | 0 | 22.8935 | 10.9067 | 87.2539 | 5132.3163
| CC5_101119SE | CC1 for MDL | 011419013.D | Cal | 5 | 2020-01-14 17:00 | y = 0.090450 * x | 0 | 22.8851 | 11.0093 | 88.0741 | 5252.7519
| CC5_101119SE | CC1 for MDL | 011419014.D | Cal | 5 | 2020-01-14 17:30 | y = 0.090450 * x | 0 | 22.8855 | 16.9801 | 135.8408 | 7794.7762
[9 empty] | sd | 2.518
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 7.91
3117
Sample | [6 empty] | 3117 Method |  | 3117 Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_101119SE | CC1 for MDL | 011419001.D | Cal | 1 | 2020-01-14 10:59 | y = 0.121410 * x | 0 | 9.5216 | 1.6072 | 91.8378 | 917.6616
| CC1_101119SE | CC1 for MDL | 011419002.D | Cal | 1 | 2020-01-14 11:29 | y = 0.121410 * x | 0 | 9.5216 | 1.5157 | 86.6091 | 904.1446
| CC1_101119SE | CC1 for MDL | 011419003.D | Cal | 1 | 2020-01-14 11:59 | y = 0.121410 * x | 0 | 9.5215 | 1.8594 | 106.2486 | 1110.2578
| CC1_101119SE | CC1 for MDL | 011419004.D | Cal | 1 | 2020-01-14 12:29 | y = 0.121410 * x | 0 | 9.5215 | 1.8766 | 107.2333 | 1133.251
| CC1_101119SE | CC1 for MDL | 011419005.D | Cal | 1 | 2020-01-14 12:59 | y = 0.121410 * x | 0 | 9.5215 | 1.5639 | 89.3663 | 1016.1946
| CC1_101119SE | CC1 for MDL | 011419006.D | Cal | 1 | 2020-01-14 13:29 | y = 0.121410 * x | 0 | 9.516 | 1.791 | 102.345 | 1064.4241
| CC1_101119SE | CC1 for MDL | 011419007.D | Cal | 1 | 2020-01-14 13:59 | y = 0.121410 * x | 0 | 9.516 | 2.0363 | 116.3599 | 1271.4287
[9 empty] | sd | 0.192
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.6
4NT
Sample | [6 empty] | 4NT Method |  | 4NT Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area
| CC1_101119SE | CC1 for MDL | 011419001.D | Cal | 1 | 2020-01-14 10:59 | y = 6.410525 * x | 0 | 10.606 | 1.5836 | 90.4933 | 82332.1628
| CC1_101119SE | CC1 for MDL | 011419002.D | Cal | 1 | 2020-01-14 11:29 | y = 6.410525 * x | 0 | 10.6009 | 1.7746 | 101.4055 | 84924.4687
| CC1_101119SE | CC1 for MDL | 011419003.D | Cal | 1 | 2020-01-14 11:59 | y = 6.410525 * x | 0 | 10.6009 | 1.6501 | 94.2929 | 85514.7787
| CC1_101119SE | CC1 for MDL | 011419004.D | Cal | 1 | 2020-01-14 12:29 | y = 6.410525 * x | 0 | 10.6009 | 1.7341 | 99.0905 | 89538.997
| CC1_101119SE | CC1 for MDL | 011419005.D | Cal | 1 | 2020-01-14 12:59 | y = 6.410525 * x | 0 | 10.6009 | 1.8571 | 106.1192 | 94693.1883
| CC1_101119SE | CC1 for MDL | 011419006.D | Cal | 1 | 2020-01-14 13:29 | y = 6.410525 * x | 0 | 10.6009 | 1.7276 | 98.7213 | 91064.0321
| CC1_101119SE | CC1 for MDL | 011419007.D | Cal | 1 | 2020-01-14 13:59 | y = 6.410525 * x | 0 | 10.6009 | 1.9229 | 109.8772 | 109228.731
[9 empty] | sd | 0.116
Student's t-value at 6 degrees of freedom (0.99 confidence interval) |  |  |  | 3.143 |  |  |  |  | MDL (nM) | 0.36
eLOQ | the lowest curve point run that was w/in 30% accuracy, where all other curve points are also w/in 30% accuracy
2019-10-29
Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results | [9 empty] | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | [8 empty] | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| UCG1CC1_8/13/2019_SE 10/15 |  | 102919001.D | Cal | 1 | 2019-10-29 11:34 | y = 0.074827 * x | 0.99397 | 9.1395 | 1.5924 | 90.9928 | 3427.028 | 13.6808 | 719040.0903 | y = 0.068145 * x | 0.99536 | 9.2158 | 1.5547 | 88.8401 | 3047.1492 | 13.6808 | 719040.0903 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 1.6106 | 92.0357 | 1006.3551 | 13.6808 | 719040.0903 | [8 empty] | y = 0.064060 * x | 0.99408 | 11.8365 | 2.627 | 150.1148 | 4840.1934 | 13.6808 | 719040.0903 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5561 | 0 | 0 | 0 | 13.6808 | 719040.0903
| UCG1CC2_8/13/2019_SE 10/15 |  | 102919002.D | Cal | 2 | 2019-10-29 12:04 | y = 0.074827 * x | 0.99397 | 9.1396 | 2.6019 | 86.7284 | 6708.7658 | 13.6804 | 861470.2526 | y = 0.068145 * x | 0.99536 | 9.2159 | 2.9372 | 97.9083 | 6897.2286 | 13.6804 | 861470.2526 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 3.5224 | 117.4137 | 2637.3779 | 13.6804 | 861470.2526 | [8 empty] | y = 0.064060 * x | 0.99408 | 11.8361 | 3.7263 | 124.2112 | 8225.6506 | 13.6804 | 861470.2526 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.6236 | 0 | 0 | 0 | 13.6804 | 861470.2526
| UCG1CC3_8/13/2019_SE 10/15 |  | 102919003.D | Cal | 3 | 2019-10-29 12:34 | y = 0.074827 * x | 0.99397 | 9.1396 | 5.0772 | 101.5446 | 15064.615 | 13.6804 | 991314.3628 | y = 0.068145 * x | 0.99536 | 9.2159 | 5.2094 | 104.1878 | 14076.4092 | 13.6804 | 991314.3628 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 5.9388 | 118.7764 | 5118.1754 | 13.6804 | 991314.3628 | [8 empty] | y = 0.064060 * x | 0.99408 | 11.8361 | 5.3748 | 107.4966 | 13652.8796 | 13.6804 | 991314.3628 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5557 | 6.1804 | 123.609 | 2085.8999 | 13.6804 | 991314.3628
2019-11-01
Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| G1+CC CC3_Nov.1 |  | 110119015.D | Cal | 3 | 2019-11-01 16:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 4.8963 | 97.925 | 14175.5513 | 13.6721 | 738431.6469 | y = 0.092110 * x | 0.99679 | 9.1996 | 4.5115 | 90.2299 | 12274.3495 | 13.6721 | 738431.6469 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 4.0411 | 80.8225 | 4004.101 | 13.6721 | 738431.6469 | y = 0.090306 * x | 0.99908 | 10.6917 | 17.4393 | 348.7858 | 89852.4773 | 10.6858 | 171162.0981 | y = 0.079567 * x | 0.99598 | 11.8193 | 4.9071 | 98.1418 | 11532.595 | 13.6721 | 738431.6469 | y = 0.012003 * x | 0.99495 | 23.4796 | 3.6383 | 72.7668 | 1289.9729 | 13.6721 | 738431.6469
| G1+CC CC2_Nov.1 |  | 110119016.D | Cal | 2 | 2019-11-01 17:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 3.3552 | 111.8388 | 8868.7775 | 13.6721 | 674192.9048 | y = 0.092110 * x | 0.99679 | 9.1996 | 3.0144 | 100.4785 | 7487.6605 | 13.6721 | 674192.9048 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 2.1169 | 70.5647 | 1915.3903 | 13.6721 | 674192.9048 | y = 0.090306 * x | 0.99908 | 10.6917 | 13.6926 | 456.4201 | 70888.9583 | 10.6858 | 171988.3186 | y = 0.079567 * x | 0.99598 | 11.8193 | 2.4418 | 81.395 | 5239.5726 | 13.6721 | 674192.9048 | y = 0.012003 * x | 0.99495 | 23.4626 | 3.0292 | 100.9717 | 980.5551 | 13.6721 | 674192.9048
| G1+CC CC1_Nov.1 |  | 110119017.D | Cal | 1 | 2019-11-01 17:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 2.2413 | 128.075 | 4287.5537 | 13.6721 | 487911.5221 | y = 0.092110 * x | 0.99679 | 9.1996 | 2.7254 | 155.736 | 4899.3174 | 13.6721 | 487911.5221 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 1.3614 | 77.7966 | 891.5218 | 13.6721 | 487911.5221 | y = 0.090306 * x | 0.99908 | 10.6917 | 11.1329 | 636.166 | 43606.9808 | 10.6858 | 130122.9376 | y = 0.079567 * x | 0.99598 | 11.8278 | 2.3361 | 133.4899 | 3627.6082 | 13.6721 | 487911.5221 | y = 0.012003 * x | 0.99495 | 23.4541 | 0 | 0 | 0 | 13.6721 | 487911.5221

## CC,eLOQ
Flag for outside eLOQ criteria | >130% | <70% | 10/29/2019 CC
|  |  | Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results |  | 4-Nitrotoluene Method |  | 4-Nitrotoluene Results |  |  |  | 4NT13C6 (ISTD) Results
Compound | eLOQ (nM) |  |  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
908 | 1.75 |  |  | UCG1CC1_8/13/2019_SE 10/15 |  | 102919001.D | Cal | 1 | 2019-10-29 11:34 | y = 0.074827 * x | 0.99397 | 9.1395 | 1.5924 | 90.9928 | 3427.028 | 13.6808 | 719040.0903 | y = 0.068145 * x | 0.99536 | 9.2158 | 1.5547 | 88.8401 | 3047.1492 | 13.6808 | 719040.0903 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 1.6106 | 92.0357 | 1006.3551 | 13.6808 | 719040.0903 | y = 0.064060 * x | 0.99408 | 11.8365 | 2.627 | 150.1148 | 4840.1934 | 13.6808 | 719040.0903 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5561 | 0 | 0 | 0 | 13.6808 | 719040.0903
909 | 1.75 |  |  | UCG1CC2_8/13/2019_SE 10/15 |  | 102919002.D | Cal | 2 | 2019-10-29 12:04 | y = 0.074827 * x | 0.99397 | 9.1396 | 2.6019 | 86.7284 | 6708.7658 | 13.6804 | 861470.2526 | y = 0.068145 * x | 0.99536 | 9.2159 | 2.9372 | 97.9083 | 6897.2286 | 13.6804 | 861470.2526 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 3.5224 | 117.4137 | 2637.3779 | 13.6804 | 861470.2526 | y = 0.064060 * x | 0.99408 | 11.8361 | 3.7263 | 124.2112 | 8225.6506 | 13.6804 | 861470.2526 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.6236 | 0 | 0 | 0 | 13.6804 | 861470.2526
916 | 3; 1.75 on 11/1 CC |  |  | UCG1CC3_8/13/2019_SE 10/15 |  | 102919003.D | Cal | 3 | 2019-10-29 12:34 | y = 0.074827 * x | 0.99397 | 9.1396 | 5.0772 | 101.5446 | 15064.615 | 13.6804 | 991314.3628 | y = 0.068145 * x | 0.99536 | 9.2159 | 5.2094 | 104.1878 | 14076.4092 | 13.6804 | 991314.3628 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 5.9388 | 118.7764 | 5118.1754 | 13.6804 | 991314.3628 | y = 0.064060 * x | 0.99408 | 11.8361 | 5.3748 | 107.4966 | 13652.8796 | 13.6804 | 991314.3628 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5557 | 6.1804 | 123.609 | 2085.8999 | 13.6804 | 991314.3628
923 | 5; 12.5 on 12/10 CC |  |  | UCG1CC4_8/13/2019_SE 10/15 |  | 102919004.D | Cal | 4 | 2019-10-29 13:04 | y = 0.074827 * x | 0.99397 | 9.1396 | 7.5916 | 101.2207 | 16745.6357 | 13.6804 | 736972.4781 | y = 0.068145 * x | 0.99536 | 9.2159 | 7.1453 | 95.2712 | 14353.8297 | 13.6804 | 736972.4781 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 7.3263 | 97.6843 | 4694.6701 | 13.6804 | 736972.4781 | y = 0.064060 * x | 0.99408 | 11.8361 | 6.7568 | 90.0909 | 12759.7267 | 13.6804 | 736972.4781 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 7.6424 | 101.8983 | 2255.5552 | 13.6804 | 736972.4781
3117 | 1.75; 3 on 1st 12/10 CC |  |  | UCG1CC5_8/13/2019_SE 10/15 | dropped | 102919005.D | Cal | 5 | 2019-10-29 13:34 | y = 0.074827 * x | 0.99397 | 9.1396 | 17.154 | 137.2317 | 31639.0042 | 13.6804 | 616224.2482 | y = 0.068145 * x | 0.99536 | 9.2159 | 14.9678 | 119.7425 | 25141.4687 | 13.6804 | 616224.2482 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 17.4023 | 139.218 | 9334.2028 | 13.6804 | 616224.2482 | y = 0.064060 * x | 0.99408 | 11.8361 | 14.0524 | 112.4189 | 22188.9271 | 13.6804 | 616224.2482 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5388 | 13.8142 | 110.5136 | 4371.1887 | 13.6804 | 616224.2482 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 11.0092 | 88.074 | 85922.8028 | 10.6858 | 142139.58
4NT | 12.5nM 10/29 & 11/1; 1.75 12/10 |  |  | UCG1CC6_8/13/2019_SE 10/15 | dropped | 102919006.D | Cal | 6 | 2019-10-29 14:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 24.4558 | 122.2789 | 65876.2888 | 13.6804 | 899969.5806 | y = 0.068145 * x | 0.99536 | 9.2159 | 25.0831 | 125.4155 | 61532.2151 | 13.6804 | 899969.5806 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 26.6824 | 133.4118 | 20922.4723 | 13.6804 | 899969.5806 | y = 0.064060 * x | 0.99408 | 11.8361 | 23.4916 | 117.4581 | 54173.7659 | 13.6804 | 899969.5806 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5643 | 25.8472 | 129.2359 | 13440.6263 | 13.6804 | 899969.5806 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 21.3217 | 106.6086 | 129222.1642 | 10.6858 | 141646.0342
|  |  |  | UCG1CC7_8/13/2019_SE 10/15 |  | 102919007.D | Cal | 7 | 2019-10-29 14:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 27.0387 | 86.524 | 71463.7088 | 13.6804 | 883037.9864 | y = 0.068145 * x | 0.99536 | 9.2159 | 25.9704 | 83.1051 | 62510.1702 | 13.6804 | 883037.9864 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 29.2994 | 93.7581 | 22548.6026 | 13.6804 | 883037.9864 | y = 0.064060 * x | 0.99408 | 11.8361 | 27.6087 | 88.3478 | 62470.2555 | 13.6804 | 883037.9864 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5642 | 34.6613 | 110.9163 | 18243.3718 | 13.6804 | 883037.9864 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 34.431 | 110.1792 | 217454.3475 | 10.6857 | 166815.9327
LOQ |  |  |  | UCG1CC8_8/13/2019_SE 10/15 |  | 102919008.D | Cal | 8 | 2019-10-29 15:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 41.2898 | 82.5796 | 113572.8108 | 13.6804 | 918992.5392 | y = 0.068145 * x | 0.99536 | 9.2159 | 43.4854 | 86.9707 | 108930.2888 | 13.6804 | 918992.5392 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 42.3932 | 84.7863 | 34001.0384 | 13.6804 | 918992.5392 | y = 0.064060 * x | 0.99408 | 11.8361 | 41.6535 | 83.307 | 98087.0024 | 13.6804 | 918992.5392 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5557 | 35.7415 | 71.4829 | 19629.9628 | 13.6804 | 918992.5392 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 48.5082 | 97.0164 | 310387.6602 | 10.6857 | 180068.8012
|  |  |  | UCG1CC9_8/13/2019_SE 10/15 |  | 102919009.D | Cal | 9 | 2019-10-29 15:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 91.8645 | 104.988 | 192073.8962 | 13.6804 | 698555.8284 | y = 0.068145 * x | 0.99536 | 9.2159 | 91.5888 | 104.6729 | 174396.1774 | 13.6804 | 698555.8284 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 91.3828 | 104.4375 | 56001.3511 | 13.6804 | 698555.8284 | y = 0.064060 * x | 0.99408 | 11.8361 | 87.8795 | 100.4337 | 157302.9047 | 13.6804 | 698555.8284 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 70.4223 | 80.4826 | 30547.5546 | 13.6804 | 698555.8284 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6865 | 86.2481 | 98.5692 | 482156.3722 | 10.6857 | 169170.1173
|  |  |  | UCG1CC10_8/13/2019_SE 10/15 |  | 102919010.D | Cal | 10 | 2019-10-29 16:05 | y = 0.074827 * x | 0.99397 | 9.1397 | 129.3084 | 103.4467 | 387525.4519 | 13.6805 | 1001276.0262 | y = 0.068145 * x | 0.99536 | 9.216 | 130.6627 | 104.5302 | 356614.4799 | 13.6805 | 1001276.0262 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 127.1415 | 101.7132 | 112100.6642 | 13.6805 | 1001276.0262 | y = 0.064060 * x | 0.99408 | 11.8362 | 126.4875 | 101.19 | 324526.1277 | 13.6805 | 1001276.0262 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5474 | 138.7736 | 111.0189 | 87226.0027 | 13.6805 | 1001276.0262 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 123.0436 | 98.4348 | 529585.9447 | 10.6858 | 134128.3581
|  |  |  | UCG1CC11_8/13/2019_SE 10/15 |  | 102919011.D | Cal | 11 | 2019-10-29 16:35 | y = 0.074827 * x | 0.99397 | 9.1395 | 190.859 | 95.4295 | 331860.3707 | 13.6808 | 580929.1207 | y = 0.068145 * x | 0.99536 | 9.2158 | 192.1874 | 96.0937 | 304327.731 | 13.6808 | 580929.1207 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 191.9128 | 95.9564 | 98841.233 | 13.6808 | 580929.1207 | y = 0.064060 * x | 0.99408 | 11.8365 | 168.8277 | 84.4139 | 251313.047 | 13.6808 | 580929.1207 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5477 | 201.4772 | 100.7386 | 73253.2349 | 13.6808 | 580929.1207 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 198.7426 | 99.3713 | 846983.6499 | 10.6858 | 136440.3565
|  |  |  | UCG1CC12_8/13/2019_SE 10/15 |  | 102919012.D | Cal | 12 | 2019-10-29 17:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 341.4471 | 91.0526 | 773830.0929 | 13.6804 | 757185.9152 | y = 0.068145 * x | 0.99536 | 9.2159 | 345.0649 | 92.0173 | 712191.4673 | 13.6804 | 757185.9152 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 379.9858 | 101.3296 | 260086.7356 | 13.6804 | 757185.9152 | y = 0.064060 * x | 0.99408 | 11.8361 | 345.6993 | 92.1865 | 670732.1088 | 13.6804 | 757185.9152 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5388 | 375.297 | 100.0792 | 174199.8837 | 13.6804 | 757185.9152 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 377.2542 | 100.6011 | 1646726.6012 | 10.6858 | 142750.4376
|  |  |  | UCG1CC13_8/13/2019_SE 10/15 |  | 102919013.D | Cal | 13 | 2019-10-29 17:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 541.171 | 86.5874 | 1479118.0833 | 13.6804 | 913163.8009 | y = 0.068145 * x | 0.99536 | 9.2159 | 570.7341 | 91.3175 | 1420613.7996 | 13.6804 | 913163.8009 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 545.3287 | 87.2526 | 457762.0072 | 13.6804 | 913163.8009 | y = 0.064060 * x | 0.99408 | 11.8361 | 539.3824 | 86.3012 | 1262099.4482 | 13.6804 | 913163.8009 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 445.0657 | 71.2105 | 246645.047 | 13.6804 | 913163.8009 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 632.5709 | 101.2113 | 3934261.086 | 10.6857 | 205379.3612
|  |  |  | UCG1CC14_8/13/2019_SE 10/15 |  | 102919014.D | Cal | 14 | 2019-10-29 18:05 | y = 0.074827 * x | 0.99397 | 9.1397 | 937.9818 | 107.1979 | 1834222.6826 | 13.6805 | 653338.0475 | y = 0.068145 * x | 0.99536 | 9.216 | 947.5996 | 108.2971 | 1687548.6307 | 13.6805 | 653338.0475 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 873.9171 | 99.8762 | 542206.2546 | 13.6805 | 653338.0475 | y = 0.064060 * x | 0.99408 | 11.8362 | 926.0159 | 105.8304 | 1550259.8615 | 13.6805 | 653338.0475 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5389 | 943.6588 | 107.8467 | 344956.4835 | 13.6805 | 653338.0475 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 889.1237 | 101.6141 | 5205424.3285 | 10.6857 | 194135.7217
|  |  |  | UCG1CC15_8/13/2019_SE 10/15 |  | 102919015.D | Cal | 15 | 2019-10-29 18:35 | y = 0.074827 * x | 0.99397 | 9.1397 | 1258.9555 | 100.7164 | 2293484.3484 | 13.6805 | 608647.3341 | y = 0.068145 * x | 0.99536 | 9.216 | 1235.9796 | 98.8784 | 2050550.688 | 13.6805 | 608647.3341 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 1250.2608 | 100.0209 | 749124.555 | 13.6805 | 608647.3341 | y = 0.064060 * x | 0.99408 | 11.8362 | 1271.0537 | 101.6843 | 1982338.1001 | 13.6805 | 608647.3341 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5389 | 1249.8533 | 99.9883 | 402813.9981 | 13.6805 | 608647.3341 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 1228.9968 | 98.3197 | 7127691.7007 | 10.6907 | 192862.2567
11/1/2019 CC
|  |  | Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
|  |  |  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
|  |  |  | G1+CC CC15_Nov.1 |  | 110119003.D | Cal | 15 | 2019-11-01 10:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 1282.0888 | 102.5671 | 3088291.3528 | 13.6805 | 614375.897 | y = 0.092110 * x | 0.99679 | 9.2049 | 1291.9138 | 103.3531 | 2924389.391 | 13.6805 | 614375.897 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 1241.5357 | 99.3229 | 914479.021 | 13.6805 | 614375.897 | y = 0.079567 * x | 0.99598 | 11.8276 | 1308.9286 | 104.7143 | 2559428.3304 | 13.6805 | 614375.897 | y = 0.012003 * x | 0.99495 | 23.4964 | 1269.2163 | 101.5373 | 374400.6196 | 13.6805 | 614375.897 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 1228.9968 | 98.3197 | 7127691.7007 | 10.6907 | 192862.2567
|  |  |  | G1+CC CC14_Nov.1 |  | 110119004.D | Cal | 14 | 2019-11-01 11:22 | y = 0.098018 * x | 0.99826 | 9.1287 | 858.4608 | 98.1098 | 2367255.0818 | 13.672 | 703329.149 | y = 0.092110 * x | 0.99679 | 9.205 | 869.3948 | 99.3594 | 2252906.3189 | 13.672 | 703329.149 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 889.9314 | 101.7064 | 775821.5446 | 13.672 | 703329.149 | y = 0.079567 * x | 0.99598 | 11.8276 | 828.827 | 94.7231 | 1855305.0945 | 13.672 | 703329.149 | y = 0.012003 * x | 0.99495 | 23.4879 | 830.1339 | 94.8724 | 280332.5333 | 13.672 | 703329.149 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 889.1237 | 101.6141 | 5205424.3285 | 10.6857 | 194135.7217
|  |  |  | G1+CC CC13_Nov.1 |  | 110119005.D | Cal | 13 | 2019-11-01 11:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 578.8779 | 92.6205 | 1830704.2438 | 13.672 | 806612.9952 | y = 0.092110 * x | 0.99679 | 9.2049 | 556.9159 | 89.1065 | 1655092.6592 | 13.672 | 806612.9952 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 557.7707 | 89.2433 | 574916.8592 | 13.672 | 806612.9952 | y = 0.079567 * x | 0.99598 | 11.8276 | 573.6178 | 91.7788 | 1472585.919 | 13.672 | 806612.9952 | y = 0.012003 * x | 0.99495 | 23.4964 | 677.8209 | 108.4513 | 262510.6297 | 13.672 | 806612.9952 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 632.5709 | 101.2113 | 3934261.086 | 10.6857 | 205379.3612
|  |  |  | G1+CC CC12_Nov.1 |  | 110119006.D | Cal | 12 | 2019-11-01 12:22 | y = 0.098018 * x | 0.99826 | 9.1287 | 378.0035 | 100.8009 | 744473.877 | 13.6806 | 502328.3423 | y = 0.092110 * x | 0.99679 | 9.205 | 363.6109 | 96.9629 | 672964.2874 | 13.6806 | 502328.3423 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5819 | 387.7014 | 103.387 | 252693.3098 | 13.6806 | 502328.3423 | y = 0.079567 * x | 0.99598 | 11.8277 | 370.8606 | 98.8962 | 592913.3437 | 13.6806 | 502328.3423 | y = 0.012003 * x | 0.99495 | 23.4965 | 315.4637 | 84.1236 | 76085.8268 | 13.6806 | 502328.3423 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 377.2542 | 100.6011 | 1646726.6012 | 10.6858 | 142750.4376
|  |  |  | G1+CC CC11_Nov.1 |  | 110119007.D | Cal | 11 | 2019-11-01 12:52 | y = 0.098018 * x | 0.99826 | 9.1288 | 205.9028 | 102.9514 | 396309.4317 | 13.6806 | 490914.9508 | y = 0.092110 * x | 0.99679 | 9.2051 | 198.5623 | 99.2812 | 359145.6143 | 13.6806 | 490914.9508 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 187.8395 | 93.9198 | 121775.5564 | 13.6806 | 490914.9508 | y = 0.079567 * x | 0.99598 | 11.8278 | 202.8357 | 101.4179 | 316915.5223 | 13.6806 | 490914.9508 | y = 0.012003 * x | 0.99495 | 23.4796 | 234.5584 | 117.2792 | 55287.1183 | 13.6806 | 490914.9508 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 198.7426 | 99.3713 | 846983.6499 | 10.6858 | 136440.3565
|  |  |  | G1+CC CC10_Nov.1 |  | 110119008.D | Cal | 10 | 2019-11-01 13:22 | y = 0.098018 * x | 0.99826 | 9.1288 | 132.6477 | 106.1182 | 242497.257 | 13.6721 | 466273.8622 | y = 0.092110 * x | 0.99679 | 9.205 | 127.9846 | 102.3877 | 219870.2161 | 13.6721 | 466273.8622 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 99.9609 | 79.9687 | 62024.455 | 13.6721 | 466273.8622 | y = 0.079567 * x | 0.99598 | 11.8277 | 126.1178 | 100.8942 | 187158.8389 | 13.6721 | 466273.8622 | y = 0.012003 * x | 0.99495 | 23.488 | 107.3064 | 85.8451 | 24023.3393 | 13.6721 | 466273.8622 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 123.0436 | 98.4348 | 529585.9447 | 10.6858 | 134128.3581
|  |  |  | G1+CC CC9_Nov.1 |  | 110119009.D | Cal | 9 | 2019-11-01 13:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 88.3173 | 100.934 | 229080.1853 | 13.672 | 661570.0892 | y = 0.092110 * x | 0.99679 | 9.2049 | 82.9002 | 94.7431 | 202068.6875 | 13.672 | 661570.0892 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 74.1712 | 84.7671 | 65444.6188 | 13.672 | 661570.0892 | y = 0.079567 * x | 0.99598 | 11.8276 | 85.6722 | 97.9111 | 180388.4598 | 13.672 | 661570.0892 | y = 0.012003 * x | 0.99495 | 23.4794 | 91.7192 | 104.8219 | 29134.1824 | 13.672 | 661570.0892 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6865 | 86.2481 | 98.5692 | 482156.3722 | 10.6857 | 169170.1173
|  |  |  | G1+CC CC8_Nov.1 |  | 110119010.D | Cal | 8 | 2019-11-01 14:22 | y = 0.098018 * x | 0.99826 | 9.1232 | 47.6901 | 95.3802 | 124575.268 | 13.672 | 666250.3013 | y = 0.092110 * x | 0.99679 | 9.205 | 46.2056 | 92.4113 | 113422.6796 | 13.672 | 666250.3013 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5819 | 47.6566 | 95.3132 | 42444.3011 | 13.672 | 666250.3013 | y = 0.079567 * x | 0.99598 | 11.8277 | 49.9394 | 99.8788 | 105894.5384 | 13.672 | 666250.3013 | y = 0.012003 * x | 0.99495 | 23.4794 | 44.9861 | 89.9721 | 14390.7049 | 13.672 | 666250.3013 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 48.5082 | 97.0164 | 310387.6602 | 10.6857 | 180068.8012
|  |  |  | G1+CC CC7_Nov.1 |  | 110119011.D | Cal | 7 | 2019-11-01 14:52 | y = 0.098018 * x | 0.99826 | 9.1287 | 29.5408 | 94.5306 | 73952.892 | 13.672 | 638508.322 | y = 0.092110 * x | 0.99679 | 9.205 | 27.5792 | 88.2535 | 64880.8037 | 13.672 | 638508.322 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 20.0521 | 64.1666 | 17156.116 | 13.672 | 638508.322 | y = 0.079567 * x | 0.99598 | 11.8276 | 28.3697 | 90.783 | 57651.9656 | 13.672 | 638508.322 | y = 0.012003 * x | 0.99495 | 23.488 | 24.4713 | 78.3081 | 7502.2233 | 13.672 | 638508.322 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 34.431 | 110.1792 | 217454.3475 | 10.6857 | 166815.9327
|  |  |  | G1+CC CC6_Nov.1 |  | 110119012.D | Cal | 6 | 2019-11-01 15:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 19.4559 | 97.2797 | 39682.6904 | 13.6721 | 520215.186 | y = 0.092110 * x | 0.99679 | 9.205 | 18.1166 | 90.5828 | 34723.7034 | 13.6721 | 520215.186 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 12.3925 | 61.9627 | 8644.1693 | 13.6721 | 520215.186 | y = 0.079567 * x | 0.99598 | 11.8277 | 18.2282 | 91.1411 | 30180.0501 | 13.6721 | 520215.186 | y = 0.012003 * x | 0.99495 | 23.4626 | 14.8962 | 74.4811 | 3720.7116 | 13.6721 | 520215.186 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 21.3217 | 106.6086 | 129222.1642 | 10.6858 | 141646.0342
|  |  |  | G1+CC CC5_Nov.1 |  | 110119013.D | Cal | 5 | 2019-11-01 15:52 | y = 0.098018 * x | 0.99826 | 9.1288 | 13.9431 | 111.5448 | 26946.7662 | 13.6721 | 492925.4119 | y = 0.092110 * x | 0.99679 | 9.1996 | 12.5319 | 100.2556 | 22759.7392 | 13.6721 | 492925.4119 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 11.2377 | 89.9013 | 7428.144 | 13.6721 | 492925.4119 | y = 0.079567 * x | 0.99598 | 11.8193 | 12.1024 | 96.8192 | 18986.5356 | 13.6721 | 492925.4119 | y = 0.012003 * x | 0.99495 | 23.4626 | 15.9709 | 127.7676 | 3779.8855 | 13.6721 | 492925.4119 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 11.0092 | 88.074 | 85922.8028 | 10.6858 | 142139.58
|  |  |  | G1+CC CC4_Nov.1 |  | 110119014.D | Cal | 4 | 2019-11-01 16:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 9.2407 | 123.2093 | 19541.8442 | 13.6721 | 539380.1551 | y = 0.092110 * x | 0.99679 | 9.1996 | 6.8546 | 91.3947 | 13622.1159 | 13.6721 | 539380.1551 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 6.3326 | 84.4349 | 4582.3136 | 13.6721 | 539380.1551 | y = 0.079567 * x | 0.99598 | 11.8192 | 6.7585 | 90.1139 | 11602.2092 | 13.6721 | 539380.1551 | y = 0.012003 * x | 0.99495 | 23.4371 | 9.1581 | 122.1077 | 2371.7349 | 13.6721 | 539380.1551
|  |  |  | G1+CC CC3_Nov.1 |  | 110119015.D | Cal | 3 | 2019-11-01 16:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 4.8963 | 97.925 | 14175.5513 | 13.6721 | 738431.6469 | y = 0.092110 * x | 0.99679 | 9.1996 | 4.5115 | 90.2299 | 12274.3495 | 13.6721 | 738431.6469 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 4.0411 | 80.8225 | 4004.101 | 13.6721 | 738431.6469 | y = 0.079567 * x | 0.99598 | 11.8193 | 4.9071 | 98.1418 | 11532.595 | 13.6721 | 738431.6469 | y = 0.012003 * x | 0.99495 | 23.4796 | 3.6383 | 72.7668 | 1289.9729 | 13.6721 | 738431.6469
|  |  |  | G1+CC CC2_Nov.1 |  | 110119016.D | Cal | 2 | 2019-11-01 17:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 3.3552 | 111.8388 | 8868.7775 | 13.6721 | 674192.9048 | y = 0.092110 * x | 0.99679 | 9.1996 | 3.0144 | 100.4785 | 7487.6605 | 13.6721 | 674192.9048 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 2.1169 | 70.5647 | 1915.3903 | 13.6721 | 674192.9048 | y = 0.079567 * x | 0.99598 | 11.8193 | 2.4418 | 81.395 | 5239.5726 | 13.6721 | 674192.9048 | y = 0.012003 * x | 0.99495 | 23.4626 | 3.0292 | 100.9717 | 980.5551 | 13.6721 | 674192.9048
|  |  |  | G1+CC CC1_Nov.1 |  | 110119017.D | Cal | 1 | 2019-11-01 17:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 2.2413 | 128.075 | 4287.5537 | 13.6721 | 487911.5221 | y = 0.092110 * x | 0.99679 | 9.1996 | 2.7254 | 155.736 | 4899.3174 | 13.6721 | 487911.5221 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 1.3614 | 77.7966 | 891.5218 | 13.6721 | 487911.5221 | y = 0.079567 * x | 0.99598 | 11.8278 | 2.3361 | 133.4899 | 3627.6082 | 13.6721 | 487911.5221 | y = 0.012003 * x | 0.99495 | 23.4541 | 0 | 0 | 0 | 13.6721 | 487911.5221
12/10/2019 CC's
|  |  | Sample | [22 empty] | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results
|  |  |  | Name | Comment | Data File | Type | Level | Acq. Date-Time | [16 empty] | CF Formula | CF R2 | RT | Final Conc. (nM) | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. (nM) | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. (nM) | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. (nM) | Accuracy | Area | RT | Area
|  |  | Initial CC | CC1 | 923 dropped | 121019001.D | Cal | 1 | 2019-12-10 12:19 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 2.5571 | 146.1191 | 1885.7691 | 13.6549 | 979157.4259 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7936 | 2.7165 | 155.2282 | 5483.255 | 13.6549 | 979157.4259 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.3266 | 0 | 0 | 0 | 13.6549 | 979157.4259 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6664 | 1.7489 | 99.9357 | 18833.0041 | 10.6554 | 247582.6763
|  |  |  | CC2 | 923 dropped | 121019002.D | Cal | 2 | 2019-12-10 12:49 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 3.8896 | 129.6533 | 3199.0797 | 13.655 | 953360.9038 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 3.0018 | 100.0602 | 6068.8063 | 13.655 | 953360.9038 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2758 | 8.0549 | 268.4955 | 818.7871 | 13.655 | 953360.9038 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 3.4807 | 116.0226 | 27327.5595 | 10.6504 | 215120.7671
|  |  |  | CC3 | 923 dropped | 121019003.D | Cal | 3 | 2019-12-10 13:19 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 5.1468 | 102.9368 | 3413.9759 | 13.655 | 725684.6679 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 6.2341 | 124.6825 | 10914.7611 | 13.655 | 725684.6679 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.3267 | 8.9021 | 178.0417 | 889.8228 | 13.655 | 725684.6679 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6564 | 3.7488 | 74.9759 | 27119.8678 | 10.6505 | 201000.9099
|  |  |  | CC4 | 923 dropped | 121019004.D | Cal | 4 | 2019-12-10 13:49 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 6.5038 | 86.7173 | 4084.3982 | 13.655 | 663010.6359 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 7.8112 | 104.1493 | 12778.3441 | 13.655 | 663010.6359 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 11.1702 | 148.9365 | 1465.0001 | 13.655 | 663010.6359 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 7.7455 | 103.2737 | 37729.5151 | 10.6504 | 149396.6597
|  |  |  | CC5 |  | 121019005.D | Cal | 5 | 2019-12-10 14:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 12.1269 | 97.0152 | 10962.979 | 13.655 | 899085.7336 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 11.61 | 92.8799 | 26494.6673 | 13.655 | 899085.7336 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 14.4892 | 115.9138 | 3280.4734 | 13.655 | 899085.7336 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 12.0743 | 96.5944 | 83410.5058 | 10.6504 | 219536.3448
|  |  |  | CC6 |  | 121019006.D | Cal | 6 | 2019-12-10 14:50 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 19.5691 | 97.8457 | 17467.3029 | 13.655 | 865646.067 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 18.8054 | 94.0268 | 42225.7558 | 13.655 | 865646.067 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 18.3178 | 91.5889 | 4595.4328 | 13.655 | 865646.067 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 20.1498 | 100.749 | 131467.2163 | 10.6505 | 212868.9604
|  |  |  | CC7 |  | 121019007.D | Cal | 7 | 2019-12-10 15:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 24.9585 | 79.8672 | 26536.6045 | 13.655 | 1022186.2036 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 29.6791 | 94.973 | 79692.1788 | 13.655 | 1022186.2036 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 29.4397 | 94.2071 | 10355.7322 | 13.655 | 1022186.2036 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 31.131 | 99.6193 | 217947.4038 | 10.6505 | 231669.4056
|  |  |  | CC8 |  | 121019008.D | Cal | 8 | 2019-12-10 15:50 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 39.0485 | 78.0971 | 40379.0409 | 13.655 | 982984.4922 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 43.9717 | 87.9435 | 114342.0544 | 13.655 | 982984.4922 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 47.5962 | 95.1925 | 17696.9954 | 13.655 | 982984.4922 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 48.9734 | 97.9468 | 317843.3051 | 10.6504 | 216830.6594
|  |  |  | CC9 | Dropped | 121019009.D | Cal | 9 | 2019-12-10 16:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 50.253 | 57.4321 | 50092.2303 | 13.655 | 943366.2225 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 46.9298 | 53.634 | 117222.8243 | 13.655 | 943366.2225 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 49.3522 | 56.4025 | 17701.9684 | 13.655 | 943366.2225 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 49.5604 | 56.6405 | 305486.2456 | 10.6505 | 205973.228
|  |  |  | CC10 |  | 121019010.D | Cal | 10 | 2019-12-10 16:50 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 103.6655 | 82.9324 | 106195.0945 | 13.655 | 961857.3725 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 126.1786 | 100.9429 | 324097.2095 | 13.655 | 961857.3725 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 127.9118 | 102.3295 | 50811.8843 | 13.655 | 961857.3725 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 139.4622 | 111.5698 | 847265.8404 | 10.6504 | 205203.8892
|  |  |  | CC11 |  | 121019011.D | Cal | 11 | 2019-12-10 17:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 189.0365 | 94.5182 | 207854.7327 | 13.6465 | 1028974.6696 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 193.5381 | 96.769 | 532730.5919 | 13.6465 | 1028974.6696 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 196.7748 | 98.3874 | 85080.4331 | 13.6465 | 1028974.6696 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 201.4194 | 100.7097 | 1379082.3207 | 10.6504 | 231690.294
|  |  |  | CC12 | Dropped | 121019012.D | Cal | 12 | 2019-12-10 17:50 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 42.9126 | 11.4434 | 2448.6758 | 13.621 | 54145.6959 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 58.6272 | 15.6339 | 8427.9853 | 13.621 | 54145.6959 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 62.082 | 16.5552 | 1314.8797 | 13.621 | 54145.6959 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.611 | 282.8531 | 75.4275 | 20074.1248 | 10.6051 | 2404.4276
|  |  |  | CC13 | Dropped | 121019013.D | Cal | 13 | 2019-12-10 18:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 728.9511 | 116.6322 | 332636.1852 | 13.6295 | 425756.6328 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 713.7281 | 114.1965 | 814822.5925 | 13.6295 | 425756.6328 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 360.9727 | 57.7556 | 65514.6711 | 13.6295 | 425756.6328 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 660.0401 | 105.6064 | 3278789.8082 | 10.6504 | 168581.9948
|  |  |  | CC14 |  | 121019014.D | Cal | 14 | 2019-12-10 18:50 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 884.853 | 101.1261 | 340217.1873 | 13.638 | 358670.3096 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 912.4574 | 104.2808 | 877728.9015 | 13.638 | 358670.3096 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 934.5931 | 106.8106 | 144397.2764 | 13.638 | 358670.3096 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 866.7687 | 99.0593 | 3338181.9287 | 10.6454 | 130739.1745
|  |  |  | CC15 |  | 121019015.D | Cal | 15 | 2019-12-10 19:20 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 1289.6447 | 103.1716 | 665791.4071 | 13.638 | 481460.7816 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 1224.9961 | 97.9997 | 1582065.8988 | 13.638 | 481460.7816 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 1194.6273 | 95.5702 | 248114.6014 | 13.638 | 481460.7816 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 1244.2973 | 99.5438 | 5975081.8387 | 10.6454 | 163059.3166
|  |  | 2nd CC | CC1 |  | 121019060.D | CC | 1 | 2019-12-11 12:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 1.9538 | 111.6446 | 1210.7023 | 13.6465 | 946883.1815 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 2.9408 | 168.0484 | 5872.6617 | 13.6465 | 946883.1815 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2078 | 8.2624 | 472.1388 | 898.4394 | 13.6465 | 946883.1815 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 1.928 | 110.1721 | 16370.2822 | 10.6353 | 201258.9146
|  |  |  | CC2 |  | 121019061.D | CC | 2 | 2019-12-11 12:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 3.2296 | 107.6549 | 2358.7416 | 13.6465 | 890928.7989 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 2.8541 | 95.1361 | 5318.1709 | 13.6465 | 890928.7989 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 9.4487 | 314.9555 | 1303.5816 | 13.6465 | 890928.7989 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 3.9651 | 132.1702 | 32462.5653 | 10.6353 | 229758.3584
|  |  |  | CC3 |  | 121019062.D | CC | 3 | 2019-12-11 13:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5546 | 3.7806 | 75.6117 | 2317.8667 | 13.6467 | 715695.8278 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 4.5308 | 90.6156 | 7492.7536 | 13.6467 | 715695.8278 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2249 | 8.9833 | 179.6665 | 902.7842 | 13.6467 | 715695.8278 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6414 | 3.8177 | 76.3549 | 23038.4085 | 10.6355 | 168220.8911
|  |  |  | CC4 |  | 121019063.D | CC | 4 | 2019-12-11 13:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 6.7118 | 89.4904 | 4837.6756 | 13.6466 | 757836.8888 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 8.3593 | 111.457 | 15720.6734 | 13.6466 | 757836.8888 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2163 | 11.0064 | 146.752 | 1620.6956 | 13.6466 | 757836.8888 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 7.7858 | 103.8102 | 42252.0949 | 10.6354 | 166523.8981
|  |  |  | CC5 |  | 121019064.D | CC | 5 | 2019-12-11 14:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 8.9355 | 71.4842 | 7520.9532 | 13.6466 | 857634.7381 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 9.6941 | 77.5524 | 20863.2004 | 13.6466 | 857634.7381 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 14.7258 | 117.8066 | 3217.2095 | 13.6466 | 857634.7381 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 12.1984 | 97.5869 | 73000.8697 | 10.6354 | 190309.3673
|  |  |  | CC6 |  | 121019065.D | CC | 6 | 2019-12-11 14:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5491 | 17.4257 | 87.1287 | 15603.2199 | 13.6466 | 872728.6505 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 20.1966 | 100.9831 | 45829.9043 | 13.6466 | 872728.6505 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2079 | 17.9197 | 89.5986 | 4482.402 | 13.6466 | 872728.6505 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 19.9157 | 99.5787 | 125514.273 | 10.6355 | 205522.3436
|  |  |  | CC7 |  | 121019066.D | CC | 7 | 2019-12-11 15:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5491 | 23.8508 | 76.3226 | 24817.8692 | 13.6466 | 1001844.4368 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 29.5067 | 94.4216 | 77642.9349 | 13.6466 | 1001844.4368 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2079 | 32.1211 | 102.7876 | 11314.4049 | 13.6466 | 1001844.4368 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 30.6813 | 98.1803 | 226452.9718 | 10.6355 | 244145.0078
|  |  |  | CC8 |  | 121019067.D | CC | 8 | 2019-12-11 15:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 40.8962 | 81.7924 | 29348.4849 | 13.6465 | 681565.6496 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 48.3483 | 96.6965 | 87286.1334 | 13.6465 | 681565.6496 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2077 | 38.7716 | 77.5432 | 9662.6286 | 13.6465 | 681565.6496 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6361 | 49.7231 | 99.4462 | 236718.2708 | 10.6353 | 159093.1011
|  |  |  | CC9 |  | 121019068.D | CC | 9 | 2019-12-11 16:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 41.9885 | 47.9869 | 40275.4253 | 13.6465 | 910541.7496 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 46.0462 | 52.6243 | 110984.9473 | 13.6465 | 910541.7496 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2077 | 42.6225 | 48.7114 | 14429.1588 | 13.6465 | 910541.7496 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 51.0969 | 58.3965 | 304988.6139 | 10.6353 | 199553.6735
|  |  |  | CC10 |  | 121019069.D | CC | 10 | 2019-12-11 16:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 93.0881 | 74.4704 | 84856.5413 | 13.6465 | 856638.1403 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 134.6618 | 107.7294 | 308146.8992 | 13.6465 | 856638.1403 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1993 | 95.1573 | 76.1258 | 33087.6487 | 13.6465 | 856638.1403 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 138.702 | 110.9616 | 819833.4928 | 10.6303 | 199641.6519
|  |  |  | CC11 |  | 121019070.D | CC | 11 | 2019-12-11 17:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 179.0725 | 89.5362 | 223268.9407 | 13.638 | 1167045.1184 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 177.1575 | 88.5787 | 552907.5218 | 13.638 | 1167045.1184 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 129.5798 | 64.7899 | 62495.3266 | 13.638 | 1167045.1184 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 199.4197 | 99.7099 | 1536077.0339 | 10.6303 | 260642.9205
|  |  |  | CC12 |  | 121019071.D | CC | 12 | 2019-12-11 17:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5545 | 294.1101 | 78.4294 | 31879.1902 | 13.6465 | 101288.4992 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 313.2502 | 83.5334 | 84982.5265 | 13.6465 | 101288.4992 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1823 | 130.4361 | 34.783 | 5461.6095 | 13.6465 | 101288.4992 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6261 | 390.0651 | 104.0174 | 240276.7157 | 10.6202 | 20886.3483
|  |  |  | CC13 |  | 121019072.D | CC | 13 | 2019-12-11 18:22 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 536.5405 | 85.8465 | 303854.3201 | 13.6296 | 528586.7676 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 698.3102 | 111.7296 | 989748.9132 | 13.6296 | 528586.7676 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1909 | 399.3619 | 63.8979 | 90136.2645 | 13.6296 | 528586.7676 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 664.9071 | 106.3851 | 3709734.8432 | 10.6303 | 189344.9976
|  |  |  | CC14 |  | 121019073.D | CC | 14 | 2019-12-11 18:52 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 675.7668 | 77.2305 | 452350.7583 | 13.6211 | 624603.803 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 851.3115 | 97.2927 | 1426014.6517 | 13.6211 | 624603.803 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 368.475 | 42.1114 | 98144.6679 | 13.6211 | 624603.803 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6312 | 881.7777 | 100.7746 | 4740765.3003 | 10.6303 | 182513.6768
|  |  |  | CC15 |  | 121019074.D | CC | 15 | 2019-12-11 19:23 | [16 empty] | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5434 | 942.1452 | 75.3716 | 1077231.3403 | 13.621 | 1066543.7172 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 1092.0583 | 87.3647 | 3124109.2286 | 13.621 | 1066543.7172 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1907 | 825.9585 | 66.0767 | 379143.8841 | 13.621 | 1066543.7172 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6311 | 1259.457 | 100.7566 | 9467701.7714 | 10.6303 | 255264.6389

## 102919RawOutput
Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| MTBE |  | 102919000.D | Blank |  | 2019-10-29 11:04 | y = 0.074827 * x | 0.99397 | 9.1013 | 0 | 0 | 0 | 13.706 | 0 | y = 0.068145 * x | 0.99536 | 9.2873 | 0 | 0 | 0 | 13.706 | 0 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.6477 | 0 | 0 | 0 | 13.706 | 0 | y = 0.064060 * x | 0.99408 | 11.8362 | 0 | 0 | 0 | 13.706 | 0 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4539 | 0 | 0 | 0 | 13.706 | 0
| MTBE |  | 102919022.D | Blank |  | 2019-10-29 22:06 | y = 0.074827 * x | 0.99397 | 9.1287 | 0 | 0 | 0 | 13.689 | 0 | y = 0.068145 * x | 0.99536 | 9.1338 | 0 | 0 | 0 | 13.689 | 0 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 0 | 0 | 0 | 13.689 | 0 | y = 0.064060 * x | 0.99408 | 11.8702 | 0 | 0 | 0 | 13.689 | 0 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5559 | 0 | 0 | 0 | 13.689 | 0
| MTBE |  | 102919034.D | Blank |  | 2019-10-30 04:07 | y = 0.074827 * x | 0.99397 | 9.485 | 0 | 0 | 0 | 13.689 | 0 | y = 0.068145 * x | 0.99536 | 9.3311 | 0 | 0 | 0 | 13.689 | 0 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5435 | 0 | 0 | 0 | 13.689 | 0 | y = 0.064060 * x | 0.99408 | 11.7088 | 0 | 0 | 0 | 13.689 | 0 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.471 | 0 | 0 | 0 | 13.689 | 0
| MTBE |  | 102919048.D | Blank |  | 2019-10-30 11:09 | y = 0.074827 * x | 0.99397 | 9.1231 | 0 | 0 | 0 | 13.6889 | 0 | y = 0.068145 * x | 0.99536 | 9.1337 | 0 | 0 | 0 | 13.6889 | 0 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5654 | 0 | 0 | 0 | 13.6889 | 0 | y = 0.064060 * x | 0.99408 | 12.7954 | 0 | 0 | 0 | 13.6889 | 0 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5982 | 0 | 0 | 0 | 13.6889 | 0
| UCG1CC1_8/13/2019_SE 10/15 |  | 102919001.D | Cal | 1 | 2019-10-29 11:34 | y = 0.074827 * x | 0.99397 | 9.1395 | 1.5924 | 90.9928 | 3427.028 | 13.6808 | 719040.0903 | y = 0.068145 * x | 0.99536 | 9.2158 | 1.5547 | 88.8401 | 3047.1492 | 13.6808 | 719040.0903 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 1.6106 | 92.0357 | 1006.3551 | 13.6808 | 719040.0903 | y = 0.064060 * x | 0.99408 | 11.8365 | 2.627 | 150.1148 | 4840.1934 | 13.6808 | 719040.0903 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5561 | 0 | 0 | 0 | 13.6808 | 719040.0903
| UCG1CC2_8/13/2019_SE 10/15 |  | 102919002.D | Cal | 2 | 2019-10-29 12:04 | y = 0.074827 * x | 0.99397 | 9.1396 | 2.6019 | 86.7284 | 6708.7658 | 13.6804 | 861470.2526 | y = 0.068145 * x | 0.99536 | 9.2159 | 2.9372 | 97.9083 | 6897.2286 | 13.6804 | 861470.2526 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 3.5224 | 117.4137 | 2637.3779 | 13.6804 | 861470.2526 | y = 0.064060 * x | 0.99408 | 11.8361 | 3.7263 | 124.2112 | 8225.6506 | 13.6804 | 861470.2526 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.6236 | 0 | 0 | 0 | 13.6804 | 861470.2526
| UCG1CC3_8/13/2019_SE 10/15 |  | 102919003.D | Cal | 3 | 2019-10-29 12:34 | y = 0.074827 * x | 0.99397 | 9.1396 | 5.0772 | 101.5446 | 15064.615 | 13.6804 | 991314.3628 | y = 0.068145 * x | 0.99536 | 9.2159 | 5.2094 | 104.1878 | 14076.4092 | 13.6804 | 991314.3628 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 5.9388 | 118.7764 | 5118.1754 | 13.6804 | 991314.3628 | y = 0.064060 * x | 0.99408 | 11.8361 | 5.3748 | 107.4966 | 13652.8796 | 13.6804 | 991314.3628 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5557 | 6.1804 | 123.609 | 2085.8999 | 13.6804 | 991314.3628
| UCG1CC4_8/13/2019_SE 10/15 |  | 102919004.D | Cal | 4 | 2019-10-29 13:04 | y = 0.074827 * x | 0.99397 | 9.1396 | 7.5916 | 101.2207 | 16745.6357 | 13.6804 | 736972.4781 | y = 0.068145 * x | 0.99536 | 9.2159 | 7.1453 | 95.2712 | 14353.8297 | 13.6804 | 736972.4781 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 7.3263 | 97.6843 | 4694.6701 | 13.6804 | 736972.4781 | y = 0.064060 * x | 0.99408 | 11.8361 | 6.7568 | 90.0909 | 12759.7267 | 13.6804 | 736972.4781 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 7.6424 | 101.8983 | 2255.5552 | 13.6804 | 736972.4781
| UCG1CC5_8/13/2019_SE 10/15 |  | 102919005.D | Cal | 5 | 2019-10-29 13:34 | y = 0.074827 * x | 0.99397 | 9.1396 | 17.154 | 137.2317 | 31639.0042 | 13.6804 | 616224.2482 | y = 0.068145 * x | 0.99536 | 9.2159 | 14.9678 | 119.7425 | 25141.4687 | 13.6804 | 616224.2482 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 17.4023 | 139.218 | 9334.2028 | 13.6804 | 616224.2482 | y = 0.064060 * x | 0.99408 | 11.8361 | 14.0524 | 112.4189 | 22188.9271 | 13.6804 | 616224.2482 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5388 | 13.8142 | 110.5136 | 4371.1887 | 13.6804 | 616224.2482
| UCG1CC6_8/13/2019_SE 10/15 |  | 102919006.D | Cal | 6 | 2019-10-29 14:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 24.4558 | 122.2789 | 65876.2888 | 13.6804 | 899969.5806 | y = 0.068145 * x | 0.99536 | 9.2159 | 25.0831 | 125.4155 | 61532.2151 | 13.6804 | 899969.5806 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 26.6824 | 133.4118 | 20922.4723 | 13.6804 | 899969.5806 | y = 0.064060 * x | 0.99408 | 11.8361 | 23.4916 | 117.4581 | 54173.7659 | 13.6804 | 899969.5806 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5643 | 25.8472 | 129.2359 | 13440.6263 | 13.6804 | 899969.5806
| UCG1CC7_8/13/2019_SE 10/15 |  | 102919007.D | Cal | 7 | 2019-10-29 14:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 27.0387 | 86.524 | 71463.7088 | 13.6804 | 883037.9864 | y = 0.068145 * x | 0.99536 | 9.2159 | 25.9704 | 83.1051 | 62510.1702 | 13.6804 | 883037.9864 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 29.2994 | 93.7581 | 22548.6026 | 13.6804 | 883037.9864 | y = 0.064060 * x | 0.99408 | 11.8361 | 27.6087 | 88.3478 | 62470.2555 | 13.6804 | 883037.9864 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5642 | 34.6613 | 110.9163 | 18243.3718 | 13.6804 | 883037.9864
| UCG1CC8_8/13/2019_SE 10/15 |  | 102919008.D | Cal | 8 | 2019-10-29 15:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 41.2898 | 82.5796 | 113572.8108 | 13.6804 | 918992.5392 | y = 0.068145 * x | 0.99536 | 9.2159 | 43.4854 | 86.9707 | 108930.2888 | 13.6804 | 918992.5392 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 42.3932 | 84.7863 | 34001.0384 | 13.6804 | 918992.5392 | y = 0.064060 * x | 0.99408 | 11.8361 | 41.6535 | 83.307 | 98087.0024 | 13.6804 | 918992.5392 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5557 | 35.7415 | 71.4829 | 19629.9628 | 13.6804 | 918992.5392
| UCG1CC9_8/13/2019_SE 10/15 |  | 102919009.D | Cal | 9 | 2019-10-29 15:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 91.8645 | 104.988 | 192073.8962 | 13.6804 | 698555.8284 | y = 0.068145 * x | 0.99536 | 9.2159 | 91.5888 | 104.6729 | 174396.1774 | 13.6804 | 698555.8284 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5928 | 91.3828 | 104.4375 | 56001.3511 | 13.6804 | 698555.8284 | y = 0.064060 * x | 0.99408 | 11.8361 | 87.8795 | 100.4337 | 157302.9047 | 13.6804 | 698555.8284 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 70.4223 | 80.4826 | 30547.5546 | 13.6804 | 698555.8284
| UCG1CC10_8/13/2019_SE 10/15 |  | 102919010.D | Cal | 10 | 2019-10-29 16:05 | y = 0.074827 * x | 0.99397 | 9.1397 | 129.3084 | 103.4467 | 387525.4519 | 13.6805 | 1001276.0262 | y = 0.068145 * x | 0.99536 | 9.216 | 130.6627 | 104.5302 | 356614.4799 | 13.6805 | 1001276.0262 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 127.1415 | 101.7132 | 112100.6642 | 13.6805 | 1001276.0262 | y = 0.064060 * x | 0.99408 | 11.8362 | 126.4875 | 101.19 | 324526.1277 | 13.6805 | 1001276.0262 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5474 | 138.7736 | 111.0189 | 87226.0027 | 13.6805 | 1001276.0262
| UCG1CC11_8/13/2019_SE 10/15 |  | 102919011.D | Cal | 11 | 2019-10-29 16:35 | y = 0.074827 * x | 0.99397 | 9.1395 | 190.859 | 95.4295 | 331860.3707 | 13.6808 | 580929.1207 | y = 0.068145 * x | 0.99536 | 9.2158 | 192.1874 | 96.0937 | 304327.731 | 13.6808 | 580929.1207 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 191.9128 | 95.9564 | 98841.233 | 13.6808 | 580929.1207 | y = 0.064060 * x | 0.99408 | 11.8365 | 168.8277 | 84.4139 | 251313.047 | 13.6808 | 580929.1207 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5477 | 201.4772 | 100.7386 | 73253.2349 | 13.6808 | 580929.1207
| UCG1CC12_8/13/2019_SE 10/15 |  | 102919012.D | Cal | 12 | 2019-10-29 17:05 | y = 0.074827 * x | 0.99397 | 9.1396 | 341.4471 | 91.0526 | 773830.0929 | 13.6804 | 757185.9152 | y = 0.068145 * x | 0.99536 | 9.2159 | 345.0649 | 92.0173 | 712191.4673 | 13.6804 | 757185.9152 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 379.9858 | 101.3296 | 260086.7356 | 13.6804 | 757185.9152 | y = 0.064060 * x | 0.99408 | 11.8361 | 345.6993 | 92.1865 | 670732.1088 | 13.6804 | 757185.9152 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5388 | 375.297 | 100.0792 | 174199.8837 | 13.6804 | 757185.9152
| UCG1CC13_8/13/2019_SE 10/15 |  | 102919013.D | Cal | 13 | 2019-10-29 17:35 | y = 0.074827 * x | 0.99397 | 9.1396 | 541.171 | 86.5874 | 1479118.0833 | 13.6804 | 913163.8009 | y = 0.068145 * x | 0.99536 | 9.2159 | 570.7341 | 91.3175 | 1420613.7996 | 13.6804 | 913163.8009 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5873 | 545.3287 | 87.2526 | 457762.0072 | 13.6804 | 913163.8009 | y = 0.064060 * x | 0.99408 | 11.8361 | 539.3824 | 86.3012 | 1262099.4482 | 13.6804 | 913163.8009 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5472 | 445.0657 | 71.2105 | 246645.047 | 13.6804 | 913163.8009
| UCG1CC14_8/13/2019_SE 10/15 |  | 102919014.D | Cal | 14 | 2019-10-29 18:05 | y = 0.074827 * x | 0.99397 | 9.1397 | 937.9818 | 107.1979 | 1834222.6826 | 13.6805 | 653338.0475 | y = 0.068145 * x | 0.99536 | 9.216 | 947.5996 | 108.2971 | 1687548.6307 | 13.6805 | 653338.0475 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 873.9171 | 99.8762 | 542206.2546 | 13.6805 | 653338.0475 | y = 0.064060 * x | 0.99408 | 11.8362 | 926.0159 | 105.8304 | 1550259.8615 | 13.6805 | 653338.0475 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5389 | 943.6588 | 107.8467 | 344956.4835 | 13.6805 | 653338.0475
| UCG1CC15_8/13/2019_SE 10/15 |  | 102919015.D | Cal | 15 | 2019-10-29 18:35 | y = 0.074827 * x | 0.99397 | 9.1397 | 1258.9555 | 100.7164 | 2293484.3484 | 13.6805 | 608647.3341 | y = 0.068145 * x | 0.99536 | 9.216 | 1235.9796 | 98.8784 | 2050550.688 | 13.6805 | 608647.3341 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 1250.2608 | 100.0209 | 749124.555 | 13.6805 | 608647.3341 | y = 0.064060 * x | 0.99408 | 11.8362 | 1271.0537 | 101.6843 | 1982338.1001 | 13.6805 | 608647.3341 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5389 | 1249.8533 | 99.9883 | 402813.9981 | 13.6805 | 608647.3341
| UCG1CC13_8/13/2019_SE 10/15 |  | 102919047.D | Cal |  | 2019-10-30 10:38 | y = 0.074827 * x | 0.99397 | 9.1286 | 520.6648 |  | 1485369.8037 | 13.6723 | 953139.9075 | y = 0.068145 * x | 0.99536 | 9.2049 | 509.0559 |  | 1322560.6072 | 13.6723 | 953139.9075 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5763 | 551.3972 |  | 483413.7138 | 13.6723 | 953139.9075 | y = 0.064060 * x | 0.99408 | 11.828 | 519.4134 |  | 1268580.1978 | 13.6723 | 953139.9075 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4882 | 455.9354 |  | 263304.6942 | 13.6723 | 953139.9075
| UCG1CC8_8/13/2019_SE 10/15 |  | 102919028.D | QC | 8 | 2019-10-30 01:07 | y = 0.074827 * x | 0.99397 | 9.1342 | 41.0464 | 82.0928 | 97980.0797 | 13.672 | 797522.2599 | y = 0.068145 * x | 0.99536 | 9.2105 | 40.4463 | 80.8925 | 87925.4608 | 13.672 | 797522.2599 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 40.791 | 81.582 | 28386.8988 | 13.672 | 797522.2599 | y = 0.064060 * x | 0.99408 | 11.8362 | 39.7745 | 79.549 | 81282.2669 | 13.672 | 797522.2599 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5303 | 26.1377 | 52.2754 | 12061.3077 | 13.672 | 797522.2599
| UCG1CC4_8/13/2019_SE 10/15 |  | 102919040.D | QC | 4 | 2019-10-30 07:08 | y = 0.074827 * x | 0.99397 | 9.1287 | 7.3466 | 97.9542 | 16959.3937 | 13.672 | 771269.7136 | y = 0.068145 * x | 0.99536 | 9.205 | 6.3367 | 84.4889 | 13321.7336 | 13.672 | 771269.7136 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 8.0178 | 106.904 | 5377.2603 | 13.672 | 771269.7136 | y = 0.064060 * x | 0.99408 | 11.8277 | 6.0682 | 80.9098 | 11992.7006 | 13.672 | 771269.7136 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5134 | 7.9885 | 106.5131 | 2535.1137 | 13.672 | 771269.7136
| UCG1T1hS1A_SE 10/11 |  | 102919016.D | Sample |  | 2019-10-29 19:05 | y = 0.074827 * x | 0.99397 | 9.1342 | 251.0068 |  | 536695.0318 | 13.6805 | 714368.7788 | y = 0.068145 * x | 0.99536 | 9.2105 | 252.8435 |  | 492343.0577 | 13.6805 | 714368.7788 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 0 |  | 0 | 13.6805 | 714368.7788 | y = 0.064060 * x | 0.99408 | 11.2504 | 0 |  | 0 | 13.6805 | 714368.7788 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5304 | 10.1973 |  | 3379.6595 | 13.6805 | 714368.7788
| UCG1T1hS1B_SE 10/11 |  | 102919017.D | Sample |  | 2019-10-29 19:35 | y = 0.074827 * x | 0.99397 | 9.1397 | 217.9529 |  | 488838.2047 | 13.6805 | 749346.7367 | y = 0.068145 * x | 0.99536 | 9.2105 | 218.3639 |  | 446022.9759 | 13.6805 | 749346.7367 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0 |  | 0 | 13.6805 | 749346.7367 | y = 0.064060 * x | 0.99408 | 11.2504 | 0 |  | 0 | 13.6805 | 749346.7367 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4031 | 0 |  | 0 | 13.6805 | 749346.7367
| UCG1T1hS2B_SE 10/11 |  | 102919018.D | Sample |  | 2019-10-29 20:05 | y = 0.074827 * x | 0.99397 | 9.2055 | 0 |  | 0 | 13.6805 | 634966.2531 | y = 0.068145 * x | 0.99536 | 9.3201 | 0 |  | 0 | 13.6805 | 634966.2531 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 0 |  | 0 | 13.6805 | 634966.2531 | y = 0.064060 * x | 0.99408 | 11.8362 | 104.3026 |  | 169704.7742 | 13.6805 | 634966.2531 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5219 | 85.4047 |  | 33855.8411 | 13.6805 | 634966.2531
| UCG1T1hS2C_SE 10/11 |  | 102919019.D | Sample |  | 2019-10-29 20:35 | y = 0.074827 * x | 0.99397 | 9.0849 | 0 |  | 0 | 13.6805 | 782386.9322 | y = 0.068145 * x | 0.99536 | 9.216 | 0 |  | 0 | 13.6805 | 782386.9322 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 0 |  | 0 | 13.6805 | 782386.9322 | y = 0.064060 * x | 0.99408 | 11.8362 | 105.4396 |  | 211384.5608 | 13.6805 | 782386.9322 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5304 | 96.5129 |  | 47256.3023 | 13.6805 | 782386.9322
| UCG1T1hS3A_SE 10/11 |  | 102919020.D | Sample |  | 2019-10-29 21:06 | y = 0.074827 * x | 0.99397 | 9.5837 | 0 |  | 0 | 13.6805 | 749514.2523 | y = 0.068145 * x | 0.99536 | 9.5832 | 0 |  | 0 | 13.6805 | 749514.2523 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 208.1333 |  | 138537.3171 | 13.6805 | 749514.2523 | y = 0.064060 * x | 0.99408 | 11.6494 | 0 |  | 0 | 13.6805 | 749514.2523 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5049 | 0 |  | 0 | 13.6805 | 749514.2523
| UCG1T1hS3B_SE 10/11 |  | 102919021.D | Sample |  | 2019-10-29 21:36 | y = 0.074827 * x | 0.99397 | 9.5837 | 0 |  | 0 | 13.655 | 434357.2951 | y = 0.068145 * x | 0.99536 | 9.5832 | 0 |  | 0 | 13.655 | 434357.2951 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 295.0724 |  | 114851.0425 | 13.655 | 434357.2951 | y = 0.064060 * x | 0.99408 | 12.2097 | 0 |  | 0 | 13.655 | 434357.2951 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5219 | 0 |  | 0 | 13.655 | 434357.2951
| UCG1T5hS1B_SE 10/11 |  | 102919023.D | Sample |  | 2019-10-29 22:36 | y = 0.074827 * x | 0.99397 | 9.1342 | 215.3843 |  | 340363.774 | 13.6805 | 527970.3481 | y = 0.068145 * x | 0.99536 | 9.2105 | 221.5918 |  | 318901.561 | 13.6805 | 527970.3481 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 0 |  | 0 | 13.6805 | 527970.3481 | y = 0.064060 * x | 0.99408 | 11.2504 | 0 |  | 0 | 13.6805 | 527970.3481 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4964 | 0 |  | 0 | 13.6805 | 527970.3481
| UCG1T5hS1C_SE 10/11 |  | 102919024.D | Sample |  | 2019-10-29 23:06 | y = 0.074827 * x | 0.99397 | 9.1342 | 231.9324 |  | 434477.9547 | 13.672 | 625873.5638 | y = 0.068145 * x | 0.99536 | 9.2105 | 242.4423 |  | 413607.6166 | 13.672 | 625873.5638 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5161 | 0 |  | 0 | 13.672 | 625873.5638 | y = 0.064060 * x | 0.99408 | 11.4542 | 0 |  | 0 | 13.672 | 625873.5638 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.089 | 0 |  | 0 | 13.672 | 625873.5638
| UCG1T5hS3B_SE 10/11 |  | 102919025.D | Sample |  | 2019-10-29 23:36 | y = 0.074827 * x | 0.99397 | 9.5837 | 0 |  | 0 | 13.672 | 843178.2604 | y = 0.068145 * x | 0.99536 | 9.5832 | 0 |  | 0 | 13.672 | 843178.2604 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 180.971 |  | 135127.4084 | 13.672 | 843178.2604 | y = 0.064060 * x | 0.99408 | 11.2504 | 0 |  | 0 | 13.672 | 843178.2604 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.488 | 0 |  | 0 | 13.672 | 843178.2604
| UCG1T5hS3C_SE 10/11 |  | 102919026.D | Sample |  | 2019-10-30 00:06 | y = 0.074827 * x | 0.99397 | 9.5837 | 0 |  | 0 | 13.672 | 571516.0405 | y = 0.068145 * x | 0.99536 | 9.5832 | 0 |  | 0 | 13.672 | 571516.0405 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 180.1358 |  | 91160.2875 | 13.672 | 571516.0405 | y = 0.064060 * x | 0.99408 | 11.573 | 0 |  | 0 | 13.672 | 571516.0405 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4794 | 0 |  | 0 | 13.672 | 571516.0405
| UCG1T5hS2A_SE 10/11 |  | 102919027.D | Sample |  | 2019-10-30 00:36 | y = 0.074827 * x | 0.99397 | 9.0849 | 0 |  | 0 | 13.672 | 818523.2678 | y = 0.068145 * x | 0.99536 | 9.2763 | 0 |  | 0 | 13.672 | 818523.2678 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.549 | 0 |  | 0 | 13.672 | 818523.2678 | y = 0.064060 * x | 0.99408 | 11.8362 | 91.6396 |  | 192203.9308 | 13.672 | 818523.2678 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5134 | 82.5032 |  | 42125.7141 | 13.672 | 818523.2678
| UCG1AFS1A_SE 10/11 |  | 102919029.D | Sample |  | 2019-10-30 01:37 | y = 0.074827 * x | 0.99397 | 9.1342 | 752.3411 |  | 1343293.3381 | 13.672 | 596535.6631 | y = 0.068145 * x | 0.99536 | 9.2105 | 257.0915 |  | 418039.832 | 13.672 | 596535.6631 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0 |  | 0 | 13.672 | 596535.6631 | y = 0.064060 * x | 0.99408 |  |  |  |  | 13.672 | 596535.6631 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.6068 | 0 |  | 0 | 13.672 | 596535.6631
| UCG1AFS1B_SE 10/11 |  | 102919030.D | Sample |  | 2019-10-30 02:07 | y = 0.074827 * x | 0.99397 | 9.1342 | 614.2034 |  | 1221970.228 | 13.672 | 664704.6284 | y = 0.068145 * x | 0.99536 | 9.2105 | 197.7298 |  | 358256.7429 | 13.672 | 664704.6284 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0 |  | 0 | 13.672 | 664704.6284 | y = 0.064060 * x | 0.99408 | 12.3116 | 0 |  | 0 | 13.672 | 664704.6284 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.2417 | 0 |  | 0 | 13.672 | 664704.6284
| UCG1AFS3A_SE 10/11 |  | 102919031.D | Sample |  | 2019-10-30 02:37 | y = 0.074827 * x | 0.99397 | 9.1287 | 3.5686 |  | 6167.9476 | 13.672 | 577462.2216 | y = 0.068145 * x | 0.99536 | 9.4572 | 0 |  | 0 | 13.672 | 577462.2216 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 23.2242 |  | 11680.6053 | 13.672 | 577462.2216 | y = 0.064060 * x | 0.99408 | 11.7598 | 0 |  | 0 | 13.672 | 577462.2216 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4794 | 0 |  | 0 | 13.672 | 577462.2216
| UCG1AFS3B_SE 10/11 |  | 102919032.D | Sample |  | 2019-10-30 03:07 | y = 0.074827 * x | 0.99397 | 9.1342 | 5.1219 |  | 7925.8706 | 13.6805 | 517009.4195 | y = 0.068145 * x | 0.99536 | 9.4572 | 0 |  | 0 | 13.6805 | 517009.4195 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 26.0463 |  | 11732.1032 | 13.6805 | 517009.4195 | y = 0.064060 * x | 0.99408 | 11.4542 | 0 |  | 0 | 13.6805 | 517009.4195 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 24.252 | 0 |  | 0 | 13.6805 | 517009.4195
| UCG1T5hS2C_SE 10/11 |  | 102919033.D | Sample |  | 2019-10-30 03:37 | y = 0.074827 * x | 0.99397 | 9.0904 | 0 |  | 0 | 13.672 | 816873.3158 | y = 0.068145 * x | 0.99536 | 9.2982 | 0 |  | 0 | 13.672 | 816873.3158 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.7189 | 0 |  | 0 | 13.672 | 816873.3158 | y = 0.064060 * x | 0.99408 | 11.8277 | 97.5127 |  | 204109.7864 | 13.672 | 816873.3158 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5049 | 69.4956 |  | 35235.8127 | 13.672 | 816873.3158
| UCG1T5hS3A_SE 10/11 |  | 102919035.D | Sample |  | 2019-10-30 04:37 | y = 0.074827 * x | 0.99397 | 9.5782 | 0 |  | 0 | 13.672 | 826905.0143 | y = 0.068145 * x | 0.99536 | 9.5777 | 0 |  | 0 | 13.672 | 826905.0143 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 222.6971 |  | 163784.7585 | 13.672 | 826905.0143 | y = 0.064060 * x | 0.99408 | 11.2504 | 0 |  | 0 | 13.672 | 826905.0143 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.7426 | 0 |  | 0 | 13.672 | 826905.0143
| UCG1AFS1C_SE 10/11 |  | 102919036.D | Sample |  | 2019-10-30 05:08 | y = 0.074827 * x | 0.99397 | 9.1342 | 655.459 |  | 1094688.8628 | 13.672 | 557988.8506 | y = 0.068145 * x | 0.99536 | 9.205 | 231.6119 |  | 352273.4133 | 13.672 | 557988.8506 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0 |  | 0 | 13.672 | 557988.8506 | y = 0.064060 * x | 0.99408 | 11.5815 | 0 |  | 0 | 13.672 | 557988.8506 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5134 | 0 |  | 0 | 13.672 | 557988.8506
| UCG1T1hS3C_SE 10/11 |  | 102919037.D | Sample |  | 2019-10-30 05:38 | y = 0.074827 * x | 0.99397 | 9.5782 | 0 |  | 0 | 13.672 | 627458.8341 | y = 0.068145 * x | 0.99536 | 9.5777 | 0 |  | 0 | 13.672 | 627458.8341 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 216.9131 |  | 120979.9055 | 13.672 | 627458.8341 | y = 0.064060 * x | 0.99408 | 12.1248 | 0 |  | 0 | 13.672 | 627458.8341 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.2587 | 0 |  | 0 | 13.672 | 627458.8341
| UCG1T5hS1A_SE 10/11 |  | 102919038.D | Sample |  | 2019-10-30 06:08 | y = 0.074827 * x | 0.99397 | 9.1287 | 225.5136 |  | 373042.1079 | 13.6805 | 552669.425 | y = 0.068145 * x | 0.99536 | 9.205 | 225.0489 |  | 339028.2634 | 13.6805 | 552669.425 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 0 |  | 0 | 13.6805 | 552669.425 | y = 0.064060 * x | 0.99408 | 11.8022 | 0 |  | 0 | 13.6805 | 552669.425 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5219 | 0 |  | 0 | 13.6805 | 552669.425
| UCG1AFS3C_SE 10/11 |  | 102919039.D | Sample |  | 2019-10-30 06:38 | y = 0.074827 * x | 0.99397 | 9.1287 | 6.7806 |  | 11074.4346 | 13.6805 | 545678.0939 | y = 0.068145 * x | 0.99536 | 9.205 | 4.3242 |  | 6431.786 | 13.6805 | 545678.0939 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 27.2635 |  | 12963.0098 | 13.6805 | 545678.0939 | y = 0.064060 * x | 0.99408 |  |  |  |  | 13.6805 | 545678.0939 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.3436 | 0 |  | 0 | 13.6805 | 545678.0939
| UCG1T1hS1C_SE 10/11 |  | 102919041.D | Sample |  | 2019-10-30 07:38 | y = 0.074827 * x | 0.99397 | 9.1287 | 231.2477 |  | 362841.8664 | 13.6805 | 524228.1999 | y = 0.068145 * x | 0.99536 | 9.205 | 237.7014 |  | 339660.975 | 13.6805 | 524228.1999 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0 |  | 0 | 13.6805 | 524228.1999 | y = 0.064060 * x | 0.99408 | 11.2419 | 0 |  | 0 | 13.6805 | 524228.1999 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.471 | 0 |  | 0 | 13.6805 | 524228.1999
| UCG1T1hS2A_SE 10/11 |  | 102919042.D | Sample |  | 2019-10-30 08:08 | y = 0.074827 * x | 0.99397 | 9.1178 | 0 |  | 0 | 13.6805 | 688971.1368 | y = 0.068145 * x | 0.99536 | 9.4572 | 0 |  | 0 | 13.6805 | 688971.1368 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5874 | 0 |  | 0 | 13.6805 | 688971.1368 | y = 0.064060 * x | 0.99408 | 11.8277 | 102.3976 |  | 180775.3125 | 13.6805 | 688971.1368 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5049 | 63.8171 |  | 27205.8449 | 13.6805 | 688971.1368
| UCG1AFS2B_SE 10/11 |  | 102919043.D | Sample |  | 2019-10-30 08:38 | y = 0.074827 * x | 0.99397 | 9.1287 | 5.8163 |  | 12461.2746 | 13.672 | 715807.7031 | y = 0.068145 * x | 0.99536 | 9.205 | 3.8621 |  | 7535.6126 | 13.672 | 715807.7031 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 0 |  | 0 | 13.672 | 715807.7031 | y = 0.064060 * x | 0.99408 | 11.8277 | 283.3687 |  | 519752.1948 | 13.672 | 715807.7031 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4964 | 247.6244 |  | 110438.6045 | 13.672 | 715807.7031
| UCG1AFS2C_SE 10/11 |  | 102919044.D | Sample |  | 2019-10-30 09:08 | y = 0.074827 * x | 0.99397 | 9.1287 | 5.2572 |  | 9021.7043 | 13.672 | 573341.5777 | y = 0.068145 * x | 0.99536 | 9.205 | 3.7654 |  | 5884.6443 | 13.672 | 573341.5777 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 0.64857 |  | 323.0947 | 13.672 | 573341.5777 | y = 0.064060 * x | 0.99408 | 11.8277 | 807.2929 |  | 1186021.8821 | 13.672 | 573341.5777 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.5049 | 892.5226 |  | 288861.2835 | 13.672 | 573341.5777
| UCG1T5hS2B_SE 10/11 |  | 102919045.D | Sample |  | 2019-10-30 09:38 | y = 0.074827 * x | 0.99397 | 10.0429 | 0 |  | 0 | 13.672 | 743496.3375 | y = 0.068145 * x | 0.99536 | 9.2982 | 0 |  | 0 | 13.672 | 743496.3375 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5764 | 0 |  | 0 | 13.672 | 743496.3375 | y = 0.064060 * x | 0.99408 | 11.8277 | 98.0796 |  | 186855.3265 | 13.672 | 743496.3375 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4964 | 78.2934 |  | 36262.6329 | 13.672 | 743496.3375
| UCG1AFS2A_SE 10/11 |  | 102919046.D | Sample |  | 2019-10-30 10:08 | y = 0.074827 * x | 0.99397 | 9.1287 | 5.1738 |  | 8044.7932 | 13.6805 | 519504.8035 | y = 0.068145 * x | 0.99536 | 9.205 | 0 |  | 0 | 13.6805 | 519504.8035 | y = 5.779670E-005 * x ^ 2 + 0.021720 * x | 0.99989 | 9.5819 | 0.82738 |  | 373.4769 | 13.6805 | 519504.8035 | y = 0.064060 * x | 0.99408 | 11.8277 | 286.4343 |  | 381296.434 | 13.6805 | 519504.8035 | y = -6.226089E-005 * x ^ 2 + 0.016389 * x - 0.001944 | 0.99926 | 23.4964 | 290.2147 |  | 93471.7478 | 13.6805 | 519504.8035

## 102919Raw4NT
[6 empty] | 4-Nitrotuluene Method |  | 4-Nitrotuluene Results |  |  |  | 4NT13C6 (ISTD) Results |  | Note: 11/1 data used to create curve
Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
G1+CC CC5_Nov.1 |  | 110119013.D | Cal | 5 | 2019-11-01 15:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 11.0092 | 88.074 | 85922.8028 | 10.6858 | 142139.58
G1+CC CC6_Nov.1 |  | 110119012.D | Cal | 6 | 2019-11-01 15:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 21.3217 | 106.6086 | 129222.1642 | 10.6858 | 141646.0342
G1+CC CC7_Nov.1 |  | 110119011.D | Cal | 7 | 2019-11-01 14:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 34.431 | 110.1792 | 217454.3475 | 10.6857 | 166815.9327
G1+CC CC8_Nov.1 |  | 110119010.D | Cal | 8 | 2019-11-01 14:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 48.5082 | 97.0164 | 310387.6602 | 10.6857 | 180068.8012
G1+CC CC9_Nov.1 |  | 110119009.D | Cal | 9 | 2019-11-01 13:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6865 | 86.2481 | 98.5692 | 482156.3722 | 10.6857 | 169170.1173
G1+CC CC10_Nov.1 |  | 110119008.D | Cal | 10 | 2019-11-01 13:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 123.0436 | 98.4348 | 529585.9447 | 10.6858 | 134128.3581
G1+CC CC11_Nov.1 |  | 110119007.D | Cal | 11 | 2019-11-01 12:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 198.7426 | 99.3713 | 846983.6499 | 10.6858 | 136440.3565
G1+CC CC12_Nov.1 |  | 110119006.D | Cal | 12 | 2019-11-01 12:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 377.2542 | 100.6011 | 1646726.6012 | 10.6858 | 142750.4376
G1+CC CC13_Nov.1 |  | 110119005.D | Cal | 13 | 2019-11-01 11:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 632.5709 | 101.2113 | 3934261.086 | 10.6857 | 205379.3612
G1+CC CC14_Nov.1 |  | 110119004.D | Cal | 14 | 2019-11-01 11:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 889.1237 | 101.6141 | 5205424.3285 | 10.6857 | 194135.7217
G1+CC CC15_Nov.1 |  | 110119003.D | Cal | 15 | 2019-11-01 10:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 1228.9968 | 98.3197 | 7127691.7007 | 10.6907 | 192862.2567
G1+CC Spiked MM_Nov.1 |  | 110119002.D | MatrixBlank |  | 2019-11-01 10:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7118 | 0 |  | 31115.1083 | 10.6908 | 212346.1131
G1+CC Spiked MM_Nov.1 |  | 110119001.D | MatrixBlank |  | 2019-11-01 09:27 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7067 | 0 |  | 27819.2267 | 10.6908 | 195788.9811
MTBE |  | 102919048.D | Blank |  | 2019-10-30 11:09 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6966 | 0 | 0 | 0 | 10.6151 | 0
UCG1AFS2A_SE 10/11 |  | 102919046.D | Sample |  | 2019-10-30 10:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 29.2628 |  | 137650.5214 | 10.6908 | 119768.454
UCG1T5hS2B_SE 10/11 |  | 102919045.D | Sample |  | 2019-10-30 09:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 105.973 |  | 535954.894 | 10.6908 | 155852.859
UCG1AFS2C_SE 10/11 |  | 102919044.D | Sample |  | 2019-10-30 09:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 70.1815 |  | 300186.8179 | 10.6908 | 126629.4829
UCG1AFS2B_SE 10/11 |  | 102919043.D | Sample |  | 2019-10-30 08:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 32.342 |  | 203318.7442 | 10.6908 | 163807.2727
UCG1T1hS2A_SE 10/11 |  | 102919042.D | Sample |  | 2019-10-30 08:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 120.2512 |  | 631905.8769 | 10.6908 | 163493.9884
UCG1T1hS1C_SE 10/11 |  | 102919041.D | Sample |  | 2019-10-30 07:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 204.9475 |  | 751962.6504 | 10.6908 | 117624.389
UCG1AFS3C_SE 10/11 |  | 102919039.D | Sample |  | 2019-10-30 06:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 45.4613 |  | 236287.5488 | 10.6908 | 144715.2044
UCG1T5hS1A_SE 10/11 |  | 102919038.D | Sample |  | 2019-10-30 06:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 158.7909 |  | 639507.296 | 10.6908 | 127511.4805
UCG1T1hS3C_SE 10/11 |  | 102919037.D | Sample |  | 2019-10-30 05:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 194.1016 |  | 813175.2011 | 10.6958 | 133983.9266
UCG1AFS1C_SE 10/11 |  | 102919036.D | Sample |  | 2019-10-30 05:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 49.7417 |  | 247997.1067 | 10.6908 | 140864.9458
UCG1T5hS3A_SE 10/11 |  | 102919035.D | Sample |  | 2019-10-30 04:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 156.8419 |  | 955242.6169 | 10.6908 | 192701.0513
MTBE |  | 102919034.D | Blank |  | 2019-10-30 04:07 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7873 | 0 |  | 0 | 10.8016 | 0
UCG1T5hS2C_SE 10/11 |  | 102919033.D | Sample |  | 2019-10-30 03:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 105.3093 |  | 615327.7822 | 10.6958 | 179970.7696
UCG1AFS3B_SE 10/11 |  | 102919032.D | Sample |  | 2019-10-30 03:07 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 51.3178 |  | 239496.6075 | 10.6958 | 132496.1929
UCG1AFS3A_SE 10/11 |  | 102919031.D | Sample |  | 2019-10-30 02:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 49.7484 |  | 206419.3933 | 10.6958 | 117234.8739
UCG1AFS1B_SE 10/11 |  | 102919030.D | Sample |  | 2019-10-30 02:07 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 42.7224 |  | 235048.9524 | 10.6958 | 151543.8603
UCG1AFS1A_SE 10/11 |  | 102919029.D | Sample |  | 2019-10-30 01:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 55.5445 |  | 288965.2862 | 10.6958 | 149434.3378
UCG1T5hS2A_SE 10/11 |  | 102919027.D | Sample |  | 2019-10-30 00:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 108.0025 |  | 589490.5923 | 10.6958 | 168453.4576
UCG1T5hS3C_SE 10/11 |  | 102919026.D | Sample |  | 2019-10-30 00:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 155.0856 |  | 678759.6727 | 10.6958 | 138389.5867
UCG1T5hS3B_SE 10/11 |  | 102919025.D | Sample |  | 2019-10-29 23:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 159.0605 |  | 918682.4946 | 10.6958 | 182882.7801
UCG1T5hS1C_SE 10/11 |  | 102919024.D | Sample |  | 2019-10-29 23:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 160.9751 |  | 791660.6949 | 10.6958 | 155823.8592
UCG1T5hS1B_SE 10/11 |  | 102919023.D | Sample |  | 2019-10-29 22:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 161.3961 |  | 598028.657 | 10.6958 | 117420.5272
MTBE |  | 102919022.D | Blank |  | 2019-10-29 22:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7672 | 0 |  | 0 | 10.6958 | 0
UCG1T1hS3B_SE 10/11 |  | 102919021.D | Sample |  | 2019-10-29 21:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 190.2535 |  | 976246.0272 | 10.6958 | 163955.1388
UCG1T1hS3A_SE 10/11 |  | 102919020.D | Sample |  | 2019-10-29 21:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 186.651 |  | 1013827.9057 | 10.6958 | 173397.994
UCG1T1hS2C_SE 10/11 |  | 102919019.D | Sample |  | 2019-10-29 20:35 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 126.5147 |  | 699851.9262 | 10.6958 | 172719.6256
UCG1T1hS2B_SE 10/11 |  | 102919018.D | Sample |  | 2019-10-29 20:05 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 124.6609 |  | 569660.7411 | 10.6958 | 142535.4979
UCG1T1hS1B_SE 10/11 |  | 102919017.D | Sample |  | 2019-10-29 19:35 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 191.6225 |  | 1079948.8643 | 10.6958 | 180135.344
UCG1T1hS1A_SE 10/11 |  | 102919016.D | Sample |  | 2019-10-29 19:05 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 204.1802 |  | 1121180.8762 | 10.6958 | 176009.205
MTBE |  | 102919000.D | Blank |  | 2019-10-29 11:04 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.5557 | 0 |  | 0 | 10.585 | 0

## 110119RawOutput
Sample | [6 empty] | 908 Method |  | 908 Results |  |  |  | M8FOSA (ISTD) Results |  | 909 Method |  | 909 Results |  |  |  | M8FOSA (ISTD) Results |  | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| MTBE |  | 110119000.D | Blank |  | 2019-11-01 08:57 | y = 0.098018 * x | 0.99826 | 9.1342 | 0 | 0 | 0 | 13.689 | 0 | y = 0.092110 * x | 0.99679 | 9.2763 | 0 | 0 | 0 | 13.689 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5655 | 0 | 0 | 0 | 13.689 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.8377 | 0 | 0 | 0 | 10.7915 | 0 | y = 0.079567 * x | 0.99598 | 12.804 | 0 | 0 | 0 | 13.689 | 0 | y = 0.012003 * x | 0.99495 | 23.4965 | 0 | 0 | 0 | 13.689 | 0
| MTBE |  | 110119022.D | Blank |  | 2019-11-01 20:23 | y = 0.098018 * x | 0.99826 | 9.0959 | 0 | 0 | 0 | 13.6806 | 0 | y = 0.092110 * x | 0.99679 | 9.2434 | 0 |  | 0 | 13.6806 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5601 | 0 | 0 | 0 | 13.6806 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.591 | 0 | 0 | 0 | 10.7865 | 244.9984 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.6806 | 0 | y = 0.012003 * x | 0.99495 | 23.3693 | 0 | 0 | 0 | 13.6806 | 0
| MTBE |  | 110119036.D | Blank |  | 2019-11-02 03:24 | y = 0.098018 * x | 0.99826 | 9.1013 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.092110 * x | 0.99679 | 9.1227 | 0 |  | 0 | 13.6805 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5819 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7268 | 0 | 0 | 0 | 10.7864 | 443.0716 | y = 0.079567 * x | 0.99598 | 11.6918 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.012003 * x | 0.99495 | 23.4964 | 0 | 0 | 0 | 13.6805 | 0
| MTBE |  | 110119049.D | Blank |  | 2019-11-02 09:55 | y = 0.098018 * x | 0.99826 | 9.0958 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.092110 * x | 0.99679 | 9.2762 | 0 |  | 0 | 13.6805 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5928 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 0 | 0 | 0 | 10.5447 | 0 | y = 0.079567 * x | 0.99598 | 11.8871 | 0 | 0 | 0 | 13.6805 | 0 | y = 0.012003 * x | 0.99495 | 23.5134 | 0 | 0 | 0 | 13.6805 | 0
| MTBE |  | 110119061.D | Blank |  | 2019-11-02 15:55 | y = 0.098018 * x | 0.99826 | 9.1232 | 0 | 0 | 0 | 13.672 | 0 | y = 0.092110 * x | 0.99679 | 9.0953 | 0 |  | 0 | 13.672 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5051 | 0 | 0 | 0 | 13.672 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.8427 | 0 | 0 | 0 | 10.7713 | 0 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.672 | 0 | y = 0.012003 * x | 0.99495 | 23.5559 | 0 | 0 | 0 | 13.672 | 0
| MTBE |  | 110119069.D | Blank |  | 2019-11-02 19:56 | y = 0.098018 * x | 0.99826 | 9.0519 | 0 | 0 | 0 | 13.672 | 0 | y = 0.092110 * x | 0.99679 | 9.4352 | 0 |  | 0 | 13.672 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 0 | 0 | 0 | 13.672 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7218 | 0 | 0 | 0 | 10.7562 | 0 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.672 | 0 | y = 0.012003 * x | 0.99495 | 22.7918 | 0 | 0 | 0 | 13.672 | 0
| MTBE |  | 110119080.D | Blank |  | 2019-11-03 01:27 | y = 0.098018 * x | 0.99826 | 9.0904 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.092110 * x | 0.99679 | 9.2927 | 0 | 0 | 0 | 13.6721 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5107 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.8377 | 0 | 0 | 0 | 10.5599 | 0 | y = 0.079567 * x | 0.99598 | 12.7955 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.012003 * x | 0.99495 | 23.4286 | 0 | 0 | 0 | 13.6721 | 0
| MTBE |  | 110119088.D | Blank |  | 2019-11-03 05:02 | y = 0.098018 * x | 0.99826 | 9.0849 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.092110 * x | 0.99679 | 9.2818 | 0 | 0 | 0 | 13.6721 | 0 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6477 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.8377 | 0 | 0 | 0 | 10.6707 | 0 | y = 0.079567 * x | 0.99598 | 11.191 | 0 | 0 | 0 | 13.6721 | 0 | y = 0.012003 * x | 0.99495 | 23.5304 | 0 | 0 | 0 | 13.6721 | 0
| G1+CC CC15_Nov.1 |  | 110119003.D | Cal | 15 | 2019-11-01 10:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 1282.0888 | 102.5671 | 3088291.3528 | 13.6805 | 614375.897 | y = 0.092110 * x | 0.99679 | 9.2049 | 1291.9138 | 103.3531 | 2924389.391 | 13.6805 | 614375.897 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 1241.5357 | 99.3229 | 914479.021 | 13.6805 | 614375.897 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 1228.9968 | 98.3197 | 7127691.7007 | 10.6907 | 192862.2567 | y = 0.079567 * x | 0.99598 | 11.8276 | 1308.9286 | 104.7143 | 2559428.3304 | 13.6805 | 614375.897 | y = 0.012003 * x | 0.99495 | 23.4964 | 1269.2163 | 101.5373 | 374400.6196 | 13.6805 | 614375.897
| G1+CC CC14_Nov.1 |  | 110119004.D | Cal | 14 | 2019-11-01 11:22 | y = 0.098018 * x | 0.99826 | 9.1287 | 858.4608 | 98.1098 | 2367255.0818 | 13.672 | 703329.149 | y = 0.092110 * x | 0.99679 | 9.205 | 869.3948 | 99.3594 | 2252906.3189 | 13.672 | 703329.149 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 889.9314 | 101.7064 | 775821.5446 | 13.672 | 703329.149 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 889.1237 | 101.6141 | 5205424.3285 | 10.6857 | 194135.7217 | y = 0.079567 * x | 0.99598 | 11.8276 | 828.827 | 94.7231 | 1855305.0945 | 13.672 | 703329.149 | y = 0.012003 * x | 0.99495 | 23.4879 | 830.1339 | 94.8724 | 280332.5333 | 13.672 | 703329.149
| G1+CC CC13_Nov.1 |  | 110119005.D | Cal | 13 | 2019-11-01 11:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 578.8779 | 92.6205 | 1830704.2438 | 13.672 | 806612.9952 | y = 0.092110 * x | 0.99679 | 9.2049 | 556.9159 | 89.1065 | 1655092.6592 | 13.672 | 806612.9952 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 557.7707 | 89.2433 | 574916.8592 | 13.672 | 806612.9952 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 632.5709 | 101.2113 | 3934261.086 | 10.6857 | 205379.3612 | y = 0.079567 * x | 0.99598 | 11.8276 | 573.6178 | 91.7788 | 1472585.919 | 13.672 | 806612.9952 | y = 0.012003 * x | 0.99495 | 23.4964 | 677.8209 | 108.4513 | 262510.6297 | 13.672 | 806612.9952
| G1+CC CC12_Nov.1 |  | 110119006.D | Cal | 12 | 2019-11-01 12:22 | y = 0.098018 * x | 0.99826 | 9.1287 | 378.0035 | 100.8009 | 744473.877 | 13.6806 | 502328.3423 | y = 0.092110 * x | 0.99679 | 9.205 | 363.6109 | 96.9629 | 672964.2874 | 13.6806 | 502328.3423 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5819 | 387.7014 | 103.387 | 252693.3098 | 13.6806 | 502328.3423 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 377.2542 | 100.6011 | 1646726.6012 | 10.6858 | 142750.4376 | y = 0.079567 * x | 0.99598 | 11.8277 | 370.8606 | 98.8962 | 592913.3437 | 13.6806 | 502328.3423 | y = 0.012003 * x | 0.99495 | 23.4965 | 315.4637 | 84.1236 | 76085.8268 | 13.6806 | 502328.3423
| G1+CC CC11_Nov.1 |  | 110119007.D | Cal | 11 | 2019-11-01 12:52 | y = 0.098018 * x | 0.99826 | 9.1288 | 205.9028 | 102.9514 | 396309.4317 | 13.6806 | 490914.9508 | y = 0.092110 * x | 0.99679 | 9.2051 | 198.5623 | 99.2812 | 359145.6143 | 13.6806 | 490914.9508 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 187.8395 | 93.9198 | 121775.5564 | 13.6806 | 490914.9508 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 198.7426 | 99.3713 | 846983.6499 | 10.6858 | 136440.3565 | y = 0.079567 * x | 0.99598 | 11.8278 | 202.8357 | 101.4179 | 316915.5223 | 13.6806 | 490914.9508 | y = 0.012003 * x | 0.99495 | 23.4796 | 234.5584 | 117.2792 | 55287.1183 | 13.6806 | 490914.9508
| G1+CC CC10_Nov.1 |  | 110119008.D | Cal | 10 | 2019-11-01 13:22 | y = 0.098018 * x | 0.99826 | 9.1288 | 132.6477 | 106.1182 | 242497.257 | 13.6721 | 466273.8622 | y = 0.092110 * x | 0.99679 | 9.205 | 127.9846 | 102.3877 | 219870.2161 | 13.6721 | 466273.8622 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 99.9609 | 79.9687 | 62024.455 | 13.6721 | 466273.8622 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 123.0436 | 98.4348 | 529585.9447 | 10.6858 | 134128.3581 | y = 0.079567 * x | 0.99598 | 11.8277 | 126.1178 | 100.8942 | 187158.8389 | 13.6721 | 466273.8622 | y = 0.012003 * x | 0.99495 | 23.488 | 107.3064 | 85.8451 | 24023.3393 | 13.6721 | 466273.8622
| G1+CC CC9_Nov.1 |  | 110119009.D | Cal | 9 | 2019-11-01 13:52 | y = 0.098018 * x | 0.99826 | 9.1286 | 88.3173 | 100.934 | 229080.1853 | 13.672 | 661570.0892 | y = 0.092110 * x | 0.99679 | 9.2049 | 82.9002 | 94.7431 | 202068.6875 | 13.672 | 661570.0892 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 74.1712 | 84.7671 | 65444.6188 | 13.672 | 661570.0892 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6865 | 86.2481 | 98.5692 | 482156.3722 | 10.6857 | 169170.1173 | y = 0.079567 * x | 0.99598 | 11.8276 | 85.6722 | 97.9111 | 180388.4598 | 13.672 | 661570.0892 | y = 0.012003 * x | 0.99495 | 23.4794 | 91.7192 | 104.8219 | 29134.1824 | 13.672 | 661570.0892
| G1+CC CC8_Nov.1 |  | 110119010.D | Cal | 8 | 2019-11-01 14:22 | y = 0.098018 * x | 0.99826 | 9.1232 | 47.6901 | 95.3802 | 124575.268 | 13.672 | 666250.3013 | y = 0.092110 * x | 0.99679 | 9.205 | 46.2056 | 92.4113 | 113422.6796 | 13.672 | 666250.3013 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5819 | 47.6566 | 95.3132 | 42444.3011 | 13.672 | 666250.3013 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 48.5082 | 97.0164 | 310387.6602 | 10.6857 | 180068.8012 | y = 0.079567 * x | 0.99598 | 11.8277 | 49.9394 | 99.8788 | 105894.5384 | 13.672 | 666250.3013 | y = 0.012003 * x | 0.99495 | 23.4794 | 44.9861 | 89.9721 | 14390.7049 | 13.672 | 666250.3013
| G1+CC CC7_Nov.1 |  | 110119011.D | Cal | 7 | 2019-11-01 14:52 | y = 0.098018 * x | 0.99826 | 9.1287 | 29.5408 | 94.5306 | 73952.892 | 13.672 | 638508.322 | y = 0.092110 * x | 0.99679 | 9.205 | 27.5792 | 88.2535 | 64880.8037 | 13.672 | 638508.322 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5818 | 20.0521 | 64.1666 | 17156.116 | 13.672 | 638508.322 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 34.431 | 110.1792 | 217454.3475 | 10.6857 | 166815.9327 | y = 0.079567 * x | 0.99598 | 11.8276 | 28.3697 | 90.783 | 57651.9656 | 13.672 | 638508.322 | y = 0.012003 * x | 0.99495 | 23.488 | 24.4713 | 78.3081 | 7502.2233 | 13.672 | 638508.322
| G1+CC CC6_Nov.1 |  | 110119012.D | Cal | 6 | 2019-11-01 15:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 19.4559 | 97.2797 | 39682.6904 | 13.6721 | 520215.186 | y = 0.092110 * x | 0.99679 | 9.205 | 18.1166 | 90.5828 | 34723.7034 | 13.6721 | 520215.186 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 12.3925 | 61.9627 | 8644.1693 | 13.6721 | 520215.186 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 21.3217 | 106.6086 | 129222.1642 | 10.6858 | 141646.0342 | y = 0.079567 * x | 0.99598 | 11.8277 | 18.2282 | 91.1411 | 30180.0501 | 13.6721 | 520215.186 | y = 0.012003 * x | 0.99495 | 23.4626 | 14.8962 | 74.4811 | 3720.7116 | 13.6721 | 520215.186
| G1+CC CC5_Nov.1 |  | 110119013.D | Cal | 5 | 2019-11-01 15:52 | y = 0.098018 * x | 0.99826 | 9.1288 | 13.9431 | 111.5448 | 26946.7662 | 13.6721 | 492925.4119 | y = 0.092110 * x | 0.99679 | 9.1996 | 12.5319 | 100.2556 | 22759.7392 | 13.6721 | 492925.4119 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 11.2377 | 89.9013 | 7428.144 | 13.6721 | 492925.4119 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6917 | 11.0092 | 88.074 | 85922.8028 | 10.6858 | 142139.58 | y = 0.079567 * x | 0.99598 | 11.8193 | 12.1024 | 96.8192 | 18986.5356 | 13.6721 | 492925.4119 | y = 0.012003 * x | 0.99495 | 23.4626 | 15.9709 | 127.7676 | 3779.8855 | 13.6721 | 492925.4119
| G1+CC CC4_Nov.1 |  | 110119014.D | Cal | 4 | 2019-11-01 16:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 9.2407 | 123.2093 | 19541.8442 | 13.6721 | 539380.1551 | y = 0.092110 * x | 0.99679 | 9.1996 | 6.8546 | 91.3947 | 13622.1159 | 13.6721 | 539380.1551 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 6.3326 | 84.4349 | 4582.3136 | 13.6721 | 539380.1551 | [8 empty] | y = 0.079567 * x | 0.99598 | 11.8192 | 6.7585 | 90.1139 | 11602.2092 | 13.6721 | 539380.1551 | y = 0.012003 * x | 0.99495 | 23.4371 | 9.1581 | 122.1077 | 2371.7349 | 13.6721 | 539380.1551
| G1+CC CC3_Nov.1 |  | 110119015.D | Cal | 3 | 2019-11-01 16:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 4.8963 | 97.925 | 14175.5513 | 13.6721 | 738431.6469 | y = 0.092110 * x | 0.99679 | 9.1996 | 4.5115 | 90.2299 | 12274.3495 | 13.6721 | 738431.6469 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 4.0411 | 80.8225 | 4004.101 | 13.6721 | 738431.6469 | [8 empty] | y = 0.079567 * x | 0.99598 | 11.8193 | 4.9071 | 98.1418 | 11532.595 | 13.6721 | 738431.6469 | y = 0.012003 * x | 0.99495 | 23.4796 | 3.6383 | 72.7668 | 1289.9729 | 13.6721 | 738431.6469
| G1+CC CC2_Nov.1 |  | 110119016.D | Cal | 2 | 2019-11-01 17:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 3.3552 | 111.8388 | 8868.7775 | 13.6721 | 674192.9048 | y = 0.092110 * x | 0.99679 | 9.1996 | 3.0144 | 100.4785 | 7487.6605 | 13.6721 | 674192.9048 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 2.1169 | 70.5647 | 1915.3903 | 13.6721 | 674192.9048 | [8 empty] | y = 0.079567 * x | 0.99598 | 11.8193 | 2.4418 | 81.395 | 5239.5726 | 13.6721 | 674192.9048 | y = 0.012003 * x | 0.99495 | 23.4626 | 3.0292 | 100.9717 | 980.5551 | 13.6721 | 674192.9048
| G1+CC CC1_Nov.1 |  | 110119017.D | Cal | 1 | 2019-11-01 17:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 2.2413 | 128.075 | 4287.5537 | 13.6721 | 487911.5221 | y = 0.092110 * x | 0.99679 | 9.1996 | 2.7254 | 155.736 | 4899.3174 | 13.6721 | 487911.5221 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 1.3614 | 77.7966 | 891.5218 | 13.6721 | 487911.5221 | [8 empty] | y = 0.079567 * x | 0.99598 | 11.8278 | 2.3361 | 133.4899 | 3627.6082 | 13.6721 | 487911.5221 | y = 0.012003 * x | 0.99495 | 23.4541 | 0 | 0 | 0 | 13.6721 | 487911.5221
| G1+CC Spiked MM_Nov.1 |  | 110119001.D | MatrixBlank |  | 2019-11-01 09:27 | y = 0.098018 * x | 0.99826 | 9.0904 | 0 | 0 | 0 | 13.6805 | 693049.9593 | y = 0.092110 * x | 0.99679 | 9.2818 | 0 | 0 | 0 | 13.6805 | 693049.9593 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5271 | 0 | 0 | 0 | 13.6805 | 693049.9593 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7067 | 0 |  | 27819.2267 | 10.6908 | 195788.9811 | y = 0.079567 * x | 0.99598 | 11.7598 | 0 | 0 | 0 | 13.6805 | 693049.9593 | y = 0.012003 * x | 0.99495 | 22.707 | 0 | 0 | 0 | 13.6805 | 693049.9593
| G1+CC Spiked MM_Nov.1 |  | 110119002.D | MatrixBlank |  | 2019-11-01 10:22 | y = 0.098018 * x | 0.99826 | 9.1123 | 0 | 0 | 0 | 13.6806 | 638474.5558 | y = 0.092110 * x | 0.99679 | 9.1064 | 0 | 0 | 0 | 13.6806 | 638474.5558 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 0 | 0 | 0 | 13.6806 | 638474.5558 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7118 | 0 |  | 31115.1083 | 10.6908 | 212346.1131 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.6806 | 638474.5558 | y = 0.012003 * x | 0.99495 | 23.471 | 0 | 0 | 0 | 13.6806 | 638474.5558
| G1+CC CC10_Nov.1 |  | 110119029.D | QC | 10 | 2019-11-01 23:53 | y = 0.098018 * x | 0.99826 | 9.1233 | 123.4516 | 98.7613 | 264896.4917 | 13.6721 | 547284.7764 | y = 0.092110 * x | 0.99679 | 9.1996 | 117.2017 | 93.7614 | 236327.7859 | 13.6721 | 547284.7764 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 103.0165 | 82.4132 | 75006.1811 | 13.6721 | 547284.7764 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6816 | 129.0087 | 103.207 | 554073.6634 | 10.6808 | 134275.5661 | y = 0.079567 * x | 0.99598 | 11.8193 | 124.8739 | 99.8991 | 217509.2912 | 13.6721 | 547284.7764 | y = 0.012003 * x | 0.99495 | 23.4456 | 140.9835 | 112.7868 | 37046.6172 | 13.6721 | 547284.7764
| QCGCCC11_1015 |  | 110119030.D | QC | 11 | 2019-11-02 00:23 | y = 0.098018 * x | 0.99826 | 9.1178 | 167.1372 | 83.5686 | 392118.8132 | 13.6721 | 598381.9479 | y = 0.092110 * x | 0.99679 | 9.1996 | 161.2066 | 80.6033 | 355409.2372 | 13.6721 | 598381.9479 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 148.3675 | 74.1837 | 117646.9397 | 13.6721 | 598381.9479 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 165.0864 | 82.5432 | 997886.0055 | 10.6808 | 191783.3587 | y = 0.079567 * x | 0.99598 | 11.8193 | 172.6403 | 86.3202 | 328786.2125 | 13.6721 | 598381.9479 | y = 0.012003 * x | 0.99495 | 23.4456 | 185.2108 | 92.6054 | 53212.2256 | 13.6721 | 598381.9479
| G1+CC CC6_Nov.1 |  | 110119042.D | QC | 6 | 2019-11-02 06:24 | y = 0.098018 * x | 0.99826 | 9.1177 | 17.5151 | 87.5756 | 39552.5315 | 13.672 | 575963.7133 | y = 0.092110 * x | 0.99679 | 9.194 | 15.1668 | 75.834 | 32185.218 | 13.672 | 575963.7133 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 16.3332 | 81.6658 | 12609.5058 | 13.672 | 575963.7133 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 23.8823 | 119.4114 | 131142.185 | 10.6756 | 132639.188 | y = 0.079567 * x | 0.99598 | 11.8107 | 16.5151 | 82.5753 | 30273.8789 | 13.672 | 575963.7133 | y = 0.012003 * x | 0.99495 | 23.4031 | 17.404 | 87.0202 | 4812.9591 | 13.672 | 575963.7133
| QCGCCC8_1015 |  | 110119048.D | QC | 8 | 2019-11-02 09:25 | y = 0.098018 * x | 0.99826 | 9.1177 | 39.2391 | 78.4783 | 108541.2968 | 13.672 | 705520.3756 | y = 0.092110 * x | 0.99679 | 9.194 | 37.5047 | 75.0094 | 97490.56 | 13.672 | 705520.3756 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 33.6473 | 67.2945 | 31771.9275 | 13.672 | 705520.3756 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 57.122 | 114.2441 | 419062.2539 | 10.6756 | 211560.9112 | y = 0.079567 * x | 0.99598 | 11.8107 | 38.1391 | 76.2781 | 85639.1589 | 13.672 | 705520.3756 | y = 0.012003 * x | 0.99495 | 23.4285 | 38.5289 | 77.0579 | 13051.5882 | 13.672 | 705520.3756
| G1+CC CC11_Nov.1 |  | 110119055.D | QC | 11 | 2019-11-02 12:55 | y = 0.098018 * x | 0.99826 | 9.1122 | 195.0314 | 97.5157 | 424859.4293 | 13.672 | 555616.0971 | y = 0.092110 * x | 0.99679 | 9.1885 | 185.0293 | 92.5146 | 378776.2619 | 13.672 | 555616.0971 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 152.1121 | 76.0561 | 111959.3672 | 13.672 | 555616.0971 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 205.3934 | 102.6967 | 901407.5382 | 10.6756 | 140708.0983 | y = 0.079567 * x | 0.99598 | 11.8107 | 209.5168 | 104.7584 | 370498.5185 | 13.672 | 555616.0971 | y = 0.012003 * x | 0.99495 | 23.4115 | 174.1538 | 87.0769 | 46459.4884 | 13.672 | 555616.0971
| G1+CC CC5_Nov.1 |  | 110119066.D | QC | 5 | 2019-11-02 18:26 | y = 0.098018 * x | 0.99826 | 9.1122 | 14.6709 | 117.3673 | 33331.3588 | 13.6635 | 579468.4184 | y = 0.092110 * x | 0.99679 | 9.1885 | 11.7252 | 93.8019 | 25033.3336 | 13.6635 | 579468.4184 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 10.4799 | 83.8395 | 8144.0438 | 13.6635 | 579468.4184 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 11.6745 | 93.3958 | 86723.3439 | 10.6706 | 138901.6434 | y = 0.079567 * x | 0.99598 | 11.8107 | 12.0411 | 96.329 | 22206.9989 | 13.6635 | 579468.4184 | y = 0.012003 * x | 0.99495 | 23.3776 | 11.5567 | 92.4537 | 3215.3696 | 13.6635 | 579468.4184
| QCGCCC4_1015 |  | 110119070.D | QC | 4 | 2019-11-02 20:26 | y = 0.098018 * x | 0.99826 | 9.1122 | 6.3185 | 84.2462 | 20710.1883 | 13.6635 | 836000.6785 | y = 0.092110 * x | 0.99679 | 9.1885 | 5.6775 | 75.7003 | 17487.7228 | 13.6635 | 836000.6785 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 5.5429 | 73.9051 | 6216.9662 | 13.6635 | 836000.6785 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6865 | 29.9583 | 399.4445 | 295570.3274 | 10.6706 | 252610.4257 | y = 0.079567 * x | 0.99598 | 11.8021 | 6.696 | 89.2799 | 17816.173 | 13.6635 | 836000.6785 | y = 0.012003 * x | 0.99495 | 23.3775 | 5.3552 | 71.4023 | 2149.5468 | 13.6635 | 836000.6785
| G1+CC CC9_Nov.1 |  | 110119076.D | QC | 9 | 2019-11-02 23:27 | y = 0.098018 * x | 0.99826 | 9.1068 | 79.061 | 90.3554 | 228213.4197 | 13.6636 | 736228.8787 | y = 0.092110 * x | 0.99679 | 9.1831 | 74.2709 | 84.881 | 201464.6792 | 13.6636 | 736228.8787 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.56 | 76.6974 | 87.6542 | 75294.1734 | 13.6636 | 736228.8787 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6715 | 90.0703 | 102.9375 | 499857.2217 | 10.6656 | 168630.9184 | y = 0.079567 * x | 0.99598 | 11.8023 | 90.9518 | 103.9449 | 213116.423 | 13.6636 | 736228.8787 | y = 0.012003 * x | 0.99495 | 23.3691 | 112.1114 | 128.1273 | 39630.4806 | 13.6636 | 736228.8787
| G1+CC CC4_Nov.1 |  | 110119087.D | QC | 4 | 2019-11-03 04:32 | y = 0.098018 * x | 0.99826 | 9.1068 | 8.2957 | 110.6093 | 18445.3915 | 13.6636 | 567111.9673 | y = 0.092110 * x | 0.99679 | 9.1831 | 7.3682 | 98.2427 | 15395.6396 | 13.6636 | 567111.9673 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.56 | 6.5073 | 86.764 | 4950.7364 | 13.6636 | 567111.9673 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6715 | 7.0434 | 93.9125 | 72150.0566 | 10.6656 | 148417.3429 | y = 0.079567 * x | 0.99598 | 11.8023 | 8.2615 | 110.1528 | 14911.4041 | 13.6636 | 567111.9673 | y = 0.012003 * x | 0.99495 | 23.5303 | 2.0878 | 27.8379 | 568.5041 | 13.6636 | 567111.9673
| UCG1 T1 S1A1009_SE 11/1 |  | 110119018.D | Sample |  | 2019-11-01 18:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 281.0361 |  | 630834.918 | 13.6721 | 572515.9293 | y = 0.092110 * x | 0.99679 | 9.1996 | 271.2912 |  | 572256.5843 | 13.6721 | 572515.9293 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.571 | 0 |  | 0 | 13.6721 | 572515.9293 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 195.4269 |  | 1211178.5546 | 10.6858 | 198269.1936 | y = 0.079567 * x | 0.99598 | 11.242 | 0 |  | 0 | 13.6721 | 572515.9293 | y = 0.012003 * x | 0.99495 | 23.573 | 0 |  | 0 | 13.6721 | 572515.9293
| UCG1 T1 S1B1009_SE 11/1 |  | 110119019.D | Sample |  | 2019-11-01 18:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 235.5626 |  | 548421.52 | 13.6721 | 593802.6131 | y = 0.092110 * x | 0.99679 | 9.1996 | 229.2873 |  | 501636.9374 | 13.6721 | 593802.6131 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5655 | 0 |  | 0 | 13.6721 | 593802.6131 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 181.811 |  | 1121879.2065 | 10.6858 | 196739.2672 | y = 0.079567 * x | 0.99598 | 11.242 | 0 |  | 0 | 13.6721 | 593802.6131 | y = 0.012003 * x | 0.99495 | 23.5645 | 0 |  | 0 | 13.6721 | 593802.6131
| UCG1 T1 S1C1009_SE 11/1 |  | 110119020.D | Sample |  | 2019-11-01 19:22 | y = 0.098018 * x | 0.99826 | 9.1233 | 293.9222 |  | 451729.9105 | 13.6721 | 391994.9233 | y = 0.092110 * x | 0.99679 | 9.1996 | 288.3369 |  | 416435.9378 | 13.6721 | 391994.9233 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5875 | 0 |  | 0 | 13.6721 | 391994.9233 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 190.1028 |  | 887592.675 | 10.6858 | 149178.9935 | y = 0.079567 * x | 0.99598 | 11.242 | 0 |  | 0 | 13.6721 | 391994.9233 | y = 0.012003 * x | 0.99495 | 23.0297 | 0 |  | 0 | 13.6721 | 391994.9233
| UCG1 T1 S1C1009_SE 11/1b |  | 110119021.D | Sample |  | 2019-11-01 19:52 | y = 0.098018 * x | 0.99826 | 9.1233 | 274.0997 |  | 464665.9818 | 13.6721 | 432380.734 | y = 0.092110 * x | 0.99679 | 9.1996 | 270.0455 |  | 430200.4296 | 13.6721 | 432380.734 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.571 | 0 |  | 0 | 13.6721 | 432380.734 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 195.5895 |  | 912536.0424 | 10.6808 | 149263.0458 | y = 0.079567 * x | 0.99598 | 11.242 | 0 |  | 0 | 13.6721 | 432380.734 | y = 0.012003 * x | 0.99495 |  |  |  |  | 13.6721 | 432380.734
| UCG1 T1 S1C1009_SE 11/1c |  | 110119023.D | Sample |  | 2019-11-01 20:53 | y = 0.098018 * x | 0.99826 | 9.1233 | 275.521 |  | 489321.3643 | 13.6721 | 452974.3085 | y = 0.092110 * x | 0.99679 | 9.1996 | 270.5173 |  | 451477.4472 | 13.6721 | 452974.3085 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6367 | 0 |  | 0 | 13.6721 | 452974.3085 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 194.9973 |  | 935461.7902 | 10.6807 | 153456.6273 | y = 0.079567 * x | 0.99598 | 11.5646 | 0 |  | 0 | 13.6721 | 452974.3085 | y = 0.012003 * x | 0.99495 | 23.4541 | 0 |  | 0 | 13.6721 | 452974.3085
| UCG1 AF S1A1009_SE 11/1 |  | 110119024.D | Sample |  | 2019-11-01 21:23 | y = 0.098018 * x | 0.99826 | 9.1233 | 759.023 |  | 2267924.5944 | 13.6721 | 762092.5023 | y = 0.092110 * x | 0.99679 | 9.1996 | 252.7821 |  | 709776.2041 | 13.6721 | 762092.5023 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.582 | 0 |  | 0 | 13.6721 | 762092.5023 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 50.6386 |  | 495449.9333 | 10.6807 | 277205.6544 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.6721 | 762092.5023 | y = 0.012003 * x | 0.99495 | 23.5729 | 0 |  | 0 | 13.6721 | 762092.5023
| UCG1 AF S1B1009_SE 11/1 |  | 110119025.D | Sample |  | 2019-11-01 21:53 | y = 0.098018 * x | 0.99826 | 9.1233 | 643.118 |  | 1464793.9512 | 13.6721 | 580924.7762 | y = 0.092110 * x | 0.99679 | 9.1996 | 207.5886 |  | 444314.9717 | 13.6721 | 580924.7762 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 0 |  | 0 | 13.6721 | 580924.7762 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 56.149 |  | 397409.0147 | 10.6808 | 203614.7766 | y = 0.079567 * x | 0.99598 | 12.023 | 0 |  | 0 | 13.6721 | 580924.7762 | y = 0.012003 * x | 0.99495 | 23.5474 | 0 |  | 0 | 13.6721 | 580924.7762
| UCG1 AF S1C1009_SE 11/1 |  | 110119026.D | Sample |  | 2019-11-01 22:23 | y = 0.098018 * x | 0.99826 | 9.1233 | 732.5599 |  | 1208506.7949 | 13.6721 | 420765.3534 | y = 0.092110 * x | 0.99679 | 9.1996 | 241.8328 |  | 374906.177 | 13.6721 | 420765.3534 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5765 | 0 |  | 0 | 13.6721 | 420765.3534 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 52.9511 |  | 277680.1347 | 10.6807 | 149586.1709 | y = 0.079567 * x | 0.99598 | 11.5731 | 0 |  | 0 | 13.6721 | 420765.3534 | y = 0.012003 * x | 0.99495 | 23.4201 | 0 |  | 0 | 13.6721 | 420765.3534
| UCG1 T5 S1A1009_SE 11/1 |  | 110119027.D | Sample |  | 2019-11-01 22:53 | y = 0.098018 * x | 0.99826 | 9.1233 | 210.9424 |  | 395535.1132 | 13.6721 | 478250.2926 | y = 0.092110 * x | 0.99679 | 9.1996 | 209.9587 |  | 369961.6139 | 13.6721 | 478250.2926 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6806 | 0 |  | 0 | 13.6721 | 478250.2926 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 162.4426 |  | 770511.5218 | 10.6807 | 150364.717 | y = 0.079567 * x | 0.99598 | 12.6087 | 0 |  | 0 | 13.6721 | 478250.2926 | y = 0.012003 * x | 0.99495 | 23.4795 | 0 |  | 0 | 13.6721 | 478250.2926
| UCG1 T5 S1B1009_SE 11/1 |  | 110119028.D | Sample |  | 2019-11-01 23:23 | y = 0.098018 * x | 0.99826 | 9.1233 | 231.8167 |  | 385517.1195 | 13.6721 | 424163.2534 | y = 0.092110 * x | 0.99679 | 9.1996 | 240.0174 |  | 375096.6488 | 13.6721 | 424163.2534 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.571 | 0 |  | 0 | 13.6721 | 424163.2534 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 154.0909 |  | 718306.2518 | 10.6807 | 147344.5174 | y = 0.079567 * x | 0.99598 | 11.8787 | 0 |  | 0 | 13.6721 | 424163.2534 | y = 0.012003 * x | 0.99495 | 23.4286 | 0 |  | 0 | 13.6721 | 424163.2534
| UCG1 T5 S1C1009_SE 11/1 |  | 110119031.D | Sample |  | 2019-11-02 00:54 | y = 0.098018 * x | 0.99826 | 9.1233 | 271.0289 |  | 449995.1365 | 13.6721 | 423473.522 | y = 0.092110 * x | 0.99679 | 9.1996 | 269.9581 |  | 421201.6636 | 13.6721 | 423473.522 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6093 | 0 |  | 0 | 13.6721 | 423473.522 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6816 | 172.3761 |  | 889123.1218 | 10.6807 | 164021.6837 | y = 0.079567 * x | 0.99598 | 11.1146 | 0 |  | 0 | 13.6721 | 423473.522 | y = 0.012003 * x | 0.99495 | 23.4285 | 0 |  | 0 | 13.6721 | 423473.522
| UCG1 AF S2A1009_SE 11/1 |  | 110119032.D | Sample |  | 2019-11-02 01:24 | y = 0.098018 * x | 0.99826 | 9.1178 | 6.461 |  | 8688.3589 | 13.6721 | 342982.8611 | y = 0.092110 * x | 0.99679 | 9.1941 | 5.2396 |  | 6621.2101 | 13.6721 | 342982.8611 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.571 | 1.1559 |  | 532.1175 | 13.6721 | 342982.8611 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6816 | 44.5587 |  | 242562.1505 | 10.6807 | 151050.1979 | y = 0.079567 * x | 0.99598 | 11.8192 | 430.5008 |  | 469936.6047 | 13.6721 | 342982.8611 | y = 0.012003 * x | 0.99495 | 23.4456 | 686.3662 |  | 113030.3517 | 13.6721 | 342982.8611
| UCG1 AF S2B1009_SE 11/1 |  | 110119033.D | Sample |  | 2019-11-02 01:54 | y = 0.098018 * x | 0.99826 | 9.1176 | 8.7247 |  | 12346.9907 | 13.6719 | 360946.0322 | y = 0.092110 * x | 0.99679 | 9.1939 | 6.4108 |  | 8525.5648 | 13.6719 | 360946.0322 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5763 | 1.3931 |  | 674.8524 | 13.6719 | 360946.0322 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 30.8072 |  | 180442.1819 | 10.6806 | 150947.113 | y = 0.079567 * x | 0.99598 | 11.8191 | 393.3236 |  | 451840.4532 | 13.6719 | 360946.0322 | y = 0.012003 * x | 0.99495 | 23.4369 | 582.0857 |  | 100877.8808 | 13.6719 | 360946.0322
| UCG1 AF S2C1009_SE 11/1 |  | 110119034.D | Sample |  | 2019-11-02 02:24 | y = 0.098018 * x | 0.99826 | 9.1177 | 5.9204 |  | 12399.6181 | 13.672 | 534187.2639 | y = 0.092110 * x | 0.99679 | 9.194 | 3.5444 |  | 6976.0025 | 13.672 | 534187.2639 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5764 | 0.93307 |  | 668.9819 | 13.672 | 534187.2639 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 79.6546 |  | 519311.6306 | 10.6806 | 195720.3545 | y = 0.079567 * x | 0.99598 | 11.8192 | 1035.5055 |  | 1760511.1883 | 13.672 | 534187.2639 | y = 0.012003 * x | 0.99495 | 23.437 | 1498.979 |  | 384464.1369 | 13.672 | 534187.2639
| UCG1 T5 S2A1009_SE 11/1 |  | 110119035.D | Sample |  | 2019-11-02 02:54 | y = 0.098018 * x | 0.99826 | 9.0903 | 0 |  | 0 | 13.672 | 408263.7418 | y = 0.092110 * x | 0.99679 | 9.1885 | 0 |  | 0 | 13.672 | 408263.7418 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 0 |  | 0 | 13.672 | 408263.7418 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 101.7128 |  | 497130.8226 | 10.6807 | 150113.515 | y = 0.079567 * x | 0.99598 | 11.8192 | 129.6863 |  | 168510.8134 | 13.672 | 408263.7418 | y = 0.012003 * x | 0.99495 | 23.437 | 146.0965 |  | 28638.3171 | 13.672 | 408263.7418
| UCG1 T5 S2B1009_SE 11/1 |  | 110119037.D | Sample |  | 2019-11-02 03:54 | y = 0.098018 * x | 0.99826 | 9.0958 | 0 |  | 0 | 13.672 | 369058.9658 | y = 0.092110 * x | 0.99679 | 9.194 | 0 |  | 0 | 13.672 | 369058.9658 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 0 |  | 0 | 13.672 | 369058.9658 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 98.6321 |  | 486340.4102 | 10.6806 | 151049.0995 | y = 0.079567 * x | 0.99598 | 11.8191 | 137.0784 |  | 161011.73 | 13.672 | 369058.9658 | y = 0.012003 * x | 0.99495 | 23.4369 | 136.2061 |  | 24135.6563 | 13.672 | 369058.9658
| UCG1 T5 S2B1009_SE 11/1b |  | 110119038.D | Sample |  | 2019-11-02 04:24 | y = 0.098018 * x | 0.99826 | 9.0903 | 0 |  | 0 | 13.672 | 363600.9065 | y = 0.092110 * x | 0.99679 | 9.194 | 0 |  | 0 | 13.672 | 363600.9065 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 0 |  | 0 | 13.672 | 363600.9065 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 96.9814 |  | 478385.6395 | 10.6806 | 150887.3284 | y = 0.079567 * x | 0.99598 | 11.8192 | 152.3338 |  | 176284.4367 | 13.672 | 363600.9065 | y = 0.012003 * x | 0.99495 | 23.4285 | 125.9337 |  | 21985.3672 | 13.672 | 363600.9065
| UCG1 T5 S2C1009_SE 11/1 |  | 110119039.D | Sample |  | 2019-11-02 04:54 | y = 0.098018 * x | 0.99826 | 9.1122 | 0 |  | 0 | 13.6635 | 614730.4797 | y = 0.092110 * x | 0.99679 | 9.1721 | 0 |  | 0 | 13.6635 | 614730.4797 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6202 | 0 |  | 0 | 13.6635 | 614730.4797 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 98.896 |  | 706413.6929 | 10.6756 | 218864.8269 | y = 0.079567 * x | 0.99598 | 11.8106 | 128.7544 |  | 251906.6359 | 13.6635 | 614730.4797 | y = 0.012003 * x | 0.99495 | 23.4285 | 126.5049 |  | 37338.6786 | 13.6635 | 614730.4797
| UCG1 T1 S2A1009_SE 11/1 |  | 110119040.D | Sample |  | 2019-11-02 05:24 | y = 0.098018 * x | 0.99826 | 9.0519 | 0 |  | 0 | 13.672 | 564324.0678 | y = 0.092110 * x | 0.99679 | 9.194 | 0 |  | 0 | 13.672 | 564324.0678 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6092 | 0 |  | 0 | 13.672 | 564324.0678 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 110.7876 |  | 671439.8281 | 10.6756 | 187419.3664 | y = 0.079567 * x | 0.99598 | 11.8106 | 123.6885 |  | 222152.349 | 13.672 | 564324.0678 | y = 0.012003 * x | 0.99495 | 23.437 | 102.0511 |  | 27651.1347 | 13.672 | 564324.0678
| UCG1 T1 S2B1009_SE 11/1 |  | 110119041.D | Sample |  | 2019-11-02 05:54 | y = 0.098018 * x | 0.99826 | 9.0903 | 0 |  | 0 | 13.672 | 366161.1609 | y = 0.092110 * x | 0.99679 | 9.1885 | 0 |  | 0 | 13.672 | 366161.1609 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5764 | 0 |  | 0 | 13.672 | 366161.1609 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 119.9381 |  | 571204.1421 | 10.6756 | 148146.7868 | y = 0.079567 * x | 0.99598 | 11.8107 | 148.6199 |  | 173197.7278 | 13.672 | 366161.1609 | y = 0.012003 * x | 0.99495 | 23.4286 | 126.2221 |  | 22190.8773 | 13.672 | 366161.1609
| UCG1 T1 S2C1009_SE 11/1 |  | 110119043.D | Sample |  | 2019-11-02 06:54 | y = 0.098018 * x | 0.99826 | 9.0903 | 0 |  | 0 | 13.6635 | 360011.2422 | y = 0.092110 * x | 0.99679 | 9.1885 | 0 |  | 0 | 13.6635 | 360011.2422 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 0 |  | 0 | 13.6635 | 360011.2422 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 120.811 |  | 559393.2953 | 10.6756 | 144109.7976 | y = 0.079567 * x | 0.99598 | 11.8107 | 142.4199 |  | 163184.7693 | 13.6635 | 360011.2422 | y = 0.012003 * x | 0.99495 | 23.437 | 123.5326 |  | 21353.2722 | 13.6635 | 360011.2422
| UCG1 T1 S3A1009_SE 11/1 |  | 110119044.D | Sample |  | 2019-11-02 07:24 | y = 0.098018 * x | 0.99826 | 9.5671 | 0 |  | 0 | 13.6635 | 527030.0409 | y = 0.092110 * x | 0.99679 | 9.5722 | 0 |  | 0 | 13.6635 | 527030.0409 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 257.0094 |  | 177793.581 | 13.6635 | 527030.0409 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 201.0527 |  | 1177350.1613 | 10.6756 | 187575.633 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 527030.0409 | y = 0.012003 * x | 0.99495 | 23.9973 | 0 |  | 0 | 13.6635 | 527030.0409
| UCG1 T1 S3B1009_SE 11/1 |  | 110119045.D | Sample |  | 2019-11-02 07:54 | y = 0.098018 * x | 0.99826 | 9.5726 | 0 |  | 0 | 13.6635 | 441925.4146 | y = 0.092110 * x | 0.99679 | 9.5722 | 0 |  | 0 | 13.6635 | 441925.4146 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 167.9131 |  | 98164.9868 | 13.6635 | 441925.4146 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 183.8573 |  | 766588.8586 | 10.6756 | 133008.8278 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 441925.4146 | y = 0.012003 * x | 0.99495 | 23.8105 | 0 |  | 0 | 13.6635 | 441925.4146
| UCG1 T1 S3C1009_SE 11/1 |  | 110119046.D | Sample |  | 2019-11-02 08:24 | y = 0.098018 * x | 0.99826 | 9.5726 | 0 |  | 0 | 13.6635 | 439771.673 | y = 0.092110 * x | 0.99679 | 9.5722 | 0 |  | 0 | 13.6635 | 439771.673 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 170.7314 |  | 99301.7133 | 13.6635 | 439771.673 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 184.173 |  | 838277.2871 | 10.6756 | 145209.9126 | y = 0.079567 * x | 0.99598 | 11.556 | 0 |  | 0 | 13.6635 | 439771.673 | y = 0.012003 * x | 0.99495 | 23.5134 | 0 |  | 0 | 13.6635 | 439771.673
| UCG1 AF S3A1009_SE 11/1 |  | 110119047.D | Sample |  | 2019-11-02 08:54 | y = 0.098018 * x | 0.99826 | 9.5672 | 0 |  | 0 | 13.6635 | 620126.5842 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 620126.5842 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 203.9452 |  | 166781.7408 | 13.6635 | 620126.5842 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 155.6529 |  | 933862.6116 | 10.6756 | 189746.5236 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 620126.5842 | y = 0.012003 * x | 0.99495 | 23.5388 | 0 |  | 0 | 13.6635 | 620126.5842
| UCG1 AF S3B1009_SE 11/1 |  | 110119050.D | Sample |  | 2019-11-02 10:25 | y = 0.098018 * x | 0.99826 | 9.5671 | 0 |  | 0 | 13.6635 | 651865.6895 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 651865.6895 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 208.5538 |  | 179207.2482 | 13.6635 | 651865.6895 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 163.7522 |  | 1029595.959 | 10.6756 | 199403.7787 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 651865.6895 | y = 0.012003 * x | 0.99495 | 23.5728 | 0 |  | 0 | 13.6635 | 651865.6895
| UCG1 AF S3C1009_SE 11/1 |  | 110119051.D | Sample |  | 2019-11-02 10:55 | y = 0.098018 * x | 0.99826 | 9.5671 | 0 |  | 0 | 13.6635 | 468168.9533 | y = 0.092110 * x | 0.99679 | 9.5722 | 0 |  | 0 | 13.6635 | 468168.9533 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 223.1103 |  | 137514.1801 | 13.6635 | 468168.9533 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 181.1785 |  | 887007.0383 | 10.6756 | 156067.3576 | y = 0.079567 * x | 0.99598 | 11.8192 | 0 |  | 0 | 13.6635 | 468168.9533 | y = 0.012003 * x | 0.99495 | 23.5728 | 0 |  | 0 | 13.6635 | 468168.9533
| UCG1 T5 S3A1009_SE 11/1 |  | 110119052.D | Sample |  | 2019-11-02 11:25 | y = 0.098018 * x | 0.99826 | 9.1122 | 3.2007 |  | 8863.6328 | 13.6635 | 706325.4339 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 706325.4339 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 16.6986 |  | 15808.9676 | 13.6635 | 706325.4339 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 58.8873 |  | 427192.0114 | 10.6756 | 210077.4583 | y = 0.079567 * x | 0.99598 | 11.2249 | 0 |  | 0 | 13.6635 | 706325.4339 | y = 0.012003 * x | 0.99495 | 23.5643 | 0 |  | 0 | 13.6635 | 706325.4339
| UCG1 T5 S3B1009_SE 11/1 |  | 110119053.D | Sample |  | 2019-11-02 11:55 | y = 0.098018 * x | 0.99826 | 9.1122 | 4.5012 |  | 11278.0211 | 13.6635 | 639054.8178 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 639054.8178 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 18.2583 |  | 15637.1878 | 13.6635 | 639054.8178 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 56.8794 |  | 377330.525 | 10.6756 | 191191.966 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.6635 | 639054.8178 | y = 0.012003 * x | 0.99495 | 23.4455 | 0 |  | 0 | 13.6635 | 639054.8178
| UCG1 T5 S3B1009_SE 11/1b |  | 110119054.D | Sample |  | 2019-11-02 12:25 | y = 0.098018 * x | 0.99826 | 9.1177 | 5.0556 |  | 8417.0068 | 13.6635 | 424636.5077 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 424636.5077 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 19.0984 |  | 10867.837 | 13.6635 | 424636.5077 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 54.9153 |  | 282648.9865 | 10.6756 | 147601.3645 | y = 0.079567 * x | 0.99598 | 11.7767 | 0 |  | 0 | 13.6635 | 424636.5077 | y = 0.012003 * x | 0.99495 | 23.5389 | 0 |  | 0 | 13.6635 | 424636.5077
| UCG1 T5 S3C1009_SE 11/1 |  | 110119056.D | Sample |  | 2019-11-02 13:25 | y = 0.098018 * x | 0.99826 | 9.1122 | 5.2492 |  | 12932.566 | 13.6635 | 628387.9292 | y = 0.092110 * x | 0.99679 | 9.5667 | 0 |  | 0 | 13.6635 | 628387.9292 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 19.2405 |  | 16201.9494 | 13.6635 | 628387.9292 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 57.6741 |  | 385361.9572 | 10.6706 | 192942.5712 | y = 0.079567 * x | 0.99598 | 11.1145 | 0 |  | 0 | 13.6635 | 628387.9292 | y = 0.012003 * x | 0.99495 | 23.42 | 0 |  | 0 | 13.6635 | 628387.9292
| UCG1 AF S1A809_Cr10/11_SE 11/1 |  | 110119057.D | Sample |  | 2019-11-02 13:55 | y = 0.098018 * x | 0.99826 | 9.1122 | 973.9347 |  | 1230182.7026 | 13.6635 | 322161.635 | y = 0.092110 * x | 0.99679 | 9.1885 | 401.2132 |  | 476229.7027 | 13.6635 | 322161.635 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 0.71744 |  | 310.2248 | 13.6635 | 322161.635 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 17.2796 |  | 110571.8962 | 10.6756 | 139673.3233 | y = 0.079567 * x | 0.99598 | 11.8106 | 0 |  | 0 | 13.6635 | 322161.635 | y = 0.012003 * x | 0.99495 | 23.6237 | 0 |  | 0 | 13.6635 | 322161.635
| UCG1 AF S1B809_Cr10/11_SE 11/1 |  | 110119058.D | Sample |  | 2019-11-02 14:25 | y = 0.098018 * x | 0.99826 | 9.1122 | 824.12 |  | 1034935.4501 | 13.6634 | 320299.9182 | y = 0.092110 * x | 0.99679 | 9.1885 | 305.3023 |  | 360291.8609 | 13.6634 | 320299.9182 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 3.222 |  | 1384.8729 | 13.6634 | 320299.9182 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6764 | 4.6223 |  | 61118.6726 | 10.6706 | 147676.7824 | y = 0.079567 * x | 0.99598 | 11.8106 | 0 |  | 0 | 13.6634 | 320299.9182 | y = 0.012003 * x | 0.99495 | 23.2332 | 0 |  | 0 | 13.6634 | 320299.9182
| UCG1 AF S1C809_Cr10/11_SE 11/1 |  | 110119059.D | Sample |  | 2019-11-02 14:55 | y = 0.098018 * x | 0.99826 | 9.1122 | 771.631 |  | 1613130.7681 | 13.6635 | 533204.625 | y = 0.092110 * x | 0.99679 | 9.1885 | 283.3674 |  | 556687.3293 | 13.6635 | 533204.625 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 0.69999 |  | 500.9563 | 13.6635 | 533204.625 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 1.167 |  | 59772.8471 | 10.6706 | 192358.5436 | y = 0.079567 * x | 0.99598 |  |  |  |  | 13.6635 | 533204.625 | y = 0.012003 * x | 0.99495 | 23.5813 | 0 |  | 0 | 13.6635 | 533204.625
| UCG1 T5 S1A809_Cr10/11_SE 11/1 |  | 110119060.D | Sample |  | 2019-11-02 15:25 | y = 0.098018 * x | 0.99826 | 9.0848 | 0 |  | 0 | 13.6635 | 494367.9488 | y = 0.092110 * x | 0.99679 | 9.2159 | 0 |  | 0 | 13.6635 | 494367.9488 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5215 | 0 |  | 0 | 13.6635 | 494367.9488 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 5.2859 |  | 81984.8006 | 10.6706 | 189047.6467 | y = 0.079567 * x | 0.99598 | 11.8106 | 287.7469 |  | 452745.2014 | 13.6635 | 494367.9488 | y = 0.012003 * x | 0.99495 | 23.3945 | 221.5002 |  | 52576.4299 | 13.6635 | 494367.9488
| UCG1 T5 S1B809_Cr10/11_SE 11/1 |  | 110119062.D | Sample |  | 2019-11-02 16:26 | y = 0.098018 * x | 0.99826 | 9.0903 | 0 |  | 0 | 13.6635 | 370641.3009 | y = 0.092110 * x | 0.99679 | 9.3091 | 0 |  | 0 | 13.6635 | 370641.3009 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5709 | 0 |  | 0 | 13.6635 | 370641.3009 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 0 |  | 30288.2043 | 10.6706 | 144791.0831 | y = 0.079567 * x | 0.99598 | 11.8107 | 286.4716 |  | 337931.1849 | 13.6635 | 370641.3009 | y = 0.012003 * x | 0.99495 | 23.4031 | 216.2494 |  | 38483.573 | 13.6635 | 370641.3009
| UCG1 T5 S1C809_Cr10/11_SE 11/1 |  | 110119063.D | Sample |  | 2019-11-02 16:56 | y = 0.098018 * x | 0.99826 | 8.7559 | 0 |  | 0 | 13.6635 | 580453.6819 | y = 0.092110 * x | 0.99679 | 9.1995 | 0 |  | 0 | 13.6635 | 580453.6819 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.8559 | 0 |  | 0 | 13.6635 | 580453.6819 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 0 |  | 46192.7399 | 10.6706 | 203064.8523 | y = 0.079567 * x | 0.99598 | 11.8106 | 278.5289 |  | 514553.7408 | 13.6635 | 580453.6819 | y = 0.012003 * x | 0.99495 | 23.3945 | 229.6789 |  | 64011.1285 | 13.6635 | 580453.6819
| UCG1 T1 S1A809_Cr10/11_SE 11/1 |  | 110119064.D | Sample |  | 2019-11-02 17:26 | y = 0.098018 * x | 0.99826 | 9.1122 | 233.6474 |  | 530287.974 | 13.6635 | 578875.0385 | y = 0.092110 * x | 0.99679 | 9.1885 | 231.4498 |  | 493638.6218 | 13.6635 | 578875.0385 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.4722 | 0 |  | 0 | 13.6635 | 578875.0385 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 2.7019 |  | 72463.1572 | 10.6706 | 203234.9214 | y = 0.079567 * x | 0.99598 | 11.6409 | 0 |  | 0 | 13.6635 | 578875.0385 | y = 0.012003 * x | 0.99495 | 23.7426 | 0 |  | 0 | 13.6635 | 578875.0385
| UCG1 T1 S2A809_Cr10/11_SE 11/1 |  | 110119065.D | Sample |  | 2019-11-02 17:56 | y = 0.098018 * x | 0.99826 | 9.0848 | 0 |  | 0 | 13.6635 | 401228.0652 | y = 0.092110 * x | 0.99679 | 9.1831 | 0 |  | 0 | 13.6635 | 401228.0652 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.56 | 0 |  | 0 | 13.6635 | 401228.0652 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 10.5643 |  | 84303.1892 | 10.6706 | 142592.9773 | y = 0.079567 * x | 0.99598 | 11.8107 | 263.0389 |  | 335895.4595 | 13.6635 | 401228.0652 | y = 0.012003 * x | 0.99495 | 23.3946 | 229.6356 |  | 44238.1847 | 13.6635 | 401228.0652
| UCG1 T5 S2A809_Cr10/11_SE 11/1 |  | 110119067.D | Sample |  | 2019-11-02 18:56 | y = 0.098018 * x | 0.99826 | 9.5617 | 0 |  | 0 | 13.6635 | 490105.9968 | y = 0.092110 * x | 0.99679 | 9.5612 | 0 |  | 0 | 13.6635 | 490105.9968 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 254.7921 |  | 163942.8204 | 13.6635 | 490105.9968 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 10.7441 |  | 121908.1835 | 10.6706 | 204344.6072 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 490105.9968 | y = 0.012003 * x | 0.99495 | 23.7425 | 0 |  | 0 | 13.6635 | 490105.9968
| UCG1 T5 S2B809_Cr10/11_SE 11/1 |  | 110119068.D | Sample |  | 2019-11-02 19:26 | y = 0.098018 * x | 0.99826 | 9.5617 | 0 |  | 0 | 13.6635 | 504000.8652 | y = 0.092110 * x | 0.99679 | 9.5612 | 0 |  | 0 | 13.6635 | 504000.8652 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 205.7586 |  | 136733.5598 | 13.6635 | 504000.8652 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 0 |  | 51087.1604 | 10.6706 | 190248.1809 | y = 0.079567 * x | 0.99598 | 11.2334 | 0 |  | 0 | 13.6635 | 504000.8652 | y = 0.012003 * x | 0.99495 | 23.4284 | 0 |  | 0 | 13.6635 | 504000.8652
| UCG1 T5 S2C809_Cr10/11_SE 11/1 |  | 110119071.D | Sample |  | 2019-11-02 20:56 | y = 0.098018 * x | 0.99826 | 9.5617 | 0 |  | 0 | 13.6635 | 526940.9493 | y = 0.092110 * x | 0.99679 | 9.5612 | 0 |  | 0 | 13.6635 | 526940.9493 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5654 | 239.019 |  | 165581.9737 | 13.6635 | 526940.9493 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6714 | 0 |  | 51814.9939 | 10.6706 | 198905.131 | y = 0.079567 * x | 0.99598 | 11.2249 | 0 |  | 0 | 13.6635 | 526940.9493 | y = 0.012003 * x | 0.99495 | 23.5727 | 0 |  | 0 | 13.6635 | 526940.9493
| UCG1 AF S2A809_Cr10/11_SE 11/1 |  | 110119072.D | Sample |  | 2019-11-02 21:26 | y = 0.098018 * x | 0.99826 | 9.1067 | 8.962 |  | 21776.2184 | 13.6635 | 619744.4736 | y = 0.092110 * x | 0.99679 | 9.183 | 7.0328 |  | 16058.6068 | 13.6635 | 619744.4736 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 0.68389 |  | 568.8734 | 13.6635 | 619744.4736 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6714 | 22.2736 |  | 212837.9363 | 10.6706 | 226255.2421 | y = 0.079567 * x | 0.99598 | 11.8021 | 730.7292 |  | 1441325.2033 | 13.6635 | 619744.4736 | y = 0.012003 * x | 0.99495 | 23.3775 | 907.1423 |  | 269932.2708 | 13.6635 | 619744.4736
| UCG1 AF S2B809_Cr10/11_SE 11/1 |  | 110119073.D | Sample |  | 2019-11-02 21:56 | y = 0.098018 * x | 0.99826 | 9.1067 | 9.2072 |  | 22174.1205 | 13.6635 | 614263.46 | y = 0.092110 * x | 0.99679 | 9.183 | 6.5637 |  | 14854.9411 | 13.6635 | 614263.46 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 0.52404 |  | 432.0623 | 13.6635 | 614263.46 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6714 | 0 |  | 40617.2233 | 10.6655 | 222691.6946 | y = 0.079567 * x | 0.99598 | 11.8022 | 718.2386 |  | 1404158.9553 | 13.6635 | 614263.46 | y = 0.012003 * x | 0.99495 | 23.3775 | 980.0344 |  | 289043.1965 | 13.6635 | 614263.46
| UCG1 AF S2C809_Cr10/11_SE 11/1 |  | 110119074.D | Sample |  | 2019-11-02 22:27 | y = 0.098018 * x | 0.99826 | 9.1067 | 9.6814 |  | 22673.0132 | 13.6635 | 597320.1248 | y = 0.092110 * x | 0.99679 | 9.183 | 6.4335 |  | 14158.6454 | 13.6635 | 597320.1248 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 0.83593 |  | 670.1737 | 13.6635 | 597320.1248 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6714 | 1.3364 |  | 70303.4042 | 10.6655 | 222624.948 | y = 0.079567 * x | 0.99598 | 11.8021 | 650.6462 |  | 1236929.4061 | 13.6635 | 597320.1248 | y = 0.012003 * x | 0.99495 | 23.3774 | 873.0413 |  | 250385.2156 | 13.6635 | 597320.1248
| UCG1 T1 S3A809_Cr10/11_SE 11/1 |  | 110119075.D | Sample |  | 2019-11-02 22:57 | y = 0.098018 * x | 0.99826 | 9.5617 | 0 |  | 0 | 13.6635 | 499480.2723 | y = 0.092110 * x | 0.99679 | 9.5612 | 0 |  | 0 | 13.6635 | 499480.2723 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 246.3612 |  | 161669.8939 | 13.6635 | 499480.2723 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6714 | 11.8687 |  | 110410.6593 | 10.6655 | 175213.8535 | y = 0.079567 * x | 0.99598 | 11.8021 | 0 |  | 0 | 13.6635 | 499480.2723 | y = 0.012003 * x | 0.99495 | 23.352 | 8.2755 |  | 1984.6387 | 13.6635 | 499480.2723
| UCG1 T5 S3A809_Cr10/11_SE 11/1 |  | 110119077.D | Sample |  | 2019-11-02 23:57 | y = 0.098018 * x | 0.99826 | 9.1068 | 244.0287 |  | 426539.9109 | 13.6636 | 445813.1129 | y = 0.092110 * x | 0.99679 | 9.1831 | 259.7598 |  | 426670.158 | 13.6636 | 445813.1129 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.6258 | 0 |  | 0 | 13.6636 | 445813.1129 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6715 | 10.9001 |  | 112921.4479 | 10.6656 | 187814.7128 | y = 0.079567 * x | 0.99598 | 11.5561 | 0 |  | 0 | 13.6636 | 445813.1129 | y = 0.012003 * x | 0.99495 | 23.055 | 0 |  | 0 | 13.6636 | 445813.1129
| UCG1 T5 S3B809_Cr10/11_SE 11/1 |  | 110119078.D | Sample |  | 2019-11-03 00:27 | y = 0.098018 * x | 0.99826 | 9.1068 | 248.3054 |  | 452974.0248 | 13.6636 | 465287.1648 | y = 0.092110 * x | 0.99679 | 9.1831 | 280.2647 |  | 480459.7541 | 13.6636 | 465287.1648 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5655 | 0 |  | 0 | 13.6636 | 465287.1648 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6715 | 16.495 |  | 149625.9039 | 10.6656 | 194767.9913 | y = 0.079567 * x | 0.99598 | 11.225 | 0 |  | 0 | 13.6636 | 465287.1648 | y = 0.012003 * x | 0.99495 | 23.4455 | 0 |  | 0 | 13.6636 | 465287.1648
| UCG1 T5 S3C809_Cr10/11_SE 11/1 |  | 110119079.D | Sample |  | 2019-11-03 00:57 | y = 0.098018 * x | 0.99826 | 9.1068 | 237.0079 |  | 363087.9316 | 13.6636 | 390735.6925 | y = 0.092110 * x | 0.99679 | 9.1831 | 265.7372 |  | 382563.0137 | 13.6636 | 390735.6925 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5655 | 0 |  | 0 | 13.6636 | 390735.6925 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6766 | 0.45187 |  | 41712.8673 | 10.6656 | 144139.2744 | y = 0.079567 * x | 0.99598 | 11.5561 | 0 |  | 0 | 13.6636 | 390735.6925 | y = 0.012003 * x | 0.99495 | 23.3861 | 0 |  | 0 | 13.6636 | 390735.6925
| UCG1 AF S3A809_Cr10/11_SE 11/1 |  | 110119081.D | Sample |  | 2019-11-03 01:57 | y = 0.098018 * x | 0.99826 | 9.1067 | 6.2104 |  | 13616.1604 | 13.6634 | 559202.0555 | y = 0.092110 * x | 0.99679 | 9.183 | 4.1095 |  | 8466.9253 | 13.6634 | 559202.0555 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5599 | 48.5066 |  | 36257.3716 | 13.6634 | 559202.0555 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6663 | 0.45673 |  | 64958.1919 | 10.6655 | 224351.1645 | y = 0.079567 * x | 0.99598 | 11.2248 | 0 |  | 0 | 13.6634 | 559202.0555 | y = 0.012003 * x | 0.99495 | 23.4878 | 0 |  | 0 | 13.6634 | 559202.0555
| UCG1 AF S3B809_Cr10/11_SE 11/1 |  | 110119082.D | Sample |  | 2019-11-03 02:02 | y = 0.098018 * x | 0.99826 | 9.1068 | 7.8558 |  | 17257.8605 | 13.6636 | 560309.8209 | y = 0.092110 * x | 0.99679 | 9.1831 | 4.7326 |  | 9770.0278 | 13.6636 | 560309.8209 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5655 | 18.0146 |  | 13527.6378 | 13.6636 | 560309.8209 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6715 | 16.1652 |  | 167414.3371 | 10.6656 | 220751.468 | y = 0.079567 * x | 0.99598 | 11.2335 | 0 |  | 0 | 13.6636 | 560309.8209 | y = 0.012003 * x | 0.99495 | 23.5559 | 0 |  | 0 | 13.6636 | 560309.8209
| UCG1 AF S3C809_Cr10/11_SE 11/1 |  | 110119083.D | Sample |  | 2019-11-03 02:32 | y = 0.098018 * x | 0.99826 | 9.1068 | 8.0625 |  | 16040.1912 | 13.6636 | 507424.7445 | y = 0.092110 * x | 0.99679 | 9.1831 | 5.6186 |  | 10504.2981 | 13.6636 | 507424.7445 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.56 | 30.3305 |  | 20604.3937 | 13.6636 | 507424.7445 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6665 | 8.8437 |  | 113727.8784 | 10.6656 | 210661.5768 | y = 0.079567 * x | 0.99598 | 11.225 | 0 |  | 0 | 13.6636 | 507424.7445 | y = 0.012003 * x | 0.99495 | 23.6492 | 0 |  | 0 | 13.6636 | 507424.7445
| UCG1 T5 S1B1009_SE 11/1b |  | 110119084.D | Sample |  | 2019-11-03 03:02 | y = 0.098018 * x | 0.99826 | 9.1013 | 204.8619 |  | 2160616.3235 | 13.655 | 2689989.4467 | y = 0.092110 * x | 0.99679 | 9.1831 | 184.0571 |  | 1824192.1322 | 13.655 | 2689989.4467 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.549 | 0 |  | 0 | 13.655 | 2689989.4467 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6664 | 151.5848 |  | 3322700.1887 | 10.6655 | 692198.2791 | y = 0.079567 * x | 0.99598 | 11.2249 | 0 |  | 0 | 13.655 | 2689989.4467 | y = 0.012003 * x | 0.99495 | 23.4369 | 0 |  | 0 | 13.655 | 2689989.4467
| UCG1 AF S2C1009_SE 11/1b |  | 110119085.D | Sample |  | 2019-11-03 03:32 | y = 0.098018 * x | 0.99826 | 9.1068 | 6.9597 |  | 18141.5924 | 13.6636 | 664838.2138 | y = 0.092110 * x | 0.99679 | 9.1831 | 4.6641 |  | 11424.9177 | 13.6636 | 664838.2138 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5545 | 0.95825 |  | 855.0741 | 13.6636 | 664838.2138 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6665 | 86.8912 |  | 744769.2679 | 10.6656 | 259562.6651 | y = 0.079567 * x | 0.99598 | 11.8023 | 1192.2075 |  | 2522671.3635 | 13.6636 | 664838.2138 | y = 0.012003 * x | 0.99495 | 23.3521 | 1795.0341 |  | 573001.0738 | 13.6636 | 664838.2138
| UCG1 T1 S3A1009_SE 11/1b |  | 110119086.D | Sample |  | 2019-11-03 04:02 | y = 0.098018 * x | 0.99826 | 9.5563 | 0 |  | 0 | 13.6636 | 624626.3482 | y = 0.092110 * x | 0.99679 | 9.5558 | 0 |  | 0 | 13.6636 | 624626.3482 | y = -7.218266E-005 * x ^ 2 + 0.033557 * x | 0.99906 | 9.5545 | 275.0152 |  | 225123.1725 | 13.6636 | 624626.3482 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6665 | 207.9132 |  | 1520491.7395 | 10.6656 | 234591.9793 | y = 0.079567 * x | 0.99598 | 11.5561 | 0 |  | 0 | 13.6636 | 624626.3482 | y = 0.012003 * x | 0.99495 | 23.3521 | 9.5154 |  | 2853.7303 | 13.6636 | 624626.3482
Note on D71: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S2
Note on D72: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S2
Note on D73: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S2
Note on D76: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S3
Note on D77: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S3
Note on D78: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S3
Note on D83: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S1
Note on D84: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S1
Note on D85: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this is actually S1

## 121019RawOutput
Sample | [6 empty] | 3117 Method |  | 3117 Results |  |  |  | M8FOSA (ISTD) Results |  | 4NT Method |  | 4NT Results |  |  |  | 4NT13C6 (ISTD) Results |  | 916 Method |  | 916 Results |  |  |  | M8FOSA (ISTD) Results |  | 923 Method |  | 923 Results |  |  |  | M8FOSA (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula Linear 1/X | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula Linear | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula Linear | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area | CF Formula Linear 1/X | CF R2 | RT | Final Conc. | Accuracy | Area | RT | Area
| Spike Blank |  | 121019000.D | MatrixBlank |  | 2019-12-10 11:49 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5974 | 0 | 0 | 0 | 13.6638 | 491753.8655 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6768 | 0.48523 |  | 4563.1816 | 10.6558 | 117368.6483 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7685 | 0 | 0 | 0 | 13.6638 | 491753.8655 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2506 | 0 | 0 | 0 | 13.6638 | 491753.8655
| CC1 | 923 dropped | 121019001.D | Cal | 1 | 2019-12-10 12:19 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 2.5571 | 146.1191 | 1885.7691 | 13.6549 | 979157.4259 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6664 | 1.7489 | 99.9357 | 18833.0041 | 10.6554 | 247582.6763 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7936 | 2.7165 | 155.2282 | 5483.255 | 13.6549 | 979157.4259 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.3266 | 0 | 0 | 0 | 13.6549 | 979157.4259
| CC2 | 924 dropped | 121019002.D | Cal | 2 | 2019-12-10 12:49 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 3.8896 | 129.6533 | 3199.0797 | 13.655 | 953360.9038 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 3.4807 | 116.0226 | 27327.5595 | 10.6504 | 215120.7671 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 3.0018 | 100.0602 | 6068.8063 | 13.655 | 953360.9038 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2758 | 8.0549 | 268.4955 | 818.7871 | 13.655 | 953360.9038
| CC3 | 925 dropped | 121019003.D | Cal | 3 | 2019-12-10 13:19 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 5.1468 | 102.9368 | 3413.9759 | 13.655 | 725684.6679 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6564 | 3.7488 | 74.9759 | 27119.8678 | 10.6505 | 201000.9099 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 6.2341 | 124.6825 | 10914.7611 | 13.655 | 725684.6679 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.3267 | 8.9021 | 178.0417 | 889.8228 | 13.655 | 725684.6679
| CC4 | 926 dropped | 121019004.D | Cal | 4 | 2019-12-10 13:49 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 6.5038 | 86.7173 | 4084.3982 | 13.655 | 663010.6359 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 7.7455 | 103.2737 | 37729.5151 | 10.6504 | 149396.6597 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 7.8112 | 104.1493 | 12778.3441 | 13.655 | 663010.6359 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 11.1702 | 148.9365 | 1465.0001 | 13.655 | 663010.6359
| CC5 |  | 121019005.D | Cal | 5 | 2019-12-10 14:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 12.1269 | 97.0152 | 10962.979 | 13.655 | 899085.7336 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 12.0743 | 96.5944 | 83410.5058 | 10.6504 | 219536.3448 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 11.61 | 92.8799 | 26494.6673 | 13.655 | 899085.7336 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 14.4892 | 115.9138 | 3280.4734 | 13.655 | 899085.7336
| CC6 |  | 121019006.D | Cal | 6 | 2019-12-10 14:50 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 19.5691 | 97.8457 | 17467.3029 | 13.655 | 865646.067 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6563 | 20.1498 | 100.749 | 131467.2163 | 10.6505 | 212868.9604 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 18.8054 | 94.0268 | 42225.7558 | 13.655 | 865646.067 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 18.3178 | 91.5889 | 4595.4328 | 13.655 | 865646.067
| CC7 |  | 121019007.D | Cal | 7 | 2019-12-10 15:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 24.9585 | 79.8672 | 26536.6045 | 13.655 | 1022186.2036 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 31.131 | 99.6193 | 217947.4038 | 10.6505 | 231669.4056 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 29.6791 | 94.973 | 79692.1788 | 13.655 | 1022186.2036 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 29.4397 | 94.2071 | 10355.7322 | 13.655 | 1022186.2036
| CC8 |  | 121019008.D | Cal | 8 | 2019-12-10 15:50 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 39.0485 | 78.0971 | 40379.0409 | 13.655 | 982984.4922 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 48.9734 | 97.9468 | 317843.3051 | 10.6504 | 216830.6594 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7937 | 43.9717 | 87.9435 | 114342.0544 | 13.655 | 982984.4922 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 47.5962 | 95.1925 | 17696.9954 | 13.655 | 982984.4922
| CC9 | Dropped | 121019009.D | Cal | 9 | 2019-12-10 16:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 50.253 | 57.4321 | 50092.2303 | 13.655 | 943366.2225 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 49.5604 | 56.6405 | 305486.2456 | 10.6505 | 205973.228 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 46.9298 | 53.634 | 117222.8243 | 13.655 | 943366.2225 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2927 | 49.3522 | 56.4025 | 17701.9684 | 13.655 | 943366.2225
| CC10 |  | 121019010.D | Cal | 10 | 2019-12-10 16:50 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 103.6655 | 82.9324 | 106195.0945 | 13.655 | 961857.3725 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 139.4622 | 111.5698 | 847265.8404 | 10.6504 | 205203.8892 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 126.1786 | 100.9429 | 324097.2095 | 13.655 | 961857.3725 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 127.9118 | 102.3295 | 50811.8843 | 13.655 | 961857.3725
| CC11 |  | 121019011.D | Cal | 11 | 2019-12-10 17:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 189.0365 | 94.5182 | 207854.7327 | 13.6465 | 1028974.6696 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 201.4194 | 100.7097 | 1379082.3207 | 10.6504 | 231690.294 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 193.5381 | 96.769 | 532730.5919 | 13.6465 | 1028974.6696 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 196.7748 | 98.3874 | 85080.4331 | 13.6465 | 1028974.6696
| CC12 | Dropped | 121019012.D | Cal | 12 | 2019-12-10 17:50 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 42.9126 | 11.4434 | 2448.6758 | 13.621 | 54145.6959 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.611 | 282.8531 | 75.4275 | 20074.1248 | 10.6051 | 2404.4276 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 58.6272 | 15.6339 | 8427.9853 | 13.621 | 54145.6959 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 62.082 | 16.5552 | 1314.8797 | 13.621 | 54145.6959
| CC13 | Dropped | 121019013.D | Cal | 13 | 2019-12-10 18:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 728.9511 | 116.6322 | 332636.1852 | 13.6295 | 425756.6328 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 660.0401 | 105.6064 | 3278789.8082 | 10.6504 | 168581.9948 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 713.7281 | 114.1965 | 814822.5925 | 13.6295 | 425756.6328 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 360.9727 | 57.7556 | 65514.6711 | 13.6295 | 425756.6328
| CC14 |  | 121019014.D | Cal | 14 | 2019-12-10 18:50 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 884.853 | 101.1261 | 340217.1873 | 13.638 | 358670.3096 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 866.7687 | 99.0593 | 3338181.9287 | 10.6454 | 130739.1745 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 912.4574 | 104.2808 | 877728.9015 | 13.638 | 358670.3096 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 934.5931 | 106.8106 | 144397.2764 | 13.638 | 358670.3096
| CC15 |  | 121019015.D | Cal | 15 | 2019-12-10 19:20 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 1289.6447 | 103.1716 | 665791.4071 | 13.638 | 481460.7816 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 1244.2973 | 99.5438 | 5975081.8387 | 10.6454 | 163059.3166 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 1224.9961 | 97.9997 | 1582065.8988 | 13.638 | 481460.7816 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2757 | 1194.6273 | 95.5702 | 248114.6014 | 13.638 | 481460.7816
| QCGC4 |  | 121019016.D | QC | 4 | 2019-12-10 19:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 42.5908 | 567.8776 | 25728.4385 | 13.6296 | 573290.1199 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 7.216 | 96.2132 | 45181.2747 | 10.6454 | 190669.0354 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 22.4181 | 298.9075 | 33523.3204 | 13.6296 | 573290.1199 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2588 | 33.3409 | 444.5452 | 6777.6915 | 13.6296 | 573290.1199
| QCGC8 |  | 121019017.D | QC | 8 | 2019-12-10 20:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.56 | 40.1475 | 80.295 | 19140.1518 | 13.638 | 452945.0527 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 36.7427 | 73.4853 | 161648.7573 | 10.6454 | 146167.2553 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 35.2071 | 70.4143 | 42032.7567 | 13.638 | 452945.0527 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2418 | 31.1908 | 62.3816 | 4932.6659 | 13.638 | 452945.0527
| QCGC11 |  | 121019018.D | QC | 11 | 2019-12-10 20:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 122.339 | 61.1695 | 66030.4846 | 13.638 | 506208.5634 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 170.4433 | 85.2216 | 794365.8415 | 10.6454 | 157591.8097 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7851 | 158.2722 | 79.1361 | 214167.961 | 13.638 | 506208.5634 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2587 | 163.9633 | 81.9816 | 34654.1062 | 13.638 | 506208.5634
| UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 136.7077 |  | 153663.3872 | 13.6295 | 1053516.204 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 258.2701 | 102.7193 |  | 10.6454 | 199655.0197 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 175.7662 |  | 495187.5451 | 13.6295 | 1053516.204 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 198.3034 |  | 87807.9036 | 13.6295 | 1053516.204
| UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 19.4492 |  | 16527.0979 | 13.6295 | 824306.2183 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 75.8432 | 109.1143 |  | 10.6454 | 154700.2595 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 398.7124 |  | 880671.4206 | 13.6295 | 824306.2183 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 469.9016 |  | 165774.5588 | 13.6295 | 824306.2183
| Spike Blank |  | 121019021.D | MatrixBlank |  | 2019-12-10 22:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 1.4359 |  | 388.9241 | 13.638 | 537957.8968 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6614 | 0 | 0 |  | 10.6454 | 104455.2917 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 3.0836 |  | 3542.5805 | 13.638 | 537957.8968 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2587 | 32.116 |  | 6074.2681 | 13.638 | 537957.8968
| UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 15.9634 |  | 15297.0462 | 13.6295 | 937909.4467 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6462 | 69.4295 | 99.8871 |  | 10.6454 | 176795.2524 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 376.492 |  | 946110.3486 | 13.6295 | 937909.4467 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2587 | 576.8576 |  | 232116.018 | 13.6295 | 937909.4467
| CC7 |  | 121019025.D | QC | 7 | 2019-12-11 00:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 25.5443 |  | 24659.8845 | 13.6295 | 927443.0027 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6462 | 30.5138 | 97.6441 |  | 10.6454 | 215738.9873 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 26.3752 |  | 64082.0941 | 13.6295 | 927443.0027 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2671 | 39.9429 |  | 13619.4606 | 13.6295 | 927443.0027
| UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 171.2408 |  | 169880.4215 | 13.6295 | 928772.3968 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 247.3339 | 98.3697 |  | 10.6454 | 178191.595 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 176.4107 |  | 438160.3374 | 13.6295 | 928772.3968 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2586 | 188.6131 |  | 73508.5583 | 13.6295 | 928772.3968
| UC1206_S3aT1_dup |  | 121019027.D | Sample |  | 2019-12-11 01:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 172.1949 |  | 168117.5366 | 13.6296 | 914018.486 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 253.4968 | 100.8208 |  | 10.6404 | 170107.9274 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 184.7055 |  | 451547.6482 | 13.6296 | 914018.486 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2502 | 203.5649 |  | 78266.2936 | 13.6296 | 914018.486
| Spike Blank |  | 121019028.D | MatrixBlank |  | 2019-12-11 01:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5545 | 0 |  | 0 | 13.6465 | 382151.9033 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6614 | 0 | 0 |  | 10.6454 | 84732.4752 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.539 | 0 |  | 0 | 13.6465 | 382151.9033 | y = 0.010840 * x - 0.002634 | 0.99675 | 24.1925 | 0 |  | 0 | 13.6465 | 382151.9033
| UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 19.8436 |  | 19252.6372 | 13.6295 | 940399.9193 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6462 | 63.2513 | 90.9986 |  | 10.6403 | 183447.3133 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7851 | 361.1413 |  | 909879.5196 | 13.6295 | 940399.9193 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2501 | 499.2984 |  | 201108.2176 | 13.6295 | 940399.9193
| Spike Blank |  | 121019033.D | MatrixBlank |  | 2019-12-11 04:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.379 | 0 |  | 0 | 13.655 | 408624.8485 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6614 | 0 | 0 |  | 10.6403 | 91753.622 | y = 0.067095 * x - 0.001691 | 0.99885 | 12.7275 | 0 |  | 0 | 13.655 | 408624.8485 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2162 | 0 |  | 0 | 13.655 | 408624.8485
| CC9 | Cal dropped | 121019036.D | QC | 9 | 2019-12-11 05:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 47.2946 |  | 34169.2704 | 13.6465 | 684407.3146 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 49.7642 | 56.8733 |  | 10.6403 | 173337.3193 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 48.3848 |  | 87717.2338 | 13.6465 | 684407.3146 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2332 | 49.74 |  | 12957.7599 | 13.6465 | 684407.3146
| Spike Blank |  | 121019040.D | MatrixBlank |  | 2019-12-11 07:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 0 |  | 0 | 13.6466 | 507966.304 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6564 | 0 | 0 |  | 10.6404 | 110631.853 | y = 0.067095 * x - 0.001691 | 0.99885 |  |  |  |  | 13.6466 | 507966.304 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.3606 | 0 |  | 0 | 13.6466 | 507966.304
| UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 129.6975 |  | 96476.9489 | 13.655 | 697408.8117 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 210.7526 | 100.0282 |  | 10.6353 | 102491.7623 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 168.7879 |  | 314743.9191 | 13.655 | 697408.8117 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 118.3517 |  | 33951.0776 | 13.655 | 697408.8117
| Spike Blank |  | 121019045.D | MatrixBlank |  | 2019-12-11 08:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.6586 | 0 |  | 0 | 13.655 | 685191.6856 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6614 | 0 | 0 |  | 10.6353 | 138055.8895 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.1315 | 0 |  | 0 | 13.655 | 685191.6856 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2162 | 0 |  | 0 | 13.655 | 685191.6856
| UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 140.4741 |  | 67388.5699 | 13.655 | 449560.9592 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 207.4956 | 98.4823 |  | 10.6354 | 79497.9818 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 182.8553 |  | 219861.7503 | 13.655 | 449560.9592 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2333 | 110.5156 |  | 20357.9782 | 13.655 | 449560.9592
| UC1206_S3bT5_dup |  | 121019049.D | Sample |  | 2019-12-11 09:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 164.5803 |  | 78251.4839 | 13.655 | 445211.6211 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 209.6749 |  |  | 10.6353 | 78960.3988 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 186.3704 |  | 221934.7371 | 13.655 | 445211.6211 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2333 | 140.2331 |  | 25897.5874 | 13.655 | 445211.6211
| UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 159.144 |  | 149200.0719 | 13.655 | 878010.1054 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 248.6948 | 98.911 |  | 10.6353 | 177168.0431 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 207.8853 |  | 488379.8161 | 13.655 | 878010.1054 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 151.9461 |  | 55532.1209 | 13.655 | 878010.1054
| Spike Blank |  | 121019051.D | MatrixBlank |  | 2019-12-11 10:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.582 | 0 |  | 0 | 13.6551 | 745004.261 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6564 | 0 | 0 |  | 10.6354 | 133419.6406 | y = 0.067095 * x - 0.001691 | 0.99885 |  |  |  |  | 13.6551 | 745004.261 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2163 | 0 |  | 0 | 13.6551 | 745004.261
| UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 106.9233 |  | 80848.8259 | 13.6549 | 709812.9343 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6361 | 213.8314 | 101.4895 |  | 10.6353 | 95550.725 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 156.3135 |  | 296577.9907 | 13.6549 | 709812.9343 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2331 | 128.3856 |  | 37642.9772 | 13.6549 | 709812.9343
| Spike Blank |  | 121019056.D | MatrixBlank |  | 2019-12-11 11:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5654 | 0 |  | 0 | 13.655 | 575556.8197 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.5959 | 0 | 0 |  | 10.6353 | 116301.013 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7597 | 0 |  | 0 | 13.655 | 575556.8197 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2332 | 0 |  | 0 | 13.655 | 575556.8197
| CC1 |  | 121019060.D | CC | 1 | 2019-12-11 12:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 1.9538 | 111.6446 | 1210.7023 | 13.6465 | 946883.1815 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6513 | 1.928 | 110.1721 | 16370.2822 | 10.6353 | 201258.9146 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 2.9408 | 168.0484 | 5872.6617 | 13.6465 | 946883.1815 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2078 | 8.2624 | 472.1388 | 898.4394 | 13.6465 | 946883.1815
| CC2 |  | 121019061.D | CC | 2 | 2019-12-11 12:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 3.2296 | 107.6549 | 2358.7416 | 13.6465 | 890928.7989 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 3.9651 | 132.1702 | 32462.5653 | 10.6353 | 229758.3584 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 2.8541 | 95.1361 | 5318.1709 | 13.6465 | 890928.7989 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 9.4487 | 314.9555 | 1303.5816 | 13.6465 | 890928.7989
| CC3 |  | 121019062.D | CC | 3 | 2019-12-11 13:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5546 | 3.7806 | 75.6117 | 2317.8667 | 13.6467 | 715695.8278 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6414 | 3.8177 | 76.3549 | 23038.4085 | 10.6355 | 168220.8911 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 4.5308 | 90.6156 | 7492.7536 | 13.6467 | 715695.8278 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2249 | 8.9833 | 179.6665 | 902.7842 | 13.6467 | 715695.8278
| CC4 |  | 121019063.D | CC | 4 | 2019-12-11 13:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 6.7118 | 89.4904 | 4837.6756 | 13.6466 | 757836.8888 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 7.7858 | 103.8102 | 42252.0949 | 10.6354 | 166523.8981 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 8.3593 | 111.457 | 15720.6734 | 13.6466 | 757836.8888 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2163 | 11.0064 | 146.752 | 1620.6956 | 13.6466 | 757836.8888
| CC5 |  | 121019064.D | CC | 5 | 2019-12-11 14:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 8.9355 | 71.4842 | 7520.9532 | 13.6466 | 857634.7381 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 12.1984 | 97.5869 | 73000.8697 | 10.6354 | 190309.3673 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 9.6941 | 77.5524 | 20863.2004 | 13.6466 | 857634.7381 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 14.7258 | 117.8066 | 3217.2095 | 13.6466 | 857634.7381
| CC6 |  | 121019065.D | CC | 6 | 2019-12-11 14:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5491 | 17.4257 | 87.1287 | 15603.2199 | 13.6466 | 872728.6505 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 19.9157 | 99.5787 | 125514.273 | 10.6355 | 205522.3436 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 20.1966 | 100.9831 | 45829.9043 | 13.6466 | 872728.6505 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2079 | 17.9197 | 89.5986 | 4482.402 | 13.6466 | 872728.6505
| CC7 |  | 121019066.D | CC | 7 | 2019-12-11 15:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5491 | 23.8508 | 76.3226 | 24817.8692 | 13.6466 | 1001844.4368 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 30.6813 | 98.1803 | 226452.9718 | 10.6355 | 244145.0078 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 29.5067 | 94.4216 | 77642.9349 | 13.6466 | 1001844.4368 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2079 | 32.1211 | 102.7876 | 11314.4049 | 13.6466 | 1001844.4368
| CC8 |  | 121019067.D | CC | 8 | 2019-12-11 15:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 40.8962 | 81.7924 | 29348.4849 | 13.6465 | 681565.6496 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6361 | 49.7231 | 99.4462 | 236718.2708 | 10.6353 | 159093.1011 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 48.3483 | 96.6965 | 87286.1334 | 13.6465 | 681565.6496 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2077 | 38.7716 | 77.5432 | 9662.6286 | 13.6465 | 681565.6496
| CC9 |  | 121019068.D | CC | 9 | 2019-12-11 16:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 41.9885 | 47.9869 | 40275.4253 | 13.6465 | 910541.7496 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 51.0969 | 58.3965 | 304988.6139 | 10.6353 | 199553.6735 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 46.0462 | 52.6243 | 110984.9473 | 13.6465 | 910541.7496 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2077 | 42.6225 | 48.7114 | 14429.1588 | 13.6465 | 910541.7496
| CC10 |  | 121019069.D | CC | 10 | 2019-12-11 16:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 93.0881 | 74.4704 | 84856.5413 | 13.6465 | 856638.1403 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 138.702 | 110.9616 | 819833.4928 | 10.6303 | 199641.6519 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 134.6618 | 107.7294 | 308146.8992 | 13.6465 | 856638.1403 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1993 | 95.1573 | 76.1258 | 33087.6487 | 13.6465 | 856638.1403
| CC11 |  | 121019070.D | CC | 11 | 2019-12-11 17:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 179.0725 | 89.5362 | 223268.9407 | 13.638 | 1167045.1184 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 199.4197 | 99.7099 | 1536077.0339 | 10.6303 | 260642.9205 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 177.1575 | 88.5787 | 552907.5218 | 13.638 | 1167045.1184 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 129.5798 | 64.7899 | 62495.3266 | 13.638 | 1167045.1184
| CC12 |  | 121019071.D | CC | 12 | 2019-12-11 17:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5545 | 294.1101 | 78.4294 | 31879.1902 | 13.6465 | 101288.4992 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6261 | 390.0651 | 104.0174 | 240276.7157 | 10.6202 | 20886.3483 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 313.2502 | 83.5334 | 84982.5265 | 13.6465 | 101288.4992 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1823 | 130.4361 | 34.783 | 5461.6095 | 13.6465 | 101288.4992
| CC13 |  | 121019072.D | CC | 13 | 2019-12-11 18:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 536.5405 | 85.8465 | 303854.3201 | 13.6296 | 528586.7676 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6362 | 664.9071 | 106.3851 | 3709734.8432 | 10.6303 | 189344.9976 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 698.3102 | 111.7296 | 989748.9132 | 13.6296 | 528586.7676 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1909 | 399.3619 | 63.8979 | 90136.2645 | 13.6296 | 528586.7676
| CC14 |  | 121019073.D | CC | 14 | 2019-12-11 18:52 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 675.7668 | 77.2305 | 452350.7583 | 13.6211 | 624603.803 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6312 | 881.7777 | 100.7746 | 4740765.3003 | 10.6303 | 182513.6768 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 851.3115 | 97.2927 | 1426014.6517 | 13.6211 | 624603.803 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1908 | 368.475 | 42.1114 | 98144.6679 | 13.6211 | 624603.803
| CC15 |  | 121019074.D | CC | 15 | 2019-12-11 19:23 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5434 | 942.1452 | 75.3716 | 1077231.3403 | 13.621 | 1066543.7172 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6311 | 1259.457 | 100.7566 | 9467701.7714 | 10.6303 | 255264.6389 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 1092.0583 | 87.3647 | 3124109.2286 | 13.621 | 1066543.7172 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.1907 | 825.9585 | 66.0767 | 379143.8841 | 13.621 | 1066543.7172
| QCGC4 |  | 121019075.D | QC | 4 | 2019-12-12 10:02 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5491 | 5.1872 | 69.1625 | 4219.5821 | 13.6211 | 888748.7161 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6363 | 6.4501 | 86.0009 | 50720.3666 | 10.6354 | 236545.5195 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7768 | 6.9345 | 92.4606 | 15038.0108 | 13.6211 | 888748.7161 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2079 | 31.9348 | 425.7967 | 9965.3335 | 13.6211 | 888748.7161

## 1206Analysis
All conc's listed in nM unless otherwise specified
UC Summary
Run Reason: | rerunning inconsistent PFAS
Date of UC run: | 2019-12-06
Date SE'd: | 2019-12-10
12/10 GC1 UC 12/6 Results
| Compound | Alone/Mix | Fu | SD | n | Stability | Comments
| 4NT | M | 0.065005 | 0.005034 | 3 | -16.3%
| 916 | M | 0.45042 | 0.057692 | 3 | -9.304%
| 923 | M | 0.87014 | 0.12481 | 3 | -33.7%
| 3117 | M | 0.02975 | 0.0060229 | 3 | -19.27%
12/6 assay data; 12/10 SE; 12/10 GCMS | [13 empty] | flagged if >/< 125/75% of avg |  |  |  | IS has been flagged if >/< 150/50% of avg
[14 empty] | 2X or 10X | 4X crash
Compound | 4NT
Sample | [7 empty] | 4NT Method |  | 4NT Results | [10 empty] | M8FOSA (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6462 | 68.4454 |  | 99.8847 | 136.8908 | 547.5633 | 0.064962 |  |  |  | 365589.9644 | 10.6454 | 176795.2524
|  | UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6462 | 62.2255 |  | 90.8077 | 124.451 | 497.8038 | 0.059992 |  |  |  | 345991.0443 | 10.6403 | 183447.3133
|  | UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6463 | 74.9024 | 68.5244 | 109.3076 | 149.8048 | 599.2193 | 0.07006 | 0.065005 | 0.005034 |  | 349100.184 | 10.6454 | 154700.2595
|  | UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6463 | 247.5518 |  | 98.3603 | 2475.5183 | 9902.0732 |  |  |  |  | 1301423.3336 | 10.6454 | 178191.595
|  | UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6463 | 258.5619 |  | 102.735 | 2585.6193 | 10342.4773 |  |  |  |  | 1522439.9451 | 10.6454 | 199655.0197
|  | UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6412 | 248.9219 | 251.6786 | 98.9047 | 2489.219 | 9956.8759 |  | 10067.1421 | 240.0165 |  | 1301043.3493 | 10.6353 | 177168.0431
|  | UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6412 | 210.7233 |  | 100.0284 | 2107.2333 | 8428.9333 |  |  |  |  | 638209.6613 | 10.6353 | 102491.7623
|  | UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6412 | 207.4443 |  | 98.4719 | 2074.443 | 8297.772 |  |  |  |  | 487408.7743 | 10.6354 | 79497.9818
|  | UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.087696 * x + 0.067078 | 0.99987 | 10.6361 | 213.8229 | 210.6635 | 101.4997 | 2138.2293 | 8552.9174 |  | 8426.5409 | 127.5895 | -16.3% | 603645.8742 | 10.6353 | 95550.725
Compound | 916
Sample | [7 empty] | 916 Method |  | 916 Results | [10 empty] | M8FOSA (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 376.492 |  | 99.3955 | 752.9841 | 3011.9364 | 0.44611 |  |  |  | 946110.3486 | 13.6295 | 937909.4467
|  | UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7851 | 361.1413 |  | 95.3428 | 722.2826 | 2889.1305 | 0.395 | 0.42056 | 0.03614 |  | 909879.5196 | 13.6295 | 940399.9193
|  | UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 398.7124 | 378.7819 | 105.2617 | 797.4247 | 3189.6988 | 0.51014 | 0.45042 | 0.057692 |  | 880671.4206 | 13.6295 | 824306.2183
|  | UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 176.4107 |  | 94.4952 | 1764.1073 | 7056.429 |  |  |  |  | 438160.3374 | 13.6295 | 928772.3968
|  | UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7852 | 175.7662 |  | 94.15 | 1757.6622 | 7030.6489 |  |  |  |  | 495187.5451 | 13.6295 | 1053516.204
|  | UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 207.8853 | 186.6874 | 111.3548 | 2078.8533 | 8315.4133 |  | 7467.4971 | 734.4301 |  | 488379.8161 | 13.655 | 878010.1054
|  | UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 168.7879 |  | 99.6864 | 1687.8795 | 6751.5179 |  |  |  |  | 314743.9191 | 13.655 | 697408.8117
|  | UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 182.8553 |  | 107.9946 | 1828.5526 | 7314.2106 |  |  |  |  | 219861.7503 | 13.655 | 449560.9592
|  | UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.067095 * x - 0.001691 | 0.99885 | 11.7767 | 156.3135 | 169.3189 | 92.319 | 1563.1346 | 6252.5384 |  | 6772.7556 | 531.1546 | -9.304% | 296577.9907 | 13.6549 | 709812.9343
[20 empty] | 554396.9608 |  | 824410.7773
Compound | 923
Sample | [7 empty] | 923 Method |  | 923 Results | [10 empty] | M8FOSA (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2587 | 576.8576 |  | 111.9346 | 1153.7152 | 4614.8609 | 0.97482 |  |  |  | 232116.018 | 13.6295 | 937909.4467
|  | UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2501 | 499.2984 |  | 96.8848 | 998.5968 | 3994.3871 | 0.90358 | 0.9392 | 0.050374 |  | 201108.2176 | 13.6295 | 940399.9193
|  | UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 469.9016 | 515.3525 | 91.1806 | 939.8033 | 3759.213 | 0.73202 | 0.87014 | 0.12481 |  | 165774.5588 | 13.6295 | 824306.2183
|  | UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2586 | 188.6131 |  | 105.0063 | 1886.1314 | 7544.5257 |  |  |  |  | 73508.5583 | 13.6295 | 928772.3968
|  | UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2672 | 198.3034 |  | 110.4011 | 1983.0339 | 7932.1355 |  |  |  |  | 87807.9036 | 13.6295 | 1053516.204
|  | UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 151.9461 | 179.6209 | 84.5927 | 1519.4608 | 6077.8433 |  | 7184.8348 | 978.0762 |  | 55532.1209 | 13.655 | 878010.1054
|  | UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2248 | 118.3517 |  | 99.3848 | 1183.5173 | 4734.0692 |  |  |  |  | 33951.0776 | 13.655 | 697408.8117
|  | UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2333 | 110.5156 |  | 92.8045 | 1105.1561 | 4420.6245 |  |  |  |  | 20357.9782 | 13.655 | 449560.9592
|  | UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.010840 * x - 0.002634 | 0.99675 | 23.2331 | 128.3856 | 119.0843 | 107.8107 | 1283.8557 | 5135.4227 |  | 4763.3721 | 358.2989 | -33.7% | 37642.9772 | 13.6549 | 709812.9343
[20 empty] | 100866.6011 |  | 824410.7773
Compound | 3117
Sample | [7 empty] | 3117 Method |  | 3117 Results | [10 empty] | M8FOSA (ISTD) Results
|  | Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
|  | UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 15.9634 |  | 86.6694 | 31.9268 | 127.7074 | 0.024616 |  |  |  | 15297.0462 | 13.6295 | 937909.4467
|  | UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 19.8436 |  | 107.7358 | 39.6872 | 158.7489 | 0.028252 |  |  |  | 19252.6372 | 13.6295 | 940399.9193
|  | UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 19.4492 | 18.4188 | 105.5948 | 38.8985 | 155.594 | 0.03638 | 0.02975 | 0.0060229 |  | 16527.0979 | 13.6295 | 824306.2183
|  | UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 171.2408 |  | 109.983 | 1712.4076 | 6849.6302 | 0.018644 |  |  |  | 169880.4215 | 13.6295 | 928772.3968
|  | UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5599 | 136.7077 |  | 87.8034 | 1367.0769 | 5468.3077 | 0.029031 |  |  |  | 153663.3872 | 13.6295 | 1053516.204
|  | UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5544 | 159.144 | 155.6975 | 102.2136 | 1591.4398 | 6365.7592 | 0.024442 | 6227.899 | 700.9044 |  | 149200.0719 | 13.655 | 878010.1054
|  | UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 129.6975 |  | 103.1816 | 1296.9747 | 5187.8988 |  |  |  |  | 96476.9489 | 13.655 | 697408.8117
|  | UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.549 | 140.4741 |  | 111.755 | 1404.7407 | 5618.9627 |  |  |  |  | 67388.5699 | 13.655 | 449560.9592
|  | UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.026823 * x - 8.176159E-004 | 0.99596 | 9.5489 | 106.9233 | 125.6983 | 85.0634 | 1069.2328 | 4276.931 |  | 5027.9308 | 685.1675 | -19.27% | 80848.8259 | 13.6549 | 709812.9343

## 4NT Analysis
11/1 GC
Sample | [6 empty] | 4NT Method |  | 4NT Results | [10 empty] | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
| UCG1 T1 S1A1009_SE 11/1 |  | 110119018.D | Sample |  | 2019-11-01 18:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 195.4269 |  |  | 1954.2695 | 7817.0779 |  |  |  |  | 1211178.5546 | 10.6858 | 198269.1936
| UCG1 T1 S1B1009_SE 11/1 |  | 110119019.D | Sample |  | 2019-11-01 18:52 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 181.811 |  |  | 1818.1095 | 7272.4381 |  |  |  |  | 1121879.2065 | 10.6858 | 196739.2672
| UCG1 T1 S1C1009_SE 11/1 |  | 110119020.D | Sample |  | 2019-11-01 19:22 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 190.1028 | 189.1136 |  | 1901.0281 | 7604.1126 |  | 7564.5429 | 274.4676 |  | 887592.675 | 10.6858 | 149178.9935
| UCG1 AF S1A1009_SE 11/1 |  | 110119024.D | Sample |  | 2019-11-01 21:23 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 50.6386 |  |  | 101.2771 | 405.1086 | 0.062346 |  |  |  | 495449.9333 | 10.6807 | 277205.6544
| UCG1 AF S1B1009_SE 11/1 |  | 110119025.D | Sample |  | 2019-11-01 21:53 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6867 | 56.149 |  |  | 112.298 | 449.192 | 0.072878 |  |  |  | 397409.0147 | 10.6808 | 203614.7766
| UCG1 AF S1C1009_SE 11/1 |  | 110119026.D | Sample |  | 2019-11-01 22:23 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 52.9511 | 53.2462 |  | 105.9022 | 423.6088 | 0.061437 | 0.065554 | 0.0063592 |  | 277680.1347 | 10.6807 | 149586.1709
| UCG1 T5 S1A1009_SE 11/1 |  | 110119027.D | Sample |  | 2019-11-01 22:53 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 162.4426 |  |  | 1624.4261 | 6497.7044 |  |  |  |  | 770511.5218 | 10.6807 | 150364.717
| UCG1 T5 S1B1009_SE 11/1 |  | 110119028.D | Sample |  | 2019-11-01 23:23 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6866 | 154.0909 |  |  | 1540.9086 | 6163.6343 |  |  |  |  | 718306.2518 | 10.6807 | 147344.5174
| UCG1 T5 S1C1009_SE 11/1 |  | 110119031.D | Sample |  | 2019-11-02 00:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6816 | 172.3761 | 162.9699 |  | 1723.7609 | 6895.0435 |  | 6518.7941 | 366.1604 | -13.82% | 889123.1218 | 10.6807 | 164021.6837
| UCG1 AF S2A1009_SE 11/1 |  | 110119032.D | Sample |  | 2019-11-02 01:24 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6816 | 44.5587 |  |  | 89.1175 | 356.4699 | 0.090112 |  |  |  | 242562.1505 | 10.6807 | 151050.1979
| UCG1 AF S2B1009_SE 11/1 |  | 110119033.D | Sample |  | 2019-11-02 01:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 30.8072 |  |  | 61.6144 | 246.4577 | 0.055615 |  |  |  | 180442.1819 | 10.6806 | 150947.113
| UCG1 AF S2C1009_SE 11/1 |  | 110119034.D | Sample |  | 2019-11-02 02:24 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 79.6546 | 51.6735 |  | 159.3093 | 637.237 | 0.13283 | 0.092851 | 0.038678 |  | 519311.6306 | 10.6806 | 195720.3545
| UCG1 T5 S2A1009_SE 11/1 |  | 110119035.D | Sample |  | 2019-11-02 02:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 101.7128 |  |  | 1017.1282 | 4068.5126 |  |  |  |  | 497130.8226 | 10.6807 | 150113.515
| UCG1 T5 S2B1009_SE 11/1 |  | 110119037.D | Sample |  | 2019-11-02 03:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 98.6321 |  |  | 986.3211 | 3945.2845 |  |  |  |  | 486340.4102 | 10.6806 | 151049.0995
| UCG1 T5 S2C1009_SE 11/1 |  | 110119039.D | Sample |  | 2019-11-02 04:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 98.896 | 98.764 |  | 988.9596 | 3955.8385 |  | 3950.5615 | 7.4628 |  | 706413.6929 | 10.6756 | 218864.8269
| UCG1 T1 S2A1009_SE 11/1 |  | 110119040.D | Sample |  | 2019-11-02 05:24 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 110.7876 |  |  | 1107.8764 | 4431.5057 |  |  |  |  | 671439.8281 | 10.6756 | 187419.3664
| UCG1 T1 S2B1009_SE 11/1 |  | 110119041.D | Sample |  | 2019-11-02 05:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 119.9381 |  |  | 1199.3807 | 4797.523 |  |  |  |  | 571204.1421 | 10.6756 | 148146.7868
| UCG1 T1 S2C1009_SE 11/1 |  | 110119043.D | Sample |  | 2019-11-02 06:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 120.811 | 117.1789 |  | 1208.1095 | 4832.4381 |  | 4687.1556 | 222.0865 | 18.65% | 559393.2953 | 10.6756 | 144109.7976
| UCG1 T1 S3A1009_SE 11/1 |  | 110119044.D | Sample |  | 2019-11-02 07:24 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 201.0527 |  |  | 2010.5272 | 8042.1088 |  |  |  |  | 1177350.1613 | 10.6756 | 187575.633
| UCG1 T1 S3B1009_SE 11/1 |  | 110119045.D | Sample |  | 2019-11-02 07:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 183.8573 |  |  | 1838.5734 | 7354.2937 |  |  |  |  | 766588.8586 | 10.6756 | 133008.8278
| UCG1 T1 S3C1009_SE 11/1 |  | 110119046.D | Sample |  | 2019-11-02 08:24 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 184.173 | 189.6944 |  | 1841.7302 | 7366.9209 |  | 7587.7745 | 393.5158 |  | 838277.2871 | 10.6756 | 145209.9126
| UCG1 AF S3A1009_SE 11/1 |  | 110119047.D | Sample |  | 2019-11-02 08:54 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6815 | 155.6529 |  |  | 311.3057 | 1245.2229 | 0.52865 |  |  |  | 933862.6116 | 10.6756 | 189746.5236
| UCG1 AF S3B1009_SE 11/1 |  | 110119050.D | Sample |  | 2019-11-02 10:25 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 163.7522 |  |  | 327.5045 | 1310.018 | 0.57579 |  |  |  | 1029595.959 | 10.6756 | 199403.7787
| UCG1 AF S3C1009_SE 11/1 |  | 110119051.D | Sample |  | 2019-11-02 10:55 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 181.1785 | 166.8612 |  | 362.357 | 1449.4279 | 0.62828 | 0.57757 | 0.049842 |  | 887007.0383 | 10.6756 | 156067.3576
| UCG1 T5 S3A1009_SE 11/1 |  | 110119052.D | Sample |  | 2019-11-02 11:25 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 58.8873 |  |  | 588.8727 | 2355.4909 |  |  |  |  | 427192.0114 | 10.6756 | 210077.4583
| UCG1 T5 S3B1009_SE 11/1 |  | 110119053.D | Sample |  | 2019-11-02 11:55 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 56.8794 |  |  | 568.7939 | 2275.1757 |  |  |  |  | 377330.525 | 10.6756 | 191191.966
| UCG1 T5 S3C1009_SE 11/1 |  | 110119056.D | Sample |  | 2019-11-02 13:25 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6765 | 57.6741 | 57.8136 |  | 576.741 | 2306.964 |  | 2312.5435 | 40.4473 | -69.52% | 385361.9572 | 10.6706 | 192942.5712
12/10 GC
Sample | [6 empty] | 4NT Method |  | 4NT Results | [10 empty] | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula Linear | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
| UC1206_S3aT1 |  | 121019026.D | Sample |  | 2019-12-11 00:51 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 247.3339 |  |  | 2473.3394 | 9893.3574 |  |  |  |  | 1301423.3336 | 10.6454 | 178191.595
| UC1206_S3bT1 |  | 121019019.D | Sample |  | 2019-12-10 21:21 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 258.2701 |  |  | 2582.7015 | 10330.806 |  |  |  |  | 1522439.9451 | 10.6454 | 199655.0197
| UC1206_S3cT1 |  | 121019050.D | Sample |  | 2019-12-11 10:22 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 248.6948 | 251.433 |  | 2486.9481 | 9947.7923 |  | 10057.3186 | 238.4058 |  | 1301043.3493 | 10.6353 | 177168.0431
| UC1206_S3aAF |  | 121019023.D | Sample |  | 2019-12-10 23:21 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6462 | 69.4295 |  |  | 138.859 | 555.4359 | 0.065887 |  |  |  | 365589.9644 | 10.6454 | 176795.2524
| UC1206_S3bAF |  | 121019031.D | Sample |  | 2019-12-11 03:22 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6462 | 63.2513 |  |  | 126.5026 | 506.0104 | 0.060966 |  |  |  | 345991.0443 | 10.6403 | 183447.3133
| UC1206_S3cAF |  | 121019020.D | Sample |  | 2019-12-10 21:51 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6463 | 75.8432 | 69.508 |  | 151.6863 | 606.7453 | 0.070937 | 0.06593 | 0.0049856 |  | 349100.184 | 10.6454 | 154700.2595
| UC1206_S3aT5 |  | 121019043.D | Sample |  | 2019-12-11 08:22 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 210.7526 |  |  | 2107.5259 | 8430.1037 |  |  |  |  | 638209.6613 | 10.6353 | 102491.7623
| UC1206_S3bT5 |  | 121019048.D | Sample |  | 2019-12-11 09:22 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6412 | 207.4956 |  |  | 2074.9557 | 8299.8226 |  |  |  |  | 487408.7743 | 10.6354 | 79497.9818
| UC1206_S3cT5 |  | 121019052.D | Sample |  | 2019-12-11 11:22 | y = 0.088288 * x + 0.024599 | 0.99911 | 10.6361 | 213.8314 | 210.6932 |  | 2138.3139 | 8553.2557 |  | 8427.7273 | 126.7332 | -16.2% | 603645.8742 | 10.6353 | 95550.725
10/29 GC
[5 empty] | 4-Nitrotuluene Method |  | 4-Nitrotuluene Results |  | 4NT Results | [10 empty] | 4NT13C6 (ISTD) Results
| Name | Comment | Data File | Type | Level | Acq. Date-Time | CF Formula | CF R2 | RT | Final Conc. | Rep Avg | Accuracy | Assay Diln Adj | Sample Prep Diln Adj | Fu | Avg | SD | Stability | Area | RT | Area
| UCG1T1hS1A_SE 10/11 |  | 102919016.D | Sample |  | 2019-10-29 19:05 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 204.1802 |  |  | 2041.8016 | 8167.2063 |  |  |  |  | 1121180.8762 | 10.6958 | 176009.205
| UCG1T1hS1B_SE 10/11 |  | 102919017.D | Sample |  | 2019-10-29 19:35 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 191.6225 |  |  | 1916.2251 | 7664.9002 |  |  |  |  | 1079948.8643 | 10.6958 | 180135.344
| UCG1T1hS1C_SE 10/11 |  | 102919041.D | Sample |  | 2019-10-30 07:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 204.9475 | 200.25 |  | 2049.4747 | 8197.8987 |  | 8010.0018 | 299.2604 |  | 751962.6504 | 10.6908 | 117624.389
| UCG1AFS1A_SE 10/11 |  | 102919029.D | Sample |  | 2019-10-30 01:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 55.5445 |  |  | 111.089 | 444.3562 | 0.069959 |  |  |  | 288965.2862 | 10.6958 | 149434.3378
| UCG1AFS1B_SE 10/11 |  | 102919030.D | Sample |  | 2019-10-30 02:07 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 42.7224 |  |  | 85.4448 | 341.7792 | 0.052941 |  |  |  | 235048.9524 | 10.6958 | 151543.8603
| UCG1AFS1C_SE 10/11 |  | 102919036.D | Sample |  | 2019-10-30 05:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 49.7417 | 49.3362 |  | 99.4833 | 397.9333 | 0.0618 | 0.061567 | 0.0085116 |  | 247997.1067 | 10.6908 | 140864.9458
| UCG1T5hS1A_SE 10/11 |  | 102919038.D | Sample |  | 2019-10-30 06:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 158.7909 |  |  | 1587.9088 | 6351.6352 |  |  |  |  | 639507.296 | 10.6908 | 127511.4805
| UCG1T5hS1B_SE 10/11 |  | 102919023.D | Sample |  | 2019-10-29 22:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 161.3961 |  |  | 1613.9614 | 6455.8455 |  |  |  |  | 598028.657 | 10.6958 | 117420.5272
| UCG1T5hS1C_SE 10/11 |  | 102919024.D | Sample |  | 2019-10-29 23:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 160.9751 | 160.3874 |  | 1609.7512 | 6439.005 |  | 6415.4952 | 55.9418 | -19.91% | 791660.6949 | 10.6958 | 155823.8592
| UCG1T1hS2A_SE 10/11 |  | 102919042.D | Sample |  | 2019-10-30 08:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 120.2512 |  |  | 1202.5122 | 4810.0489 |  |  |  |  | 631905.8769 | 10.6908 | 163493.9884
| UCG1T1hS2B_SE 10/11 |  | 102919018.D | Sample |  | 2019-10-29 20:05 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 124.6609 |  |  | 1246.6088 | 4986.4352 |  |  |  |  | 569660.7411 | 10.6958 | 142535.4979
| UCG1T1hS2C_SE 10/11 |  | 102919019.D | Sample |  | 2019-10-29 20:35 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 126.5147 | 123.8089 |  | 1265.1469 | 5060.5875 |  | 4952.3572 | 128.6988 |  | 699851.9262 | 10.6958 | 172719.6256
| UCG1AFS2A_SE 10/11 |  | 102919046.D | Sample |  | 2019-10-30 10:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 29.2628 |  |  | 58.5257 | 234.1026 | 0.054189 |  |  |  | 137650.5214 | 10.6908 | 119768.454
| UCG1AFS2B_SE 10/11 |  | 102919043.D | Sample |  | 2019-10-30 08:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 32.342 |  |  | 64.6839 | 258.7356 | 0.061038 |  |  |  | 203318.7442 | 10.6908 | 163807.2727
| UCG1AFS2C_SE 10/11 |  | 102919044.D | Sample |  | 2019-10-30 09:08 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 70.1815 | 51.2617 |  | 140.3629 | 561.4516 | 0.13329 | 0.082838 | 0.043824 |  | 300186.8179 | 10.6908 | 126629.4829
| UCG1T5hS2A_SE 10/11 |  | 102919027.D | Sample |  | 2019-10-30 00:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 108.0025 |  |  | 1080.0255 | 4320.1018 |  |  |  |  | 589490.5923 | 10.6958 | 168453.4576
| UCG1T5hS2B_SE 10/11 |  | 102919045.D | Sample |  | 2019-10-30 09:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6916 | 105.973 |  |  | 1059.73 | 4238.9201 |  |  |  |  | 535954.894 | 10.6908 | 155852.859
| UCG1T5hS2C_SE 10/11 |  | 102919033.D | Sample |  | 2019-10-30 03:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 105.3093 | 106.4283 |  | 1053.0931 | 4212.3725 |  | 4257.1315 | 56.1262 | -14.04% | 615327.7822 | 10.6958 | 179970.7696
| UCG1T1hS3A_SE 10/11 |  | 102919020.D | Sample |  | 2019-10-29 21:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 186.651 |  |  | 1866.5104 | 7466.0414 |  |  |  |  | 1013827.9057 | 10.6958 | 173397.994
| UCG1T1hS3B_SE 10/11 |  | 102919021.D | Sample |  | 2019-10-29 21:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.7017 | 190.2535 |  |  | 1902.5352 | 7610.1409 |  |  |  |  | 976246.0272 | 10.6958 | 163955.1388
| UCG1T1hS3C_SE 10/11 |  | 102919037.D | Sample |  | 2019-10-30 05:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 194.1016 | 190.3354 |  | 1941.0156 | 7764.0624 |  | 7613.4149 | 149.0374 |  | 813175.2011 | 10.6958 | 133983.9266
| UCG1AFS3A_SE 10/11 |  | 102919031.D | Sample |  | 2019-10-30 02:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 49.7484 |  |  | 99.4969 | 397.9876 | 0.063438 |  |  |  | 206419.3933 | 10.6958 | 117234.8739
| UCG1AFS3B_SE 10/11 |  | 102919032.D | Sample |  | 2019-10-30 03:07 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 51.3178 |  |  | 102.6356 | 410.5423 | 0.064526 |  |  |  | 239496.6075 | 10.6958 | 132496.1929
| UCG1AFS3C_SE 10/11 |  | 102919039.D | Sample |  | 2019-10-30 06:38 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 45.4613 | 48.8425 |  | 90.9226 | 363.6903 | 0.058627 | 0.062197 | 0.003139 |  | 236287.5488 | 10.6908 | 144715.2044
| UCG1T5hS3A_SE 10/11 |  | 102919035.D | Sample |  | 2019-10-30 04:37 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 156.8419 |  |  | 1568.4192 | 6273.6769 |  |  |  |  | 955242.6169 | 10.6908 | 192701.0513
| UCG1T5hS3B_SE 10/11 |  | 102919025.D | Sample |  | 2019-10-29 23:36 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 159.0605 |  |  | 1590.6055 | 6362.4219 |  |  |  |  | 918682.4946 | 10.6958 | 182882.7801
| UCG1T5hS3C_SE 10/11 |  | 102919026.D | Sample |  | 2019-10-30 00:06 | y = 0.089540 * x + 0.275906 | 0.99955 | 10.6967 | 155.0856 | 156.996 |  | 1550.8561 | 6203.4246 |  | 6279.8411 | 79.6777 | -17.52% | 678759.6727 | 10.6958 | 138389.5867

## ValueList_Helper (hidden)
Sample | 1
Blank | 2
Cal | 3
QC | 4
CC | 5
DoubleBlank | 6
MatrixSpike | 7
MatrixSpikeDup | 8
MatrixBlank | 9
TuneCheck | 10
ResponseCheck | 11
12
13
14
15