## Cover Sheet
Analytical Analysis Cover Sheet for Plasma Protein Binding Assessment of PFAS Compounds
QAPP: | In Vitro Toxicokinetic (TK) Assessments of PFAS (D-CCED-0031343-1-1)
PI: | Barbara A. Wetmore
Analysis Set ID: | 20201110_PFAS_PPB_UC_Repeats1
Analytes: | DTXSID1037303, DTXSID3031860, DTXSID3059921, DTXSID4059833, DTXSID50375114, DTXSID90868151, DTXSID3037707, DTXSID80108992
Calibration Curve Prepared Date: | 2020-11-10
Preparer: | Marci G. Smeltz
Sample Type: | Calibration Curve (CC1 - CC17), Curve Precision Check (QC-A - QC-D)
Ultracentrifugation Assay Conducted Date: | 2020-11-10
Preparer: | Marci G. Smeltz
Sample Type: | Ultracentrifugation Assay Samples (Human Plasma)
Analytical Analysis Date: | 2020-11-23
Analytical Technique: | UPLC-MS/MS
Instrument: | Waters Xevo TQ-S micro (QEB0036)
Analyst: | Marci G. Smeltz
Total Samples Number (including blanks): | 127
Actual Sample Number: | 99
Method Detection Limit (MDL) Definition: | Minimum measured concentration of a substance that can be reported with 99% confidence that the measured concentration is distinguishable from method blank results
MDLXi = method detection limit determined for target analyte X from solution i = t(n-1,1-α=0.99) x Si
t(n-1,1-α=0.99) = 3.143, the student's t-value for a 99% confidence level and a standard deviation estimate with six degrees of freedom (n=7) |  |  |  | 40 CFR Part 136 | EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
Si = standard deviation of the seven replicate analyses for solution i |  |  |  | https://www.ecfr.gov/cgi-bin/text-idx?SID=a6bb8a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
estimated Limit of Quantitation (eLOQ) Definition: | Lowest Calibration Curve point at which precision checks can be quantitated against curve ± 30%, provided other curve points have accuracy to ± 25%
QC Samples and Acceptance Criteria
QC Type | Frequency | Acceptance Criteria | Type | Corrective Action if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | 7-point curve min; r2≥0.98 | Linearity | Re-run samples and curve
Curve Precision Check (Analytical) | 2-3 samples per run | 75-125% of expected | Accuracy | Flag in comments
Blanks (Analytical) | 1 every 6-10 injections | <1/2 MDL of analytes | Specificity | Flag in comments, assess and resolve
Technical Replicate Assessment (Assay) | Every assay sample | Individual Replicates within ± 25% of Sample Mean | Precision | Flag in comments
Positive Control Reference Chemical(s) for Protein Binding (Assay) | 1 per assay batch | ± 20% of historical or published values | Binding Assay Performance | Flag in comments
Note: one of the three blanks (solvent, crash blank, crash mixed matrix) should be run every 6 injections
Analytes, Internal Standards Assessed
Analyte (Common Name) | DTXSID | Sample ID | Average Molecular Weight (g/mol) | Internal Standard | Peak Assignment from CoA | Mix | Concentration at Instrument (pg/uL ¦ µM)
Perfluoroheptanoic acid (PFHpA) | DTXSID1037303 | 1503225 | 364.062 | M4PFHpA | F | 1 | 3 ¦ 0.01
Perfluorodecanoic acid (PFDA) | DTXSID3031860 | 1504073 | 514.086 | M6PFDA | M | 1 | 3 ¦ 0.01
Perfluorotetradecanoic acid (PFTeDA) | DTXSID3059921 | 1503105 | 714.117 | M2PFTeDA | S | 1 | 3 ¦ 0.01
Octafluoroadipic acid | DTXSID4059833 | 1503115 | 290.065 | MPFBA | A | 2 | 3 ¦ 0.01
Perfluoro-3,6,9-trioxatridecanoic acid (PFPE-6) | DTXSID50375114 | 1480968 | 562.082 | M7PFUdA | Q | 2 | 3 ¦ 0.01
Perfluorotridecanoic acid (PFTrDA) | DTXSID90868151 | 1503185 | 664.109 | M2PFTeDA | S | 2 | 3 ¦ 0.01
Potassium perfluorobutanesulfonate (K+PFBS) | DTXSID3037707 | 1480942 | 338.19 | M3PFBS | C | 3 | 3 ¦ 0.01
Perfluorooctanesulfonate (PFOS-) | DTXSID80108992 | 1503156 | 499.12 | M8PFOS | K | 3 | 3 ¦ 0.01
n-Butylparaben (Reference Compound) | DTXSID3020209 | Sigma Cat#54680-50G-F, lot BCBT3342 | 194.23 | 13C6-n-Butylparaben | N/A | 1, 2, 3 | 3 ¦ 0.01
Internal Standard Information
Standard | Vendor | Catalog Number | Lot Number | Other Descriptors
MPFAC-24ES | Wellington Lab | MPFAC-24ES | MPFAC-24ES0218 | in Methanol/Isopropanol (2%)/Water (<1%), at 1 ug/mL, maintain at 4C
13C6-n-Butylparaben | Cambridge Isotope Lab | CLM-8285-S | SDEH-006 | in Methanol, prepared 9/25/2019 at 10 ug/mL, maintain at 4C
Human Plasma | BioIVT | HUMANPLK21901363 | HMN461056 | ultrafiltrate prepared 11/9/2020, stored at -80C until use in assay
An Example Certificate of Analysis for Wellington MPFAC-24ES Mix (Mass-labelled PFAS Solution)
Sample ID Key
CC# | Calibration Curve
QC-X | Quality Check
UC_AF_Mix#_X | Ultracentrifugation Assay_Crashed Aqueous Fraction_Mix_Replicate
UC_T#hr_Mix#_X | Ultracentrifugation Assay_Stability Time Point_Mix_Replicate
Action/Individual | Date
UC assay, calibration curve prepared (MGS) | 2020-11-10
Sample prep, UPLC-MS/MS analysis (MGS) | 2020-11-23
Data Processing by TargetLynx (MGS) | 2020-11-30
Initial Report Generation (MGS) | 2021-01-08
Final Report Generation (MGS) | 2021-06-28
Report Review (BAW) | 2021-06-28
*MGS = Marci Smeltz
*BAW = Barbara Wetmore
Additional notes
Concentrations listed are at the instrument, unless otherwise specified

## Executive Summary
|  |  |  | Concentration (nM) | [10 empty] | Fraction Unbound | [5 empty] | Stability (T5/T1) |  |  |  |  | Comments
Analyte | DTXSID | Mix | R2 | MDL | eLOQ | AF A | AF B | AF C | T1 A | T1 B | T1 C | T5 A | T5 B | T5 C | Fu A | Fu B | Fu C | Avg Fu | Std Dev Fu | CV (%) | A | B | C | Avg | Std Dev
Perfluoroheptanoic acid | DTXSID1037303 | 1 | 0.998 | 0.021739 | 0.17 | 3.3482 | 3.3437 | 4.5818 | 9919.8128 | 10834.7448 | 10516.6224 | 9975.384 | 10806.7272 | 10673.9024 | 0.00033564 | 0.00030941 | 0.00042925 | 0.0003581 | 6.2998e-05 | 17.59% | 1.0056 | 0.99741 | 1.015 | 1.006 | 0.0087771
Perfluorodecanoic acid | DTXSID3031860 | 1 | 0.998 | 0.036642 | 0.28 | 28.2525 | 31.6554 | 26.6525 | 9821.4552 | 10508.388 | 10818.792 | 10042.6256 | 10758.5544 | 10689.2536 | 0.0028133 | 0.0029423 | 0.0024934 | 0.0027497 | 0.00023113 | 8.406% | 1.0225 | 1.0238 | 0.98803 | 1.0115 | 0.020296
Perfluorotetradecanoic acid | DTXSID3059921 | 1 | 0.995 | 0.15064 | 1.14 | 259.495 | 290.7878 | 232.3824 | 8972.1104 | 9686.0552 | 9639.7944 | 8982.4552 | 10001.3128 | 10016.2 | 0.028889 | 0.029075 | 0.023201 | 0.027055 | 0.0033392 | 12.34% | 1.0012 | 1.0325 | 1.039 | 1.0242 | 0.020264
N-butylparaben (Ref Compound) | DTXSID3020209 | 1 | 0.977 | 0.12989 | 1.182 | 149.48 | 165.1712 | 170.0992 | 9231.224 | 9953.92 | 9681.2968 | 7501.4456 | 8218.8728 | 8181.8008 | 0.019927 | 0.020097 | 0.02079 | 0.020271 | 0.00045727 | 2.256% | 0.81262 | 0.82569 | 0.84511 | 0.82781 | 0.016352
Perfluoro-3,6,9-trioxatridecanoic acid | DTXSID50375114 | 2 | 0.997 | 0.046436 | 0.28 | 17.6762 | 17.7494 | 20.9683 | 10808.5312 | 10279.284 | 10923.3344 | 10510.596 | 10646.9056 | 10746.6696 | 0.0016817 | 0.0016671 | 0.0019511 | 0.0017667 | 0.00015993 | 9.053% | 0.97244 | 1.0358 | 0.98383 | 0.99734 | 0.033758
Perfluorotridecanoic acid | DTXSID90868151 | 2 | 0.994 | 0.16677 | 0.71 | 507.1606 | 456.6042 | 475.1165 | 10858.4648 | 9859.6992 | 10995.8264 | 10495.0744 | 12162.216 | 10817.3848 | 0.048324 | 0.037543 | 0.043922 | 0.043263 | 0.0054205 | 12.53% | 0.96653 | 1.2335 | 0.98377 | 1.0613 | 0.14942
Perfluorooctadipic acid | DTXSID4059833 | 2 | 0.019 | ND | ND | 4846.6544 | 4173.0934 | 3929.6243 | 16631.5128 | 17606.2472 | 18308.7288 | 21171.5496 | 17698.1256 | 16274.0488 | 0.22892 | 0.23579 | 0.24147 | 0.23539 | 0.0062809 | 2.668% | 1.273 | 1.0052 | 0.88887 | 1.0557 | 0.19697 | *need to repeat due to poor retention, lack of detection; try HILIC
N-butylparaben (Ref Compound) | DTXSID3020209 | 2 | 0.977 | 0.12989 | 1.182 | 174.2019 | 181.1498 | 184.5968 | 10622.7424 | 10592.184 | 10823.6528 | 8462.492 | 8585.1672 | 8620.1768 | 0.020585 | 0.0211 | 0.021415 | 0.021033 | 0.0004187 | 1.991% | 0.79664 | 0.81052 | 0.79642 | 0.80119 | 0.0080775
Potassium perfluorobutanesulfonate | DTXSID3037707 | 3 | 0.997 | 0.058931 | 0.71 | 47.5235 | 52.2803 | 58.5456 | 11267.6064 | 11341.8976 | 11489.72 | 11114.0552 | 11550.624 | 11470.4824 | 0.004276 | 0.0045262 | 0.005104 | 0.0046354 | 0.00042468 | 9.162% | 0.98637 | 1.0184 | 0.99833 | 1.001 | 0.016186
Perfluorosulfonate | DTXSID80108992 | 3 | 0.998 | 0.072928 | 0.28 | 50.0317 | 50.3146 | 54.201 | 10787.3776 | 11310.2 | 11180.5984 | 10896.1912 | 11138.7096 | 11015.9184 | 0.0045917 | 0.0045171 | 0.0049202 | 0.0046763 | 0.0002145 | 4.587% | 1.0101 | 0.98484 | 0.98527 | 0.9934 | 0.014454
N-butylparaben (Ref Compound) | DTXSID3020209 | 3 | 0.977 | 0.12989 | 1.182 | 159.7056 | 177.0912 | 175.6349 | 9863.7608 | 9743.2104 | 10105.3192 | 7866.2552 | 8035.9744 | 8001.064 | 0.020303 | 0.022037 | 0.021951 | 0.02143 | 0.00097768 | 4.562% | 0.79749 | 0.82478 | 0.79177 | 0.80468 | 0.017639

## Sample ID and Procedures
Analytical Sample Information |  |  |  |  | PROCEDURES
Analysis Set ID | 20201110_PFAS_PPB_UC_Repeats1 |  |  |  | *all samples stored in labeled 1.5 mL polypropylene microcentrifuge tubes at -80C unless noted
Total # samples | 127 |  |  |  | * final instrument dilution in mobile phase prepared day of analytical analysis
Actual # samples | 99 |  |  |  | PFAS-Plasma Calibration Curve
[6 empty] | C1 | V1 (uL) | # chems/mix | C2 (nM) | Dil | Conc (nM) at instrument | V2 (uL) | Vol diluent (uL) | Diluent
Sample ID | Sample Name | Type | PFAS Conc (nM) at Instrument | QC Flags | PFAS/Ref cmpd (3 mM) | 30 | 10 | 9 | 3 mM | 10 |  | 100 | 10 | DMSO
1 | Solvent Blank | Solvent |  | highlighted on each | Wkg Stock 30 uM | 3000000 | 12 | 1 | 30000 | 100 |  | 1200 | 1188 | MM
2 | Solvent Blank | Solvent |  | analyte sheets | Wkg Stock 10 uM | 30000 | 333.3333 |  | 10000 | 3 |  | 1000 | 666.6667 | MM
3 | Solvent Blank | Solvent |  |  | CC17 | 30000 | 46.6667 |  | 4000 | 7.5 | 250 | 350 | 303.3333 | MM
4 | Crash Blank | Blank |  |  | CC16 | 10000 | 87.5 |  | 2500 | 4 | 156.25 | 350 | 262.5 | MM
5 | Crash Mixed Matrix Blank | Blank |  |  | CC15 | 10000 | 70 |  | 2000 | 5 | 125 | 350 | 280 | MM
6 | Solvent Blank | Solvent |  |  | CC14 | 10000 | 43.75 |  | 1250 | 8 | 78.125 | 350 | 306.25 | MM
7 | CC1 | Standard | 0.17 |  | CC13 | 10000 | 27.3438 |  | 781.25 | 12.8 | 48.8281 | 350 | 322.6562 | MM
8 | CC2 | Standard | 0.28 |  | CC12 | 10000 | 17.0898 |  | 488.2812 | 20.48 | 30.5176 | 350 | 332.9102 | MM
9 | CC3 | Standard | 0.44 |  | CC11 | 10000 | 30.5176 |  | 305.1758 | 32.768 | 19.0735 | 1000 | 969.4824 | MM
10 | CC4 | Standard | 0.71 |  | CC10 | 10000 | 6.6757 |  | 190.7349 | 52.4288 | 11.9209 | 350 | 343.3243 | MM
11 | CC5 | Standard | 1.14 |  | CC9 | 10000 | 4.1723 |  | 119.2093 | 83.8861 | 7.4506 | 350 | 345.8277 | MM
12 | CC6 | Standard | 1.82 |  | CC8 | 305.18 | 85.448 |  | 74.5058 | 4.0961 | 4.6566 | 350 | 264.552 | MM
13 | Crash Blank | Blank |  |  | CC7 | 305.18 | 53.405 |  | 46.5661 | 6.5537 | 2.9104 | 350 | 296.595 | MM
14 | CC7 | Standard | 2.91 |  | CC6 | 305.18 | 33.3781 |  | 29.1038 | 10.4859 | 1.819 | 350 | 316.6219 | MM
15 | CC8 | Standard | 4.66 |  | CC5 | 305.18 | 20.8613 |  | 18.1899 | 16.7774 | 1.1369 | 350 | 329.1387 | MM
16 | CC9 | Standard | 7.45 |  | CC4 | 305.18 | 13.0383 |  | 11.3687 | 26.8439 | 0.71054 | 350 | 336.9617 | MM
17 | CC10 | Standard | 11.92 |  | CC3 | 305.18 | 8.149 |  | 7.1054 | 42.9503 | 0.44409 | 350 | 341.851 | MM
18 | CC11 | Standard | 19.07 |  | CC2 | 305.18 | 5.0931 |  | 4.4409 | 68.7204 | 0.27756 | 350 | 344.9069 | MM
19 | CC12 | Standard | 30.52 |  | CC1 | 305.18 | 3.1832 |  | 2.7756 | 109.9527 | 0.17347 | 350 | 346.8168 | MM
20 | Crash Mixed Matrix Blank | Blank |  |  | QC-D | 10000 | 14 |  | 400 | 25 | 25 | 350 | 336 | MM
21 | CC13 | Standard | 48.83 |  | QC-C | 305.18 | 114.6864 |  | 100 | 3.0518 | 6.25 | 350 | 235.3136 | MM
22 | CC14 | Standard | 78.13 |  | QC-B | 305.18 | 45.8746 |  | 40 | 7.6295 | 2.5 | 350 | 304.1254 | MM
23 | CC15 | Standard | 125 |  | QC-A | 305.18 | 11.4686 |  | 10 | 30.518 | 0.625 | 350 | 338.5314 | MM
24 | CC16 | Standard | 156.25
25 | CC17 | Standard | 250 |  | Crash Solution and Internal Standard Calculator
26 | Solvent Blank | Solvent |  |  | IS used | IS conc. | ng IS | Volume (uL)
27 | QC-A | QC | 0.63 |  | MPFAC-24ES | 1 | 320 | 320
28 | QC-B | QC | 2.5 |  | 13C6 n-butyl paraben | 10 | 320 | 32
29 | QC-C | QC | 6.25 |  | 1 M Formic Acid |  |  | 2000
30 | QC-D | QC | 25 |  | Acetonitrile |  |  | 17648
31 | Solvent Blank | Solvent
32 | CC1 | Standard | 0.17 |  | Final IS Conc. (ng/mL) | 16
33 | CC2 | Standard | 0.28 |  | Total Crash Volume (mL) | 20
34 | CC3 | Standard | 0.44
35 | CC4 | Standard | 0.71 |  | Ultracentrifugation Assay to Quantitate Plasma Protein Binding
36 | CC5 | Standard | 1.14
37 | CC6 | Standard | 1.82
38 | Crash Mixed Matrix Blank | Blank
39 | CC1 | Standard | 0.17
40 | CC2 | Standard | 0.28
41 | CC3 | Standard | 0.44
42 | CC4 | Standard | 0.71
43 | CC5 | Standard | 1.14
44 | CC6 | Standard | 1.82
45 | Solvent Blank | Solvent
46 | UC_AF_Mix1_A | Analyte
47 | UC_AF_Mix1_B | Analyte
48 | UC_AF_Mix1_C | Analyte
49 | UC_AF_Mix2_A | Analyte
50 | UC_AF_Mix2_B | Analyte
51 | UC_AF_Mix2_C | Analyte
52 | Crash Mixed Matrix Blank | Blank
53 | UC_AF_Mix3_A | Analyte
54 | UC_AF_Mix3_B | Analyte
55 | UC_AF_Mix3_C | Analyte
56 | UC_T1hr_Mix1_A | Analyte
57 | UC_T1hr_Mix1_B | Analyte
58 | UC_T1hr_Mix1_C | Analyte
59 | Solvent Blank | Solvent
60 | UC_T1hr_Mix2_A | Analyte
61 | UC_T1hr_Mix2_B | Analyte
62 | UC_T1hr_Mix2_C | Analyte
63 | UC_T1hr_Mix3_A | Analyte
64 | UC_T1hr_Mix3_B | Analyte
65 | UC_T1hr_Mix3_C | Analyte
66 | Crash Blank | Blank
67 | UC_T5hr_Mix1_A | Analyte
68 | UC_T5hr_Mix1_B | Analyte
69 | UC_T5hr_Mix1_C | Analyte
70 | UC_T5hr_Mix2_A | Analyte
71 | UC_T5hr_Mix2_B | Analyte
72 | UC_T5hr_Mix2_C | Analyte
73 | Crash Mixed Matrix Blank | Blank
74 | UC_T5hr_Mix3_A | Analyte
75 | UC_T5hr_Mix3_B | Analyte
76 | UC_T5hr_Mix3_C | Analyte
77 | Crash Mixed Matrix Blank | Blank
78 | Solvent Blank | Solvent
79 | CC1 | Standard | 0.17
80 | CC2 | Standard | 0.28
81 | CC3 | Standard | 0.44
82 | CC4 | Standard | 0.71
83 | CC5 | Standard | 1.14
84 | CC6 | Standard | 1.82
85 | Crash Blank | Blank
86 | CC7 | Standard | 2.91
87 | CC8 | Standard | 4.66
88 | CC9 | Standard | 7.45
89 | CC10 | Standard | 11.92
90 | CC11 | Standard | 19.07
91 | CC12 | Standard | 30.52
92 | Crash Mixed Matrix Blank | Blank
93 | CC13 | Standard | 48.83
94 | C14 | Standard | 78.13
95 | CC15 | Standard | 125
96 | CC16 | Standard | 156.25
97 | CC17 | Standard | 250
98 | Solvent Blank | Solvent
99 | QC-A | QC | 0.63
100 | QC-B | QC | 2.5
101 | QC-C | QC | 6.25
102 | QC-D | QC | 25
103 | Solvent Blank | Solvent
104 | Crash Mixed Matrix Blank | Blank
105 | CC1 | Standard | 0.17
106 | CC2 | Standard | 0.28
107 | CC3 | Standard | 0.44
108 | CC4 | Standard | 0.71
109 | CC5 | Standard | 1.14
110 | CC6 | Standard | 1.82
111 | Crash Mixed Matrix Blank | Blank
112 | CC1 | Standard | 0.17
113 | CC2 | Standard | 0.28
114 | CC3 | Standard | 0.44
115 | CC4 | Standard | 0.71
116 | CC5 | Standard | 1.14
117 | CC6 | Standard | 1.82
118 | Solvent Blank | Solvent
119 | CC1 | Standard | 0.17
120 | CC2 | Standard | 0.28
121 | CC3 | Standard | 0.44
122 | CC4 | Standard | 0.71
123 | CC5 | Standard | 1.14
124 | CC6 | Standard | 1.82
125 | Solvent Blank | Solvent
126 | Solvent Blank | Solvent
127 | Solvent Blank | Solvent

## Instrument Output (raw)
Quantify Compound Summary Report
Printed Fri Jan 08 13:46:09 2021
Compound 1: Perfluoroadipic acid
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 0.68 | 0.156 | 7 |  |  | 0.019 |  |  |  | MM | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 0.68 | 0.243 | 12 |  |  | 0.019 |  |  |  | MM | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 0.66 | 1393.331 | 20801 |  |  | 0.019 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 0.68 | 250.861 | 7254 | 50185.297 | 0 | 0.019 |  |  |  | bbI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 0.68 | 0.069 | 3 |  |  | 0.019 |  |  |  | MM | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 0.68 | 238.644 | 7371 | 49927.879 | 0 | 0.019 | 0.17 |  |  | bbI | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 0.68 | 242.337 | 7887 | 49278.141 | 0 | 0.019 | 0.28 |  |  | bbI | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 0.68 | 248.373 | 7951 | 50032.969 | 0 | 0.019 | 0.44 |  |  | bbI | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 0.68 | 230.388 | 7443 | 48869.313 | 0 | 0.019 | 0.71 |  |  | bbI | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 0.68 | 221.738 | 7076 | 49424.641 | 0 | 0.019 | 1.14 |  |  | bbI | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 0.68 | 233.228 | 7486 | 48781.555 | 0 | 0.019 | 1.82 |  |  | bbI | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 0.66 | 1502.766 | 25643 |  |  | 0.019 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 0.68 | 294.299 | 8434 | 50327.848 | 0 | 0.019 | 2.91 | 18.9639 | 551.68 | bb | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 0.68 | 247.44 | 7224 | 49203.574 | 0 | 0.019 | 4.66 |  |  | dbI | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 0.68 | 223.968 | 7078 | 47733.922 | 0 | 0.019 | 7.45 |  |  | bbI | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 0.68 | 296.6 | 8103 | 50520.066 | 0 | 0.019 | 11.92 | 20.5023 | 72 | bb | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 0.68 | 280.624 | 7793 | 49274.195 | 0 | 0.019 | 19.07 | 8.4712 | -55.58 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 0.68 | 300.632 | 8312 | 49934.551 | 0 | 0.019 | 30.52 | 30.0263 | -1.62 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 0.68 | 296.555 | 8557 | 50805.66 | 0 | 0.019 |  | 18.2593 |  | bb | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 0.68 | 355.055 | 7962 | 50913.094 | 0 | 0.019 | 48.83 | 80.8086 | 65.49 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 0.68 | 342.088 | 7422 | 48568.758 | 0 | 0.019 | 78.13 | 84.0448 | 7.57 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 0.68 | 452.098 | 8259 | 51406.383 | 0 | 0.019 | 125 | 153.5087 | 22.81 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 0.65 | 452.462 | 8218 | 51234.66 | 0 | 0.019 | 156.25 | 154.7787 | -0.94 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 0.65 | 563.38 | 11063 | 51287.395 | 0 | 0.019 | 250 | 221.5325 | -11.39 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 0.68 | 226.647 | 7164 | 49652.594 | 0 | 0.019 | 0.63 |  |  | bbI | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 0.68 | 295.704 | 8762 | 52099.977 | 0 | 0.019 | 2.5 | 7.0749 | 183 | bb | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 0.68 | 262.242 | 7624 | 51356.805 | 0 | 0.019 | 6.25 |  |  | MMI | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 0.68 | 309.147 | 8904 | 51164.73 | 0 | 0.019 | 25 | 31.3591 | 25.44 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 0.68 | 258.192 | 8212 | 50434.191 | 0 | 0.019 | 0.17 |  |  | bbI | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 0.68 | 281.091 | 8541 | 51354.762 | 0 | 0.019 | 0.28 |  |  | bbI | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 0.68 | 268.349 | 8280 | 51327.613 | 0 | 0.019 | 0.44 |  |  | bbI | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 0.68 | 238.684 | 7443 | 49470.234 | 0 | 0.019 | 0.71 |  |  | bbI | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 0.68 | 246.822 | 7753 | 49787.16 | 0 | 0.019 | 1.14 |  |  | bbI | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 0.68 | 242.172 | 8026 | 50241.621 | 0 | 0.019 | 1.82 |  |  | bbI | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 0.68 | 313.731 | 8913 | 53139.266 | 0 | 0.019 |  | 22.655 |  | bb | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 0.68 | 266.135 | 8182 | 51236.457 | 0 | 0.019 | 0.17 |  |  | bbI | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 0.68 | 262.542 | 7889 | 51647.742 | 0 | 0.019 | 0.28 |  |  | bbI | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 0.68 | 255.774 | 7779 | 51511.32 | 0 | 0.019 | 0.44 |  |  | bbI | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 0.68 | 248.871 | 7757 | 50949.309 | 0 | 0.019 | 0.71 |  |  | bbI | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 0.68 | 247.588 | 7721 | 50855.781 | 0 | 0.019 | 1.14 |  |  | bbI | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 0.68 | 253.325 | 8031 | 50099.988 | 0 | 0.019 | 1.82 |  |  | bbI | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 0.68 | 399.129 | 9404 | 55329.66 | 0 | 0.019 |  | 91.7593 |  | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 0.68 | 398.924 | 9809 | 55421.43 | 0 | 0.019 |  | 91.0618 |  | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 0.68 | 369.02 | 9798 | 52817.883 | 0 | 0.019 |  | 81.412 |  | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 0.68 | 464.598 | 9763 | 53182.844 | 0 | 0.019 |  | 151.458 |  | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 0.68 | 435.556 | 9527 | 53391.262 | 0 | 0.019 |  | 130.4092 |  | MM | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 0.68 | 425.336 | 9729 | 53434.508 | 0 | 0.019 |  | 122.8008 |  | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 0.68 | 318.624 | 8913 | 53001.285 | 0 | 0.019 |  | 29.4761 |  | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 0.68 | 405.582 | 9958 | 54932.02 | 0 | 0.019 |  | 99.1882 |  | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 0.68 | 381.698 | 9809 | 54176.371 | 0 | 0.019 |  | 84.1414 |  | bb | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 0.68 | 336.534 | 9306 | 52035.281 | 0 | 0.019 |  | 55.5798 |  | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 0.68 | 327.478 | 8181 | 54616.355 | 0 | 0.019 |  | 28.5026 |  | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 0.68 | 335.911 | 9067 | 53656.375 | 0 | 0.019 |  | 44.2118 |  | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 0.68 | 308.265 | 8142 | 51797.668 | 0 | 0.019 |  | 25.6937 |  | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 |  | Solvent |  |  |  | 4.41 |  | 0.019
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 0.65 | 549.797 | 10329 | 52322.02 | 0 | 0.019 |  | 207.8939 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 0.65 | 589.284 | 11168 | 53899.488 | 0 | 0.019 |  | 220.0781 |  | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 0.65 | 605.892 | 11355 | 53863.215 | 0 | 0.019 |  | 228.8591 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 0.68 | 327.587 | 9297 | 52084.168 | 0 | 0.019 |  | 45.8576 |  | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 0.68 | 318.111 | 8677 | 52470.254 | 0 | 0.019 |  | 32.6101 |  | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 0.68 | 322.782 | 8480 | 53591.719 | 0 | 0.019 |  | 30.1783 |  | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 0.65 | 1220.616 | 23424 |  |  | 0.019 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 0.68 | 384.081 | 9198 | 59098.656 | 0 | 0.019 |  | 57.2521 |  | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 0.68 | 336.718 | 8595 | 54825.996 | 0 | 0.019 |  | 37.3357 |  | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 0.68 | 352.189 | 8996 | 53523.332 | 0 | 0.019 |  | 61.4833 |  | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 0.65 | 685.725 | 14944 | 54352.184 | 0 | 0.019 |  | 264.6444 |  | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 0.65 | 603.88 | 10458 | 55028.941 | 0 | 0.019 |  | 221.2266 |  | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 0.65 | 559.077 | 9523 | 53986.629 | 0 | 0.019 |  | 203.4256 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 0.68 | 335.093 | 9093 | 53940.129 | 0 | 0.019 |  | 41.4621 |  | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 0.68 | 365.847 | 9218 | 54814.41 | 0 | 0.019 |  | 66.281 |  | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 0.68 | 336.965 | 8706 | 54693.875 | 0 | 0.019 |  | 38.4746 |  | bb | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 0.68 | 359.952 | 9353 | 55158.957 | 0 | 0.019 |  | 58.6574 |  | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 0.68 | 362.04 | 9832 | 54180.07 | 0 | 0.019 |  | 66.6769 |  | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 0.68 | 300.587 | 8760 | 53647.801 | 0 | 0.019 | 0.17 | 1.7211 | 912.44 | bb | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 0.68 | 303.198 | 8823 | 55028.125 | 0 | 0.019 | 0.28 |  |  | bbI | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 0.68 | 300.952 | 9208 | 53959.531 | 0 | 0.019 | 0.44 |  |  | bbI | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 0.68 | 318.037 | 9626 | 52453.168 | 0 | 0.019 | 0.71 | 32.6444 | 4497.8 | bb | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 0.68 | 269.703 | 8030 | 52932.809 | 0 | 0.019 | 1.14 |  |  | bbI | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 0.68 | 296.067 | 8883 | 53059.031 | 0 | 0.019 | 1.82 |  |  | bbI | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 0.65 | 1005.532 | 19754 |  |  | 0.019 |  |  |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 0.68 | 347.083 | 9945 | 54954.105 | 0 | 0.019 | 2.91 | 47.3291 | 1526.43 | bb | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 0.68 | 302.522 | 9008 | 53011.246 | 0 | 0.019 | 4.66 | 9.2974 | 99.51 | bb | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 0.68 | 281.234 | 8115 | 51497.086 | 0 | 0.019 | 7.45 |  |  | MMI | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 0.68 | 345.506 | 9428 | 54404.949 | 0 | 0.019 | 11.92 | 49.2555 | 313.22 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 0.68 | 379.932 | 9692 | 54264.176 | 0 | 0.019 | 19.07 | 82.1056 | 330.55 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 0.68 | 355.709 | 8975 | 53552.957 | 0 | 0.019 | 30.52 | 64.6595 | 111.86 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 0.68 | 359.962 | 9499 | 55176.629 | 0 | 0.019 |  | 58.5575 |  | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 0.68 | 329.829 | 8414 | 53773.059 | 0 | 0.019 | 48.83 | 36.8712 | -24.49 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 0.68 | 406.014 | 8995 | 52683.031 | 0 | 0.019 | 78.13 | 112.7151 | 44.27 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 0.68 | 507.918 | 9452 | 55300.879 | 0 | 0.019 | 125 | 166.7689 | 33.42 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 0.68 | 511.885 | 9733 | 53606.293 | 0 | 0.019 | 156.25 | 178.6402 | 14.33 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 0.65 | 616.254 | 11422 | 54873.789 | 0 | 0.019 | 250 | 228.3553 | -8.66 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 0.68 | 317.799 | 9637 | 54029.367 | 0 | 0.019 | 0.63 | 21.2253 | 3269.1 | bb | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 0.68 | 405.041 | 9916 | 61291.387 | 0 | 0.019 | 2.5 | 62.9399 | 2417.6 | bb | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 0.68 | 324.476 | 9044 | 53206.5 | 0 | 0.019 | 6.25 | 34.7677 | 456.28 | bb | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 0.68 | 338.175 | 9132 | 54984.762 | 0 | 0.019 | 25 | 37.8525 | 51.41 | bb | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 0.66 | 0.04 | 2 |  |  | 0.019 |  |  |  | MM | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 0.68 | 388.666 | 9676 | 57833.516 | 0 | 0.019 |  | 68.589 |  | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 0.68 | 348.056 | 9903 | 54130.621 | 0 | 0.019 | 0.17 | 53.5733 | 31413.73 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 0.68 | 312.158 | 8578 | 54825.285 | 0 | 0.019 | 0.28 | 8.3666 | 2888.06 | bb | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 0.68 | 324.751 | 8953 | 55636.879 | 0 | 0.019 | 0.44 | 18.2544 | 4048.73 | bb | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 0.68 | 311.763 | 9071 | 54984.301 | 0 | 0.019 | 0.71 | 6.6654 | 838.79 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 0.68 | 320.574 | 9344 | 54067.758 | 0 | 0.019 | 1.14 | 24.2768 | 2029.55 | bb | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 0.68 | 302.173 | 8886 | 53643.699 | 0 | 0.019 | 1.82 | 3.955 | 117.31 | bb | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 0.66 | 12.7 | 536 | 19403.883 | 0 | 0.019 |  |  |  | bbI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 0.68 | 336.485 | 9306 | 55147.813 | 0 | 0.019 | 0.17 | 34.9526 | 20460.34 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 0.68 | 353.841 | 9756 | 56832.938 | 0 | 0.019 | 0.28 | 42.2481 | 14988.61 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 0.68 | 341.82 | 9547 | 55401.938 | 0 | 0.019 | 0.44 | 38.9954 | 8762.59 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 0.68 | 331.014 | 9525 | 54257.391 | 0 | 0.019 | 0.71 | 34.9105 | 4816.97 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 0.68 | 314.462 | 8687 | 54348.875 | 0 | 0.019 | 1.14 | 14.8139 | 1199.46 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 0.68 | 319.549 | 9162 | 54800.516 | 0 | 0.019 | 1.82 | 17.8642 | 881.55 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 0.68 | 0.487 | 12 |  |  | 0.019 |  |  |  | MM | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 0.68 | 346.02 | 9443 | 55044.574 | 0 | 0.019 | 0.17 | 45.667 | 26762.94 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 0.68 | 374.919 | 9863 | 55853.539 | 0 | 0.019 | 0.28 | 68.1963 | 24255.81 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 0.68 | 337.014 | 9210 | 56901.824 | 0 | 0.019 | 0.44 | 23.8669 | 5324.3 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 0.68 | 334.85 | 9637 | 55579.008 | 0 | 0.019 | 0.71 | 30.2877 | 4165.88 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 0.68 | 321.177 | 8873 | 54417.223 | 0 | 0.019 | 1.14 | 22.5372 | 1876.94 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 0.68 | 353.421 | 10046 | 54960.121 | 0 | 0.019 | 1.82 | 53.6041 | 2845.28 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 0.67 | 0.042 | 2 |  |  | 0.019 |  |  |  | MM | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.019 |  |  |  |  | 2020-11-24
Compound 2: PFBS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 1.67 | 0.492 | 47 |  |  | 0.997 |  |  |  | MM | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 1.68 | 0.123 | 12 |  |  | 0.997 |  |  |  | MM | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 1.73 | 3.023 | 134 |  |  | 0.997 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [5 empty] | 0.997 |  |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.69 | 32.985 | 1541 | 21628.184 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 1.71 | 0.381 | 37 |  |  | 0.997 |  |  |  | MM | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.69 | 270.846 | 13568 | 21861.801 | 0 | 0.997 | 0.17 | 0.04426 | -73.96 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.69 | 765.518 | 39083 | 22905.938 | 0 | 0.997 | 0.28 | 0.36684 | 31.01 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1098.736 | 54638 | 22565.848 | 0 | 0.997 | 0.44 | 0.60105 | 36.6 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1545.47 | 79762 | 22329.166 | 0.001 | 0.997 | 0.71 | 0.91581 | 28.99 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1717.796 | 85594 | 22016.998 | 0.001 | 0.997 | 1.14 | 1.0509 | -7.82 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2849.414 | 143965 | 22032.42 | 0.001 | 0.997 | 1.82 | 1.8378 | 0.98 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 1.7 | 1.247 | 119 |  |  | 0.997 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.69 | 4048.259 | 203877 | 22609.912 | 0.002 | 0.997 | 2.91 | 2.6002 | -10.65 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.69 | 6474.464 | 332965 | 21791.783 | 0.003 | 0.997 | 4.66 | 4.4104 | -5.36 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.69 | 10672.343 | 534297 | 21652.168 | 0.005 | 0.997 | 7.45 | 7.4119 | -0.51 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.69 | 18260.215 | 923656 | 21857.189 | 0.008 | 0.997 | 11.92 | 12.6607 | 6.21 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.69 | 31758.162 | 1595521 | 22171.375 | 0.014 | 0.997 | 19.07 | 21.8017 | 14.32 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.69 | 44598.066 | 2214801 | 21427.695 | 0.021 | 0.997 | 30.52 | 31.7294 | 3.96 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.69 | 44.182 | 2081 | 22818.307 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.69 | 66633.859 | 3282871 | 20779.268 | 0.032 | 0.997 | 48.83 | 48.9239 | 0.19 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.69 | 101232.203 | 4914728 | 19147.74 | 0.053 | 0.997 | 78.13 | 80.6296 | 3.2 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.69 | 140618.609 | 6703322 | 18008.693 | 0.078 | 0.997 | 125 | 118.9318 | -4.85 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.69 | 174261.922 | 8197679 | 17279.416 | 0.101 | 0.997 | 156.25 | 153.3926 | -1.83 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.69 | 263390.281 | 12053817 | 15521.709 | 0.17 | 0.997 | 250 | 256.9115 | 2.76 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [5 empty] | 0.997 |  |  |  |  | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.69 | 960.737 | 45972 | 22865.168 | 0 | 0.997 | 0.63 | 0.4987 | -20.84 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.69 | 4319.657 | 220943 | 24861.91 | 0.002 | 0.997 | 2.5 | 2.5189 | 0.76 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.69 | 10008.617 | 505952 | 23299.66 | 0.004 | 0.997 | 6.25 | 6.441 | 3.06 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.69 | 37407.426 | 1871102 | 21679.879 | 0.017 | 0.997 | 25 | 26.2861 | 5.14 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 1.67 | 0.584 | 57 |  |  | 0.997 |  |  |  | MM | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.69 | 305.436 | 14419 | 23188.197 | 0 | 0.997 | 0.17 | 0.05627 | -66.9 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.69 | 794.594 | 40541 | 23955.818 | 0 | 0.997 | 0.28 | 0.36299 | 29.64 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1132.518 | 56657 | 23133.342 | 0 | 0.997 | 0.44 | 0.60512 | 37.53 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1597.693 | 81589 | 22994.361 | 0.001 | 0.997 | 0.71 | 0.91994 | 29.57 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1727.973 | 85843 | 22848.482 | 0.001 | 0.997 | 1.14 | 1.0142 | -11.04 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2846.754 | 143471 | 22995.889 | 0.001 | 0.997 | 1.82 | 1.7529 | -3.69 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.69 | 43.611 | 2163 | 22658.945 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.69 | 279.533 | 14261 | 22730.893 | 0 | 0.997 | 0.17 | 0.04286 | -74.79 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.69 | 779.21 | 39072 | 24510.729 | 0 | 0.997 | 0.28 | 0.34184 | 22.09 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1118.739 | 56886 | 24355.15 | 0 | 0.997 | 0.44 | 0.55878 | 27 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1584.861 | 79827 | 23756.479 | 0.001 | 0.997 | 0.71 | 0.87747 | 23.59 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1822.375 | 93105 | 23377.701 | 0.001 | 0.997 | 1.14 | 1.0499 | -7.91 | bb | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2931.468 | 148720 | 23030.449 | 0.001 | 0.997 | 1.82 | 1.8064 | -0.74 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 1.74 | 1.503 | 80 |  |  | 0.997 |  |  |  | bb | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.69 | 86.739 | 4274 | 25180.121 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.69 | 101.228 | 4709 | 25625.738 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.69 | 78.525 | 3265 | 24362.371 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.69 | 79.826 | 4069 | 24305.764 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.7 | 83.702 | 3666 | 24109.816 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.69 | 102.883 | 5027 | 24456.668 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.7 | 42.821 | 2080 | 23242.805 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.69 | 2653.8 | 134402 | 24957.41 | 0.001 | 0.997 |  | 1.4851 |  | MM | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.69 | 2918.392 | 146406 | 25152.914 | 0.001 | 0.997 |  | 1.6338 |  | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.69 | 3116.855 | 156086 | 24200.555 | 0.001 | 0.997 |  | 1.8296 |  | MM | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.69 | 38.017 | 1873 | 22887.045 | 0 | 0.997 |  |  |  | bdI | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.69 | 100.904 | 5253 | 22461.596 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.69 | 239.046 | 12328 | 23178.355 | 0 | 0.997 |  | 0.01242 |  | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [5 empty] | 0.997 |  |  |  |  | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.7 | 104.639 | 5040 | 22570.561 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.69 | 67.257 | 3414 | 23917.525 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.69 | 54.455 | 2625 | 23604.629 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.69 | 163018.234 | 7720060 | 17613.797 | 0.093 | 0.997 |  | 140.8451 |  | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.69 | 166278.828 | 7892064 | 17847.734 | 0.093 | 0.997 |  | 141.7737 |  | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.69 | 175612.5 | 8300231 | 18605.641 | 0.094 | 0.997 |  | 143.6215 |  | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 1.7 | 12.433 | 347 |  |  | 0.997 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.69 | 41.388 | 1913 | 26959.02 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.69 | 63.592 | 2794 | 23924.834 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.69 | 35.389 | 1826 | 24150.537 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.69 | 49.728 | 2225 | 23967.572 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.69 | 103.038 | 5567 | 23285.992 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.69 | 248.04 | 12986 | 24485.539 | 0 | 0.997 |  | 0.00961 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.69 | 46.517 | 2451 | 24092.141 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.69 | 176105.891 | 8342776 | 19292.301 | 0.091 | 0.997 |  | 138.9257 |  | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.69 | 172904.688 | 8105040 | 18221.59 | 0.095 | 0.997 |  | 144.3828 |  | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.69 | 173770.672 | 8266025 | 18441.566 | 0.094 | 0.997 |  | 143.381 |  | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.7 | 84.043 | 2658 | 24051.273 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [5 empty] | 0.997 |  |  |  |  | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.69 | 284.038 | 12562 | 24736.178 | 0 | 0.997 | 0.17 | 0.03036 | -82.14 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.69 | 849.65 | 42095 | 25389.068 | 0 | 0.997 | 0.28 | 0.36753 | 31.26 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1192.203 | 59300 | 24953.592 | 0 | 0.997 | 0.44 | 0.58704 | 33.42 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1631.263 | 82380 | 24666.17 | 0.001 | 0.997 | 0.71 | 0.86858 | 22.34 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1962.02 | 99118 | 24743.717 | 0.001 | 0.997 | 1.14 | 1.0704 | -6.1 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3003.684 | 151614 | 24207.008 | 0.001 | 0.997 | 1.82 | 1.7573 | -3.44 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 1.7 | 2.178 | 166 |  |  | 0.997 |  |  |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.69 | 4509.937 | 226581 | 24554.58 | 0.002 | 0.997 | 2.91 | 2.6711 | -8.21 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.69 | 7021.911 | 350201 | 23955.301 | 0.003 | 0.997 | 4.66 | 4.3494 | -6.67 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.69 | 11412.911 | 576123 | 23626.941 | 0.005 | 0.997 | 7.45 | 7.2608 | -2.54 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.69 | 19943.035 | 1004406 | 24218.33 | 0.008 | 0.997 | 11.92 | 12.4774 | 4.68 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.69 | 34046.984 | 1690798 | 24038.738 | 0.014 | 0.997 | 19.07 | 21.556 | 13.04 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.69 | 47851.457 | 2390021 | 23164.816 | 0.021 | 0.997 | 30.52 | 31.4904 | 3.18 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.69 | 56.11 | 2585 | 24598.893 | 0 | 0.997 |  |  |  | bdI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.69 | 72477.875 | 3580599 | 22853.115 | 0.032 | 0.997 | 48.83 | 48.3853 | -0.91 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.69 | 106142.273 | 5134447 | 20442.584 | 0.052 | 0.997 | 78.13 | 79.1885 | 1.35 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.69 | 148861.125 | 7077424 | 19474.186 | 0.076 | 0.997 | 125 | 116.4396 | -6.85 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.69 | 187885.969 | 8887854 | 18660.311 | 0.101 | 0.997 | 156.25 | 153.1479 | -1.99 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.69 | 281528.156 | 12888545 | 16898.207 | 0.167 | 0.997 | 250 | 252.2882 | 0.92 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 1.66 | 27.867 | 667 |  |  | 0.997 |  |  |  | bb | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.69 | 1062.019 | 52544 | 24217.922 | 0 | 0.997 | 0.63 | 0.52685 | -16.37 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.69 | 5109.564 | 255476 | 29108.41 | 0.002 | 0.997 | 2.5 | 2.5464 | 1.85 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.69 | 10634.096 | 531629 | 24923.5 | 0.004 | 0.997 | 6.25 | 6.3967 | 2.35 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.69 | 40760.316 | 2015584 | 24550.033 | 0.017 | 0.997 | 25 | 25.2894 | 1.16 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 1.72 | 4.584 | 217 |  |  | 0.997 |  |  |  | bb | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.69 | 42.679 | 1983 | 25048.125 | 0 | 0.997 |  |  |  | bbI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.69 | 310.445 | 16342 | 24669.846 | 0 | 0.997 | 0.17 | 0.04725 | -72.21 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.69 | 856.988 | 42860 | 26169.432 | 0 | 0.997 | 0.28 | 0.35652 | 27.33 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1220.072 | 61537 | 26166.074 | 0 | 0.997 | 0.44 | 0.56942 | 29.41 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1693.45 | 87004 | 25000.137 | 0.001 | 0.997 | 0.71 | 0.89318 | 25.8 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1954.207 | 99743 | 25457.678 | 0.001 | 0.997 | 1.14 | 1.0316 | -9.51 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2986.899 | 153470 | 24884.777 | 0.001 | 0.997 | 1.82 | 1.6951 | -6.86 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.69 | 17.689 | 933 | 8914.928 | 0 | 0.997 |  |  |  | bbI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.69 | 332.996 | 16342 | 25334.619 | 0 | 0.997 | 0.17 | 0.05584 | -67.15 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.69 | 860.977 | 41948 | 26174.143 | 0 | 0.997 | 0.28 | 0.35877 | 28.13 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1230.669 | 61669 | 26063.229 | 0 | 0.997 | 0.44 | 0.57848 | 31.47 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1718.22 | 87933 | 25239.68 | 0.001 | 0.997 | 0.71 | 0.89837 | 26.53 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.69 | 2008.324 | 103481 | 25575.846 | 0.001 | 0.997 | 1.14 | 1.0586 | -7.14 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3014.977 | 149427 | 25309.125 | 0.001 | 0.997 | 1.82 | 1.6813 | -7.62 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 1.73 | 5.926 | 246 |  |  | 0.997 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.69 | 298.9 | 15508 | 25002.83 | 0 | 0.997 | 0.17 | 0.0376 | -77.88 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.69 | 847.982 | 42497 | 26803.898 | 0 | 0.997 | 0.28 | 0.33948 | 21.24 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1225.297 | 61938 | 26872.852 | 0 | 0.997 | 0.44 | 0.55359 | 25.82 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1736.466 | 86665 | 25737.27 | 0.001 | 0.997 | 0.71 | 0.88906 | 25.22 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1967.181 | 101259 | 25205.348 | 0.001 | 0.997 | 1.14 | 1.0513 | -7.78 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3135.531 | 160453 | 25344.141 | 0.001 | 0.997 | 1.82 | 1.7517 | -3.75 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 1.69 | 2.708 | 188 |  |  | 0.997 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 1.76 | 2.691 | 182 |  |  | 0.997 |  |  |  | bd | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.997 |  |  |  |  | 2020-11-24
Compound 3: PFHpA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 1.82 | 97.424 | 2181 |  |  | 0.998 |  |  |  | bb | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 1.83 | 143.086 | 2718 |  |  | 0.998 |  |  |  | bb | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 1.83 | 82.237 | 2157 |  |  | 0.998 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 1.76 | 20.155 | 721 |  |  | 0.998 |  |  |  | MM | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.76 | 113.106 | 5453 | 66453.484 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 1.76 | 19.311 | 622 |  |  | 0.998 |  |  |  | MM | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.76 | 850.71 | 43330 | 66855.516 | 0 | 0.998 | 0.17 | 0.14193 | -16.51 | bb | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1596.522 | 80606 | 70175.805 | 0 | 0.998 | 0.28 | 0.2748 | -1.86 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2553.536 | 127184 | 69660.461 | 0 | 0.998 | 0.44 | 0.4591 | 4.34 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.76 | 3685.775 | 190921 | 69179.859 | 0.001 | 0.998 | 0.71 | 0.67937 | -4.31 | bb | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.76 | 5816.126 | 293972 | 65949.148 | 0.001 | 0.998 | 1.14 | 1.142 | 0.18 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.76 | 9600.393 | 482368 | 65820.242 | 0.001 | 0.998 | 1.82 | 1.9061 | 4.73 | bb | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 1.77 | 20.274 | 901 |  |  | 0.998 |  |  |  | MM | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.76 | 13913.541 | 696557 | 69313.656 | 0.002 | 0.998 | 2.91 | 2.6332 | -9.51 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.76 | 22342.369 | 1126751 | 67371.406 | 0.003 | 0.998 | 4.66 | 4.3672 | -6.28 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.76 | 36733.152 | 1808733 | 65924.273 | 0.006 | 0.998 | 7.45 | 7.3546 | -1.28 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.76 | 64327.168 | 3108039 | 66429.594 | 0.01 | 0.998 | 11.92 | 12.7966 | 7.35 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.76 | 110352.383 | 5346011 | 67007.094 | 0.016 | 0.998 | 19.07 | 21.7694 | 14.16 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.76 | 157745.297 | 7601439 | 65890.977 | 0.024 | 0.998 | 30.52 | 31.638 | 3.66 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.76 | 121.095 | 5130 | 70428.734 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.76 | 234145.531 | 11241628 | 64077.266 | 0.037 | 0.998 | 48.83 | 48.2531 | -1.18 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.76 | 359678.344 | 16933280 | 59706.527 | 0.06 | 0.998 | 78.13 | 79.4088 | 1.64 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.76 | 507858.75 | 23705804 | 56827.754 | 0.089 | 0.998 | 125 | 117.5312 | -5.98 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.76 | 645238.188 | 29797752 | 54128.734 | 0.119 | 0.998 | 156.25 | 156.393 | 0.09 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.76 | 984418 | 44298656 | 50352.082 | 0.196 | 0.998 | 250 | 254.9259 | 1.97 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 1.75 | 124.244 | 4163 |  |  | 0.998 |  |  |  | MM | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.76 | 3246.877 | 163527 | 69882.695 | 0 | 0.998 | 0.63 | 0.58904 | -6.5 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.76 | 15453.368 | 764193 | 74668.664 | 0.002 | 0.998 | 2.5 | 2.7157 | 8.63 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.76 | 34994.387 | 1718356 | 71009.594 | 0.005 | 0.998 | 6.25 | 6.5019 | 4.03 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.76 | 131310.563 | 6424507 | 66898.727 | 0.02 | 0.998 | 25 | 25.944 | 3.78 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 1.76 | 18.022 | 999 |  |  | 0.998 |  |  |  | MM | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.76 | 878.514 | 46919 | 70140.469 | 0 | 0.998 | 0.17 | 0.13929 | -18.07 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1666.76 | 81526 | 73291.031 | 0 | 0.998 | 0.28 | 0.27469 | -1.9 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2730.885 | 138122 | 71199 | 0 | 0.998 | 0.44 | 0.48161 | 9.46 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4133.796 | 204342 | 70969.344 | 0.001 | 0.998 | 0.71 | 0.74522 | 4.96 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6447.323 | 312789 | 69821.609 | 0.001 | 0.998 | 1.14 | 1.197 | 5 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.76 | 9697.646 | 488360 | 70147.031 | 0.001 | 0.998 | 1.82 | 1.8053 | -0.81 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.76 | 78.945 | 3535 | 70267.711 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.76 | 910.793 | 47778 | 69550.805 | 0 | 0.998 | 0.17 | 0.14684 | -13.62 | bb | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1679.931 | 85479 | 72982.797 | 0 | 0.998 | 0.28 | 0.27835 | -0.59 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2804.495 | 142603 | 73203.641 | 0 | 0.998 | 0.44 | 0.48102 | 9.32 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.76 | 3977.837 | 198806 | 71708.977 | 0.001 | 0.998 | 0.71 | 0.70844 | -0.22 | bb | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6033.968 | 310688 | 69097.773 | 0.001 | 0.998 | 1.14 | 1.1305 | -0.83 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10209.131 | 501514 | 69563.125 | 0.001 | 0.998 | 1.82 | 1.9181 | 5.39 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 1.8 | 4.15 | 234 |  |  | 0.998 |  |  |  | MM | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.76 | 751.544 | 39201 | 75833.383 | 0 | 0.998 |  | 0.10463 |  | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.76 | 777.696 | 39627 | 78561.383 | 0 | 0.998 |  | 0.10449 |  | db | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.76 | 971.488 | 49541 | 75786.555 | 0 | 0.998 |  | 0.14318 |  | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.76 | 70.381 | 3697 | 75532.336 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.76 | 115.072 | 5839 | 73999.875 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.76 | 102.199 | 4557 | 73473.383 | 0 | 0.998 |  |  |  | bdI | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.76 | 104.814 | 5277 | 71625.82 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.76 | 70.733 | 3571 | 72956.234 | 0 | 0.998 |  |  |  | ddI | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.76 | 93.345 | 4359 | 75183.023 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.76 | 103.114 | 4817 | 72624.836 | 0 | 0.998 |  |  |  | bdI | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.76 | 545587.125 | 25399062 | 57842.676 | 0.094 | 0.998 |  | 123.9977 |  | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.76 | 576372.25 | 26675222 | 55906.832 | 0.103 | 0.998 |  | 135.4343 |  | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.76 | 579618.688 | 27357730 | 57936.676 | 0.1 | 0.998 |  | 131.4578 |  | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 1.76 | 88.152 | 2940 |  |  | 0.998 |  |  |  | MM | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.76 | 140.486 | 7038 | 68748.156 | 0 | 0.998 |  | 0.00037 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.76 | 249.245 | 12045 | 73999.633 | 0 | 0.998 |  | 0.01793 |  | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.76 | 181.4 | 9688 | 71143.813 | 0 | 0.998 |  | 0.00708 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.76 | 129.401 | 6930 | 69815.391 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.76 | 105.802 | 4441 | 71290.555 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.76 | 108.512 | 4980 | 74547.953 | 0 | 0.998 |  |  |  | bdI | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 1.77 | 28.679 | 1180 |  |  | 0.998 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.76 | 616916.688 | 28676648 | 65037.824 | 0.095 | 0.998 |  | 124.6923 |  | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.76 | 599925.25 | 27823716 | 58343.555 | 0.103 | 0.998 |  | 135.0841 |  | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.76 | 600657.063 | 28247480 | 59147.715 | 0.102 | 0.998 |  | 133.4238 |  | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.76 | 254.678 | 11025 | 71500.539 | 0 | 0.998 |  | 0.0205 |  | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.76 | 239.811 | 11703 | 70718.625 | 0 | 0.998 |  | 0.01823 |  | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.76 | 172.22 | 8458 | 71977.289 | 0 | 0.998 |  | 0.005 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.76 | 142.571 | 5999 | 73164.828 | 0 | 0.998 |  |  |  | bdI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.76 | 108.144 | 4597 | 75227.086 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.76 | 92.977 | 4295 | 72306.594 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.76 | 107.86 | 5319 | 72174.523 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.76 | 122.398 | 6077 | 72886.563 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 1.76 | 17.823 | 784 |  |  | 0.998 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.76 | 965.351 | 49330 | 73650.867 | 0 | 0.998 | 0.17 | 0.147 | -13.53 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1788.739 | 90149 | 75305.336 | 0 | 0.998 | 0.28 | 0.28809 | 2.89 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3039.829 | 147500 | 75526.563 | 0 | 0.998 | 0.44 | 0.5067 | 15.16 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4200.137 | 215512 | 74249.516 | 0.001 | 0.998 | 0.71 | 0.72296 | 1.83 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6841.671 | 335362 | 74140.367 | 0.001 | 0.998 | 1.14 | 1.1962 | 4.93 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10747.685 | 524962 | 72727.977 | 0.001 | 0.998 | 1.82 | 1.9316 | 6.13 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 1.76 | 10.256 | 686 |  |  | 0.998 |  |  |  | MM | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.76 | 15120.782 | 752900 | 73489.727 | 0.002 | 0.998 | 2.91 | 2.6997 | -7.23 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.76 | 25136.83 | 1237218 | 72791.852 | 0.003 | 0.998 | 4.66 | 4.5486 | -2.39 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.76 | 38859.742 | 1924695 | 69266.383 | 0.006 | 0.998 | 7.45 | 7.4051 | -0.6 | bb | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.76 | 69992.758 | 3395560 | 72812.047 | 0.01 | 0.998 | 11.92 | 12.703 | 6.57 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.76 | 116698.813 | 5599537 | 70597.922 | 0.017 | 0.998 | 19.07 | 21.8504 | 14.58 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.76 | 165116.125 | 7948942 | 69640.266 | 0.024 | 0.998 | 30.52 | 31.3337 | 2.67 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.76 | 135.832 | 6397 | 74817.703 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.76 | 253357.641 | 12204068 | 67770.406 | 0.037 | 0.998 | 48.83 | 49.3642 | 1.09 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.76 | 367405.875 | 17545514 | 61003.473 | 0.06 | 0.998 | 78.13 | 79.3904 | 1.61 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.76 | 530252.313 | 24655996 | 59756.648 | 0.089 | 0.998 | 125 | 116.705 | -6.64 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.76 | 667760.563 | 30854474 | 57883.109 | 0.115 | 0.998 | 156.25 | 151.4011 | -3.1 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.76 | 1034223.938 | 46612284 | 53039.328 | 0.195 | 0.998 | 250 | 254.2649 | 1.71 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 1.75 | 151.183 | 4824 |  |  | 0.998 |  |  |  | MM | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.76 | 3587.323 | 179450 | 73898.461 | 0 | 0.998 | 0.63 | 0.61663 | -2.12 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.76 | 17789.33 | 869332 | 86813.352 | 0.002 | 0.998 | 2.5 | 2.6886 | 7.54 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.76 | 37541.043 | 1824743 | 74901.398 | 0.005 | 0.998 | 6.25 | 6.6131 | 5.81 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.76 | 141971.547 | 6814487 | 73049.375 | 0.019 | 0.998 | 25 | 25.6887 | 2.75 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 1.76 | 26.737 | 1220 |  |  | 0.998 |  |  |  | MM | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.76 | 131.084 | 6036 | 77838.516 | 0 | 0.998 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.76 | 914.414 | 47027 | 75592.711 | 0 | 0.998 | 0.17 | 0.13361 | -21.41 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1758.781 | 87451 | 78736.781 | 0 | 0.998 | 0.28 | 0.26933 | -3.81 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3006.335 | 153274 | 78371.031 | 0 | 0.998 | 0.44 | 0.48167 | 9.47 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4212.83 | 209553 | 74719.867 | 0.001 | 0.998 | 0.71 | 0.7205 | 1.48 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6562.722 | 325605 | 75418.586 | 0.001 | 0.998 | 1.14 | 1.1265 | -1.19 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10626.602 | 533370 | 72627.703 | 0.001 | 0.998 | 1.82 | 1.9122 | 5.07 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.76 | 43.671 | 1682 | 26529.941 | 0 | 0.998 |  |  |  | MMI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.76 | 931.982 | 45388 | 75935.922 | 0 | 0.998 | 0.17 | 0.13595 | -20.03 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1817.852 | 89845 | 77973.492 | 0 | 0.998 | 0.28 | 0.28227 | 0.81 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2829.649 | 139779 | 77142.375 | 0 | 0.998 | 0.44 | 0.45942 | 4.41 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4355.408 | 219745 | 76099.219 | 0.001 | 0.998 | 0.71 | 0.73178 | 3.07 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6592.64 | 331029 | 76403.75 | 0.001 | 0.998 | 1.14 | 1.1168 | -2.04 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10844.135 | 536457 | 72442.719 | 0.001 | 0.998 | 1.82 | 1.9569 | 7.52 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 1.81 | 127.509 | 2806 |  |  | 0.998 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.76 | 1011.845 | 51110 | 74628.094 | 0 | 0.998 | 0.17 | 0.15298 | -10.01 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1778.642 | 85922 | 79756.063 | 0 | 0.998 | 0.28 | 0.26885 | -3.98 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3093.993 | 153457 | 79699.555 | 0 | 0.998 | 0.44 | 0.48777 | 10.86 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4305.967 | 212324 | 76448.648 | 0.001 | 0.998 | 0.71 | 0.71974 | 1.37 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6798.62 | 334024 | 74293.961 | 0.001 | 0.998 | 1.14 | 1.186 | 4.04 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10520.271 | 523345 | 75961.703 | 0.001 | 0.998 | 1.82 | 1.8086 | -0.63 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 1.76 | 26.945 | 1365 |  |  | 0.998 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 1.77 | 17.491 | 895 |  |  | 0.998 |  |  |  | MM | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 1.78 | 9.957 | 400 |  |  | 0.998 |  |  |  | MM | 2020-11-24
Compound 4: PFDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.25 | 181.391 | 1809 |  |  | 0.998 |  |  |  | bb | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.28 | 46.525 | 988 |  |  | 0.998 |  |  |  | bb | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.3 | 0.717 | 60 |  |  | 0.998 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.31 | 44.99 | 710 |  |  | 0.998 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.29 | 150.075 | 2767 | 37792.707 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.28 | 43.526 | 1093 |  |  | 0.998 |  |  |  | bb | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.29 | 812.475 | 26234 | 37339.707 | 0 | 0.998 | 0.17 | 0.1442 | -15.17 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1422.246 | 46842 | 37559.73 | 0 | 0.998 | 0.28 | 0.28291 | 1.04 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2425.374 | 80815 | 38314.09 | 0.001 | 0.998 | 0.44 | 0.50194 | 14.08 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3344.134 | 108165 | 38319.367 | 0.001 | 0.998 | 0.71 | 0.70833 | -0.24 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5157.513 | 167565 | 35923.289 | 0.001 | 0.998 | 1.14 | 1.1931 | 4.66 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.29 | 7982.595 | 263229 | 36542.281 | 0.002 | 0.998 | 1.82 | 1.8378 | 0.98 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.34 | 19.874 | 743 |  |  | 0.998 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.29 | 12157.958 | 400542 | 37978.066 | 0.003 | 0.998 | 2.91 | 2.7133 | -6.76 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.29 | 19602.594 | 620168 | 37018.879 | 0.005 | 0.998 | 4.66 | 4.516 | -3.09 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.29 | 30781.998 | 977786 | 34951.309 | 0.009 | 0.998 | 7.45 | 7.5387 | 1.19 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.29 | 53970.766 | 1724786 | 36311.348 | 0.015 | 0.998 | 11.92 | 12.7499 | 6.96 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.29 | 93540.961 | 2970972 | 37384.508 | 0.025 | 0.998 | 19.07 | 21.4855 | 12.67 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.29 | 136254.234 | 4337225 | 37091.914 | 0.037 | 0.998 | 30.52 | 31.5511 | 3.38 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.29 | 103.346 | 3479 | 38336.848 | 0 | 0.998 |  |  |  | bdI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.29 | 203442.953 | 6492925 | 35457.441 | 0.057 | 0.998 | 48.83 | 49.2709 | 0.9 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.29 | 307602.375 | 9752293 | 33010.672 | 0.093 | 0.998 | 78.13 | 79.9495 | 2.33 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.29 | 429381.469 | 13474437 | 31009.475 | 0.138 | 0.998 | 125 | 118.645 | -5.08 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.29 | 533271.25 | 16663163 | 29707.164 | 0.18 | 0.998 | 156.25 | 153.6141 | -1.69 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.29 | 821408.875 | 25130938 | 27123.66 | 0.303 | 0.998 | 250 | 258.1306 | 3.25 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.29 | 205.321 | 5267 |  |  | 0.998 |  |  |  | bb | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.29 | 2693.958 | 89714 | 38590.461 | 0.001 | 0.998 | 0.63 | 0.55797 | -11.43 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.29 | 12772.691 | 406187 | 40285.297 | 0.003 | 0.998 | 2.5 | 2.6868 | 7.47 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.29 | 30260.361 | 977413 | 39404.684 | 0.008 | 0.998 | 6.25 | 6.5682 | 5.09 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.29 | 113406.469 | 3612227 | 36878.289 | 0.031 | 0.998 | 25 | 26.4108 | 5.64 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.32 | 45.597 | 1050 |  |  | 0.998 |  |  |  | bb | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.29 | 715.83 | 22620 | 37417.016 | 0 | 0.998 | 0.17 | 0.12157 | -28.49 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1486.707 | 49406 | 40856.688 | 0 | 0.998 | 0.28 | 0.27018 | -3.51 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2629.677 | 86066 | 39381.18 | 0.001 | 0.998 | 0.44 | 0.53184 | 20.87 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3486.108 | 112718 | 39899.059 | 0.001 | 0.998 | 0.71 | 0.70921 | -0.11 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5235.55 | 174282 | 39440.199 | 0.001 | 0.998 | 1.14 | 1.0999 | -3.52 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8397.799 | 272370 | 39243.359 | 0.002 | 0.998 | 1.82 | 1.7995 | -1.13 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.29 | 133.816 | 4106 | 38771.027 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.29 | 793.24 | 25147 | 38646.242 | 0 | 0.998 | 0.17 | 0.13358 | -21.42 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1580.923 | 52199 | 42048.52 | 0 | 0.998 | 0.28 | 0.28059 | 0.21 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2479.736 | 80126 | 39681.668 | 0.001 | 0.998 | 0.44 | 0.49495 | 12.49 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3425.66 | 112990 | 38738.613 | 0.001 | 0.998 | 0.71 | 0.71832 | 1.17 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5363 | 172030 | 37568.254 | 0.001 | 0.998 | 1.14 | 1.1861 | 4.04 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8747.275 | 279911 | 38877.977 | 0.002 | 0.998 | 1.82 | 1.8942 | 4.08 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.3 | 53.246 | 931 |  |  | 0.998 |  |  |  | bb | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.29 | 4536.727 | 148347 | 42185.547 | 0.001 | 0.998 |  | 0.88289 |  | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.29 | 5222.466 | 171073 | 43559.676 | 0.001 | 0.998 |  | 0.98923 |  | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.29 | 4229.53 | 140333 | 41573.668 | 0.001 | 0.998 |  | 0.83289 |  | MM | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.29 | 178.196 | 4457 | 42123.012 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.29 | 167.629 | 4287 | 41354.645 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.29 | 146.175 | 3256 | 39860.156 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.29 | 113.565 | 3097 | 38245.285 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.29 | 159.547 | 3706 | 40662.246 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.29 | 185.232 | 4299 | 41766.652 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.29 | 128.842 | 3062 | 38594.965 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.29 | 448711.094 | 13890057 | 31312.432 | 0.143 | 0.998 |  | 122.7682 |  | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.29 | 482083.438 | 14960000 | 31432.322 | 0.153 | 0.998 |  | 131.3548 |  | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.29 | 500682.344 | 15483513 | 31703.865 | 0.158 | 0.998 |  | 135.2349 |  | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.29 | 92.209 | 1708 |  |  | 0.998 |  |  |  | bb | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.29 | 146.408 | 4941 | 36568.078 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.29 | 209.484 | 7209 | 39947.965 | 0 | 0.998 |  | 0.00199 |  | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.29 | 201.569 | 6163 | 39242.902 | 0 | 0.998 |  | 0.00106 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.29 | 156.397 | 3979 | 37110.707 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.29 | 84.559 | 2439 | 38090.129 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.29 | 113.601 | 2361 | 39006.359 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.33 | 49.65 | 989 |  |  | 0.998 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.29 | 540588 | 16791414 | 36889.391 | 0.147 | 0.998 |  | 125.5328 |  | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.29 | 532659.75 | 16598030 | 33918.5 | 0.157 | 0.998 |  | 134.4819 |  | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.29 | 524709.313 | 16397762 | 33629.922 | 0.156 | 0.998 |  | 133.6157 |  | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.29 | 175.618 | 5829 | 39910.887 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.29 | 190.796 | 6047 | 39164.621 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.29 | 110.206 | 2674 | 39562.254 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.29 | 151.703 | 4265 | 41662.426 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.29 | 94.143 | 2920 | 39324.82 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.29 | 87.277 | 2357 | 37500.363 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.29 | 73.969 | 2556 | 38443.734 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.29 | 108.92 | 2283 | 40001.32 | 0 | 0.998 |  |  |  | bbI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.33 | 40.784 | 1301 |  |  | 0.998 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.29 | 775.351 | 24917 | 39172.805 | 0 | 0.998 | 0.17 | 0.12727 | -25.13 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1632.62 | 54329 | 40422.34 | 0 | 0.998 | 0.28 | 0.30463 | 8.8 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2642.395 | 87039 | 41193.734 | 0.001 | 0.998 | 0.44 | 0.5092 | 15.73 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3584.333 | 119301 | 41931.879 | 0.001 | 0.998 | 0.71 | 0.69291 | -2.41 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5677.545 | 185915 | 40251.609 | 0.001 | 0.998 | 1.14 | 1.1714 | 2.76 | bb | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8584.184 | 278202 | 39630.223 | 0.002 | 0.998 | 1.82 | 1.822 | 0.11 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.31 | 26.628 | 1282 |  |  | 0.998 |  |  |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.29 | 13500.459 | 429025 | 41927.211 | 0.003 | 0.998 | 2.91 | 2.7294 | -6.21 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.29 | 20670.557 | 664651 | 40052.781 | 0.005 | 0.998 | 4.66 | 4.4003 | -5.57 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.29 | 34151.543 | 1101646 | 40052.25 | 0.009 | 0.998 | 7.45 | 7.2975 | -2.05 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.29 | 57101.934 | 1824724 | 38973.023 | 0.015 | 0.998 | 11.92 | 12.5677 | 5.43 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.29 | 102659.703 | 3279173 | 40601.832 | 0.025 | 0.998 | 19.07 | 21.7118 | 13.85 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.29 | 142710.609 | 4517513 | 40061.527 | 0.036 | 0.998 | 30.52 | 30.5964 | 0.25 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.29 | 80.47 | 2644 | 42510.348 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.29 | 213439.125 | 6869003 | 37898.871 | 0.056 | 0.998 | 48.83 | 48.3628 | -0.96 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.29 | 321368.563 | 10134493 | 33821.711 | 0.095 | 0.998 | 78.13 | 81.5203 | 4.34 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.29 | 456410.469 | 14321911 | 33647.129 | 0.136 | 0.998 | 125 | 116.2373 | -7.01 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.29 | 570029.625 | 17655558 | 32227.65 | 0.177 | 0.998 | 156.25 | 151.3732 | -3.12 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.29 | 867564.75 | 26367316 | 29370.512 | 0.295 | 0.998 | 250 | 251.839 | 0.74 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.29 | 197.565 | 5385 |  |  | 0.998 |  |  |  | MM | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.29 | 3139.04 | 97125 | 40316.699 | 0.001 | 0.998 | 0.63 | 0.62729 | -0.43 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.29 | 14056.702 | 463083 | 46858.434 | 0.003 | 0.998 | 2.5 | 2.5398 | 1.59 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.29 | 32222.326 | 1026852 | 43284.461 | 0.007 | 0.998 | 6.25 | 6.3658 | 1.85 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.29 | 119831.102 | 3818558 | 40310.613 | 0.03 | 0.998 | 25 | 25.5303 | 2.12 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.29 | 54.728 | 1442 |  |  | 0.998 |  |  |  | bb | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.29 | 113.05 | 3512 | 42210.781 | 0 | 0.998 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.29 | 844.023 | 28432 | 41901.957 | 0 | 0.998 | 0.17 | 0.13029 | -23.36 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1627.655 | 53157 | 44425.273 | 0 | 0.998 | 0.28 | 0.27233 | -2.74 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2656.241 | 86361 | 41899.441 | 0.001 | 0.998 | 0.44 | 0.50274 | 14.26 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3708.542 | 124160 | 42107.133 | 0.001 | 0.998 | 0.71 | 0.71525 | 0.74 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5653.892 | 185675 | 41405.266 | 0.001 | 0.998 | 1.14 | 1.1327 | -0.64 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9366.021 | 295213 | 41834.133 | 0.002 | 0.998 | 1.82 | 1.8846 | 3.55 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.29 | 75.13 | 1928 | 14514.265 | 0 | 0.998 |  | 0.0014 |  | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.29 | 892.164 | 29991 | 40732.211 | 0 | 0.998 | 0.17 | 0.14545 | -14.44 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1658.027 | 53910 | 43092.957 | 0 | 0.998 | 0.28 | 0.28815 | 2.91 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2750.117 | 90688 | 42587.418 | 0.001 | 0.998 | 0.44 | 0.51291 | 16.57 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3710.15 | 123010 | 43136.676 | 0.001 | 0.998 | 0.71 | 0.69747 | -1.76 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5641.129 | 182794 | 39118.82 | 0.001 | 0.998 | 1.14 | 1.1986 | 5.14 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9061.154 | 297179 | 40277.5 | 0.002 | 0.998 | 1.82 | 1.894 | 4.06 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.27 | 20.866 | 493 |  |  | 0.998 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.29 | 729.499 | 23738 | 41387.645 | 0 | 0.998 | 0.17 | 0.10861 | -36.11 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1746.708 | 56488 | 45184.34 | 0 | 0.998 | 0.28 | 0.28972 | 3.47 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2783.193 | 90359 | 42725.34 | 0.001 | 0.998 | 0.44 | 0.51778 | 17.68 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3690.172 | 122912 | 41798.488 | 0.001 | 0.998 | 0.71 | 0.71706 | 0.99 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5829.492 | 183732 | 41191.141 | 0.001 | 0.998 | 1.14 | 1.1755 | 3.11 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9407.359 | 295921 | 42275.109 | 0.002 | 0.998 | 1.82 | 1.873 | 2.91 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.25 | 50.808 | 736 |  |  | 0.998 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 2.25 | 70.522 | 1242 |  |  | 0.998 |  |  |  | bb | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.36 | 73.566 | 965 |  |  | 0.998 |  |  |  | bb | 2020-11-24
Compound 5: PFOS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.4 | 1.81 | 69 |  |  | 0.998 |  |  |  | MM | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.38 | 1.317 | 60 |  |  | 0.998 |  |  |  | bb | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.38 | 1.662 | 81 |  |  | 0.998 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [5 empty] | 0.998 |  |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.39 | 318.799 | 5621 | 3612.967 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.38 | 1.939 | 95 |  |  | 0.998 |  |  |  | bb | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.39 | 540.821 | 10969 | 3515.653 | 0.002 | 0.998 | 0.17 | 0.12168 | -28.42 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.39 | 684.029 | 14818 | 3777.996 | 0.002 | 0.998 | 0.28 | 0.19764 | -29.41 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1153.145 | 30019 | 3699.702 | 0.003 | 0.998 | 0.44 | 0.56207 | 27.74 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1326.21 | 34770 | 3810.099 | 0.003 | 0.998 | 0.71 | 0.66358 | -6.54 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.39 | 1883.786 | 47481 | 3518.784 | 0.005 | 0.998 | 1.14 | 1.1859 | 4.02 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.39 | 2778.222 | 73282 | 3525.893 | 0.008 | 0.998 | 1.82 | 1.8901 | 3.85 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [5 empty] | 0.998 |  |  |  |  | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.39 | 4068.736 | 109708 | 3792.97 | 0.011 | 0.998 | 2.91 | 2.6836 | -7.78 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.38 | 6271.539 | 167990 | 3544.214 | 0.018 | 0.998 | 4.66 | 4.6238 | -0.78 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.39 | 9385.291 | 253109 | 3472.039 | 0.027 | 0.998 | 7.45 | 7.22 | -3.09 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.39 | 16472.125 | 450029 | 3420.079 | 0.048 | 0.998 | 11.92 | 13.0821 | 9.75 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.39 | 28893.621 | 771561 | 3621.232 | 0.08 | 0.998 | 19.07 | 21.8188 | 14.41 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.39 | 40739.523 | 1081526 | 3567.718 | 0.114 | 0.998 | 30.52 | 31.2723 | 2.46 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.38 | 263.402 | 4646 | 3638.662 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.38 | 60867.402 | 1597439 | 3390.249 | 0.18 | 0.998 | 48.83 | 49.092 | 0.54 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.38 | 93646.594 | 2432739 | 3215.925 | 0.291 | 0.998 | 78.13 | 79.1339 | 1.28 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.39 | 127253.023 | 3214305 | 2816.201 | 0.452 | 0.998 | 125 | 121.5041 | -2.8 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.39 | 161479.188 | 4011924 | 2776.837 | 0.582 | 0.998 | 156.25 | 155.0028 | -0.8 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.39 | 249333.906 | 5930930 | 2500.137 | 0.997 | 0.998 | 250 | 258.609 | 3.44 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.38 | 19.47 | 536 |  |  | 0.998 |  |  |  | MM | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.39 | 1191.374 | 29767 | 3600.895 | 0.003 | 0.998 | 0.63 | 0.61554 | -2.3 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.39 | 4194.164 | 112327 | 3990.492 | 0.011 | 0.998 | 2.5 | 2.6232 | 4.93 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.38 | 9285.781 | 249606 | 3898.205 | 0.024 | 0.998 | 6.25 | 6.3277 | 1.24 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.38 | 33847.813 | 890702 | 3650.252 | 0.093 | 0.998 | 25 | 25.3803 | 1.52 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.39 | 3.406 | 132 |  |  | 0.998 |  |  |  | bb | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.38 | 527.062 | 11033 | 3625.022 | 0.001 | 0.998 | 0.17 | 0.09814 | -42.27 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.39 | 827.368 | 19812 | 3695.023 | 0.002 | 0.998 | 0.28 | 0.31721 | 13.29 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1271.47 | 31590 | 3889.072 | 0.003 | 0.998 | 0.44 | 0.6046 | 37.41 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1399.694 | 34316 | 3767.749 | 0.004 | 0.998 | 0.71 | 0.72889 | 2.66 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2004.768 | 50060 | 3848.652 | 0.005 | 0.998 | 1.14 | 1.1456 | 0.49 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.38 | 2839.042 | 73645 | 3778.806 | 0.008 | 0.998 | 1.82 | 1.7879 | -1.76 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.38 | 278.284 | 4387 | 3744.692 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.38 | 567.657 | 13006 | 3855.243 | 0.001 | 0.998 | 0.17 | 0.1033 | -39.24 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.38 | 839.677 | 20095 | 4039.755 | 0.002 | 0.998 | 0.28 | 0.27241 | -2.71 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1328.937 | 33861 | 3796.014 | 0.004 | 0.998 | 0.44 | 0.66918 | 52.09 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1363.15 | 33780 | 3839.881 | 0.004 | 0.998 | 0.71 | 0.68288 | -3.82 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.38 | 1926.949 | 49041 | 3659.798 | 0.005 | 0.998 | 1.14 | 1.1612 | 1.86 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.38 | 2903.629 | 76605 | 3857.974 | 0.008 | 0.998 | 1.82 | 1.7916 | -1.56 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.38 | 0.68 | 36 |  |  | 0.998 |  |  |  | MM | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.38 | 295.277 | 4998 | 3968.314 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.39 | 297.621 | 5067 | 4331.905 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.39 | 316.359 | 6489 | 3953.513 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.39 | 337.693 | 6453 | 4149.038 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.39 | 340.582 | 5610 | 4039.311 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.38 | 330.39 | 6190 | 3948.045 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.39 | 316.148 | 5605 | 3766.95 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.39 | 2757.122 | 81393 | 4110.244 | 0.007 | 0.998 |  | 1.5635 |  | MM | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.38 | 2786.65 | 83571 | 4134.708 | 0.007 | 0.998 |  | 1.5723 |  | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.38 | 2731.141 | 80144 | 3806.312 | 0.007 | 0.998 |  | 1.6938 |  | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.39 | 123.017 | 2036 | 3714.114 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.31 | 91.615 | 1338 | 3808.625 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.38 | 135.134 | 2066 | 3752.886 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.38 | 0.677 | 36 |  |  | 0.998 |  |  |  | MM | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.39 | 148.575 | 2751 | 3651.146 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.39 | 168.705 | 2828 | 4014.833 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.39 | 161.723 | 2740 | 3989.71 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.39 | 146618.016 | 3627747 | 2913.662 | 0.503 | 0.998 |  | 134.8422 |  | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.39 | 150890.422 | 3728096 | 2855.056 | 0.529 | 0.998 |  | 141.3775 |  | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.38 | 155619.047 | 3811427 | 2979.928 | 0.522 | 0.998 |  | 139.7575 |  | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.38 | 9.846 | 305 |  |  | 0.998 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.38 | 125.056 | 2465 | 4423.004 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.38 | 138.453 | 2327 | 3863.019 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.38 | 170 | 3330 | 3884.237 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.38 | 153.994 | 2656 | 3750.135 | 0 | 0.998 |  |  |  | MMI | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.39 | 195.995 | 3559 | 3862.314 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.39 | 304.253 | 6793 | 3867.536 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.39 | 357.354 | 6692 | 4093.351 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.38 | 155005.031 | 3792690 | 3048.485 | 0.508 | 0.998 |  | 136.2024 |  | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.38 | 153813.781 | 3738978 | 2956.842 | 0.52 | 0.998 |  | 139.2339 |  | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.39 | 154819.5 | 3822974 | 3010.562 | 0.514 | 0.998 |  | 137.699 |  | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.39 | 302.362 | 6403 | 4037.25 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.3 | 1.455 | 61 |  |  | 0.998 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.39 | 586.011 | 12529 | 3890.618 | 0.002 | 0.998 | 0.17 | 0.11273 | -33.69 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.39 | 827.47 | 18625 | 4132.465 | 0.002 | 0.998 | 0.28 | 0.25116 | -10.3 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1365.324 | 33701 | 4043.019 | 0.003 | 0.998 | 0.44 | 0.63463 | 44.23 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1454.075 | 35802 | 3970.605 | 0.004 | 0.998 | 0.71 | 0.71415 | 0.58 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.38 | 2100.144 | 53892 | 3911.216 | 0.005 | 0.998 | 1.14 | 1.1903 | 4.41 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.39 | 2960.601 | 78339 | 3887.929 | 0.008 | 0.998 | 1.82 | 1.8163 | -0.2 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.4 | 0.946 | 45 |  |  | 0.998 |  |  |  | MM | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.38 | 4287.631 | 110532 | 4020.333 | 0.011 | 0.998 | 2.91 | 2.6663 | -8.38 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.39 | 6823.404 | 186699 | 3883.1 | 0.018 | 0.998 | 4.66 | 4.5896 | -1.51 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.38 | 10241.668 | 273649 | 3891.241 | 0.026 | 0.998 | 7.45 | 7.0223 | -5.74 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.38 | 17748.586 | 475577 | 3876.859 | 0.046 | 0.998 | 11.92 | 12.4223 | 4.21 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.39 | 30783.832 | 821268 | 3949.588 | 0.078 | 0.998 | 19.07 | 21.3096 | 11.74 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.39 | 44354.648 | 1184102 | 3972.074 | 0.112 | 0.998 | 30.52 | 30.5806 | 0.2 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.39 | 342 | 6891 | 4145.597 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.39 | 66189.313 | 1736158 | 3703.594 | 0.179 | 0.998 | 48.83 | 48.8694 | 0.08 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.39 | 97748.891 | 2529847 | 3471.925 | 0.282 | 0.998 | 78.13 | 76.5559 | -2.01 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.39 | 139309.797 | 3475948 | 3256.591 | 0.428 | 0.998 | 125 | 115.2151 | -7.83 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.39 | 176249.031 | 4387810 | 2948.698 | 0.598 | 0.998 | 156.25 | 159.1454 | 1.85 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.38 | 268494.438 | 6311834 | 2836.917 | 0.946 | 0.998 | 250 | 246.2291 | -1.51 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.39 | 35.317 | 1175 |  |  | 0.998 |  |  |  | MM | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.38 | 1205.266 | 28393 | 3799.566 | 0.003 | 0.998 | 0.63 | 0.57748 | -8.34 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.38 | 4807.84 | 127955 | 4658.844 | 0.01 | 0.998 | 2.5 | 2.5701 | 2.81 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.38 | 9814.89 | 268958 | 4262.813 | 0.023 | 0.998 | 6.25 | 6.1063 | -2.3 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.38 | 36575.203 | 979482 | 3777.862 | 0.097 | 0.998 | 25 | 26.5039 | 6.02 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.38 | 3.556 | 134 |  |  | 0.998 |  |  |  | bb | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.38 | 319.602 | 5525 | 4047.1 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.39 | 627.342 | 14206 | 4065.303 | 0.002 | 0.998 | 0.17 | 0.12304 | -27.63 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.39 | 954.839 | 21923 | 4359.24 | 0.002 | 0.998 | 0.28 | 0.30361 | 8.43 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1292.338 | 29650 | 4362.105 | 0.003 | 0.998 | 0.44 | 0.51905 | 17.97 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1482.216 | 36689 | 4099.792 | 0.004 | 0.998 | 0.71 | 0.70111 | -1.25 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2028.367 | 51723 | 4000.38 | 0.005 | 0.998 | 1.14 | 1.1069 | -2.9 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3126.272 | 82932 | 4255.308 | 0.007 | 0.998 | 1.82 | 1.7416 | -4.31 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.39 | 103.959 | 2131 | 1428.834 | 0.001 | 0.998 |  |  |  | MMI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.38 | 605.618 | 12945 | 3958.613 | 0.002 | 0.998 | 0.17 | 0.11933 | -29.81 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.38 | 864.866 | 20423 | 4343.202 | 0.002 | 0.998 | 0.28 | 0.24807 | -11.4 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1350.651 | 33993 | 4184.41 | 0.003 | 0.998 | 0.44 | 0.59302 | 34.78 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1560.742 | 39680 | 4219.208 | 0.004 | 0.998 | 0.71 | 0.72448 | 2.04 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2153.438 | 55100 | 4015.649 | 0.005 | 0.998 | 1.14 | 1.1884 | 4.25 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3208.746 | 87627 | 3982.236 | 0.008 | 0.998 | 1.82 | 1.9397 | 6.58 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.39 | 2.669 | 119 |  |  | 0.998 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.38 | 588.431 | 13426 | 3983.047 | 0.001 | 0.998 | 0.17 | 0.10467 | -38.43 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.39 | 948.439 | 22130 | 4469.724 | 0.002 | 0.998 | 0.28 | 0.28451 | 1.61 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1399.109 | 35520 | 4314.678 | 0.003 | 0.998 | 0.44 | 0.59716 | 35.72 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1508.809 | 36809 | 4151.266 | 0.004 | 0.998 | 0.71 | 0.70648 | -0.5 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2146.172 | 53294 | 4116.965 | 0.005 | 0.998 | 1.14 | 1.1467 | 0.59 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3324.506 | 86797 | 4224.377 | 0.008 | 0.998 | 1.82 | 1.8874 | 3.7 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.4 | 1.987 | 98 |  |  | 0.998 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.998 |  |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.998 |  |  |  |  | 2020-11-24
Compound 6: n-Butylparaben
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.52 | 1.753 | 72 |  |  | 0.977 |  |  |  | bb | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.48 | 5.68 | 153 |  |  | 0.977 |  |  |  | bb | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.46 | 1.126 | 41 |  |  | 0.977 |  |  |  | MM | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.49 | 12.328 | 184 |  |  | 0.977 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.49 | 86.892 | 2244 | 18857.293 | 0 | 0.977 |  |  |  | bbI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.47 | 2.664 | 90 |  |  | 0.977 |  |  |  | bb | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.49 | 275.932 | 7944 | 19211.244 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1093.283 | 29519 | 19452.498 | 0.001 | 0.977 | 0.28 | 0.5328 | 90.29 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2038.82 | 54724 | 19150.906 | 0.001 | 0.977 | 0.44 | 1.572 | 257.28 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1336.353 | 36085 | 19353 | 0.001 | 0.977 | 0.71 | 0.79849 | 12.46 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1132.776 | 30425 | 18654.168 | 0.001 | 0.977 | 1.14 | 0.62631 | -45.06 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1858.92 | 49326 | 18554.959 | 0.001 | 0.977 | 1.82 | 1.4423 | -20.76 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.49 | 11.044 | 290 |  |  | 0.977 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.49 | 3011.271 | 82197 | 19629.662 | 0.002 | 0.977 | 2.91 | 2.5427 | -12.62 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.49 | 5437.346 | 144185 | 18885.018 | 0.003 | 0.977 | 4.66 | 5.324 | 14.25 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.49 | 6844.65 | 186400 | 18374.332 | 0.004 | 0.977 | 7.45 | 7.073 | -5.06 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.49 | 11706.379 | 317014 | 19000.076 | 0.006 | 0.977 | 11.92 | 12.1094 | 1.59 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.49 | 23140.768 | 624188 | 20171.609 | 0.011 | 0.977 | 19.07 | 23.0868 | 21.06 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.49 | 30056.643 | 810873 | 19826.656 | 0.015 | 0.977 | 30.52 | 30.7079 | 0.62 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.49 | 90.725 | 2272 | 19474.645 | 0 | 0.977 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.49 | 46749.152 | 1276742 | 19678.5 | 0.024 | 0.977 | 48.83 | 48.4689 | -0.74 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.49 | 73348.117 | 1995609 | 18895.691 | 0.039 | 0.977 | 78.13 | 79.568 | 1.84 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.49 | 107727.188 | 2899075 | 18753.004 | 0.057 | 0.977 | 125 | 118.0027 | -5.6 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.49 | 141221.516 | 3782837 | 19034.314 | 0.074 | 0.977 | 156.25 | 152.5302 | -2.38 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.49 | 233946.781 | 6256998 | 19099.988 | 0.122 | 0.977 | 250 | 251.9417 | 0.78 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.47 | 50.034 | 746 |  |  | 0.977 |  |  |  | bb | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.49 | 869.421 | 23752 | 19263.697 | 0 | 0.977 | 0.63 | 0.30389 | -51.76 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.49 | 2826.188 | 75980 | 20758.027 | 0.001 | 0.977 | 2.5 | 2.1859 | -12.56 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.49 | 6402.18 | 171699 | 20168.148 | 0.003 | 0.977 | 6.25 | 5.9344 | -5.05 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.49 | 25493.525 | 686891 | 19404.105 | 0.013 | 0.977 | 25 | 26.5305 | 6.12 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.46 | 9.642 | 158 |  |  | 0.977 |  |  |  | bb | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.49 | 344.988 | 9416 | 19356.43 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1084.972 | 29756 | 20514.73 | 0.001 | 0.977 | 0.28 | 0.46425 | 65.8 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2111.016 | 57924 | 20179.088 | 0.001 | 0.977 | 0.44 | 1.5338 | 248.6 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1404.504 | 38192 | 20475.318 | 0.001 | 0.977 | 0.71 | 0.78905 | 11.13 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1214.246 | 32244 | 19498.076 | 0.001 | 0.977 | 1.14 | 0.65837 | -42.25 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1892.037 | 49662 | 19607.979 | 0.001 | 0.977 | 1.82 | 1.3659 | -24.95 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.49 | 120.839 | 3122 | 19650.477 | 0 | 0.977 |  |  |  | bbI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.49 | 255.55 | 6633 | 19171.268 | 0 | 0.977 | 0.17 |  |  | bbI | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1221.837 | 33205 | 21051.086 | 0.001 | 0.977 | 0.28 | 0.57082 | 103.87 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2038.302 | 54126 | 20352.199 | 0.001 | 0.977 | 0.44 | 1.4416 | 227.63 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1503.632 | 39949 | 19914.854 | 0.001 | 0.977 | 0.71 | 0.93189 | 31.25 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1165.913 | 32669 | 19468.383 | 0.001 | 0.977 | 1.14 | 0.60899 | -46.58 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1999.524 | 53469 | 20120.816 | 0.001 | 0.977 | 1.82 | 1.4255 | -21.67 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.6 | 4.744 | 112 |  |  | 0.977 |  |  |  | bb | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.49 | 5562.143 | 151564 | 21697.566 | 0.003 | 0.977 |  | 4.6712 |  | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.49 | 6113.616 | 167782 | 21829.346 | 0.003 | 0.977 |  | 5.1616 |  | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.49 | 5989.404 | 162539 | 20831.814 | 0.003 | 0.977 |  | 5.3156 |  | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.49 | 6199.086 | 166477 | 21105.879 | 0.003 | 0.977 |  | 5.4438 |  | MM | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.49 | 6240.639 | 170442 | 20513.939 | 0.003 | 0.977 |  | 5.6609 |  | MM | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.49 | 6418.389 | 173326 | 20742.979 | 0.003 | 0.977 |  | 5.7687 |  | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.49 | 121.111 | 3289 | 20304.375 | 0 | 0.977 |  |  |  | MMI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.49 | 5727.25 | 153109 | 21071.273 | 0.003 | 0.977 |  | 4.9908 |  | MM | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.49 | 6183.206 | 167058 | 20743.402 | 0.003 | 0.977 |  | 5.5341 |  | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.49 | 6032.018 | 162409 | 20386.736 | 0.003 | 0.977 |  | 5.4886 |  | MM | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.49 | 107377.172 | 2903668 | 19113.52 | 0.056 | 0.977 |  | 115.3903 |  | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.49 | 115516.563 | 3114122 | 19075.039 | 0.061 | 0.977 |  | 124.424 |  | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.49 | 115451.219 | 3119983 | 19599.061 | 0.059 | 0.977 |  | 121.0162 |  | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.47 | 11.253 | 324 |  |  | 0.977 |  |  |  | MM | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.49 | 120116.875 | 3285115 | 18590.02 | 0.065 | 0.977 |  | 132.7843 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.49 | 130066.008 | 3487232 | 20187.695 | 0.064 | 0.977 |  | 132.4023 |  | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.49 | 130411.273 | 3512177 | 19809.779 | 0.066 | 0.977 |  | 135.2957 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.49 | 114309.352 | 3075295 | 19047.59 | 0.06 | 0.977 |  | 123.297 |  | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.49 | 116265.188 | 3146169 | 19612.295 | 0.059 | 0.977 |  | 121.7901 |  | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.49 | 121336.031 | 3245035 | 19736.887 | 0.061 | 0.977 |  | 126.3165 |  | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.47 | 11.231 | 317 |  |  | 0.977 |  |  |  | MM | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.49 | 104149.32 | 2813923 | 22790.889 | 0.046 | 0.977 |  | 93.7681 |  | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.49 | 100502.93 | 2698115 | 20082.863 | 0.05 | 0.977 |  | 102.7359 |  | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.49 | 100130.891 | 2728622 | 20098.73 | 0.05 | 0.977 |  | 102.2725 |  | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.49 | 101750.938 | 2724955 | 19749.709 | 0.052 | 0.977 |  | 105.7811 |  | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.49 | 104112.445 | 2802601 | 19920.66 | 0.052 | 0.977 |  | 107.3146 |  | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.49 | 101660.32 | 2720809 | 19372.84 | 0.052 | 0.977 |  | 107.7522 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.49 | 140.072 | 3686 | 20612.568 | 0 | 0.977 |  |  |  | MMI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.49 | 99647.773 | 2684259 | 20799.916 | 0.048 | 0.977 |  | 98.3282 |  | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.49 | 98506.859 | 2675978 | 20129.732 | 0.049 | 0.977 |  | 100.4497 |  | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.49 | 98968.75 | 2698679 | 20311.908 | 0.049 | 0.977 |  | 100.0133 |  | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.5 | 111.485 | 2843 | 20241.559 | 0 | 0.977 |  |  |  | MMI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.47 | 7.267 | 249 |  |  | 0.977 |  |  |  | MM | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.49 | 299.582 | 7591 | 20027.883 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1179.964 | 32207 | 21530.92 | 0.001 | 0.977 | 0.28 | 0.50386 | 79.95 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2052.207 | 59116 | 21003.221 | 0.001 | 0.977 | 0.44 | 1.3911 | 216.15 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1596.819 | 42966 | 21356.199 | 0.001 | 0.977 | 0.71 | 0.91675 | 29.12 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1243.231 | 34288 | 20342.879 | 0.001 | 0.977 | 1.14 | 0.63435 | -44.36 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1923.861 | 51393 | 20303.764 | 0.001 | 0.977 | 1.82 | 1.33 | -26.93 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.46 | 9.078 | 178 |  |  | 0.977 |  |  |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.49 | 3185.441 | 87389 | 20294.021 | 0.002 | 0.977 | 2.91 | 2.6163 | -10.09 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.49 | 5724.029 | 158430 | 19831.742 | 0.003 | 0.977 | 4.66 | 5.3387 | 14.56 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.49 | 7168.389 | 194432 | 20010.529 | 0.004 | 0.977 | 7.45 | 6.7777 | -9.02 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.49 | 12422.017 | 335605 | 20213.203 | 0.006 | 0.977 | 11.92 | 12.0769 | 1.32 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.49 | 25001.709 | 672642 | 21640.061 | 0.012 | 0.977 | 19.07 | 23.2552 | 21.95 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.49 | 31641.816 | 849207 | 21196.979 | 0.015 | 0.977 | 30.52 | 30.2281 | -0.96 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.49 | 125.845 | 2872 | 21380.4 | 0 | 0.977 |  |  |  | bbI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.49 | 50444.727 | 1365900 | 20941.443 | 0.024 | 0.977 | 48.83 | 49.1547 | 0.66 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.49 | 77764 | 2124020 | 20097.592 | 0.039 | 0.977 | 78.13 | 79.3116 | 1.51 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.49 | 116241.914 | 3148214 | 20623.277 | 0.056 | 0.977 | 125 | 115.7735 | -7.38 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.49 | 150393.859 | 4087704 | 19677.9 | 0.076 | 0.977 | 156.25 | 157.1352 | 0.57 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.49 | 253237.719 | 6808499 | 20306.045 | 0.125 | 0.977 | 250 | 256.5171 | 2.61 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.5 | 1.847 | 91 |  |  | 0.977 |  |  |  | MM | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.49 | 885.964 | 23830 | 20563.285 | 0 | 0.977 | 0.63 | 0.26154 | -58.49 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.49 | 3360.3 | 91191 | 24810.557 | 0.001 | 0.977 | 2.5 | 2.1712 | -13.15 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.49 | 6762.778 | 184250 | 21870.688 | 0.003 | 0.977 | 6.25 | 5.7644 | -7.77 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.49 | 27081.482 | 739194 | 21461.814 | 0.013 | 0.977 | 25 | 25.4563 | 1.83 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.52 | 2.993 | 128 |  |  | 0.977 |  |  |  | MM | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.49 | 113.136 | 3323 | 21765.705 | 0 | 0.977 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.49 | 324.78 | 8199 | 21216.219 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1213.55 | 32660 | 21769.814 | 0.001 | 0.977 | 0.28 | 0.52333 | 86.9 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2281.747 | 61964 | 21732.258 | 0.001 | 0.977 | 0.44 | 1.5417 | 250.38 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1593.851 | 44044 | 21173.01 | 0.001 | 0.977 | 0.71 | 0.92723 | 30.6 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1309.173 | 35064 | 21096.451 | 0.001 | 0.977 | 1.14 | 0.65385 | -42.65 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2031.443 | 55841 | 20774.008 | 0.001 | 0.977 | 1.82 | 1.3927 | -23.48 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.49 | 55.68 | 1125 | 7133.3 | 0 | 0.977 |  |  |  | bbI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.49 | 361.179 | 9824 | 20622.34 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1313.061 | 35458 | 22279.238 | 0.001 | 0.977 | 0.28 | 0.58933 | 110.47 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2139.417 | 57603 | 22107.174 | 0.001 | 0.977 | 0.44 | 1.3718 | 211.76 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1595.844 | 43525 | 22507.645 | 0.001 | 0.977 | 0.71 | 0.83676 | 17.85 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1263.111 | 35567 | 21428.309 | 0.001 | 0.977 | 1.14 | 0.58952 | -48.29 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2034.314 | 53494 | 21171.861 | 0.001 | 0.977 | 1.82 | 1.3575 | -25.41 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.5 | 8.466 | 186 |  |  | 0.977 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.49 | 358.748 | 10079 | 20838.236 | 0 | 0.977 | 0.17 |  |  | MMI | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1252.409 | 34936 | 22277.256 | 0.001 | 0.977 | 0.28 | 0.53314 | 90.41 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2225.525 | 59572 | 22667.871 | 0.001 | 0.977 | 0.44 | 1.4008 | 218.36 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1545.231 | 41500 | 21603.74 | 0.001 | 0.977 | 0.71 | 0.84966 | 19.67 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1230.308 | 33452 | 21934.615 | 0.001 | 0.977 | 1.14 | 0.53046 | -53.47 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2038.945 | 56389 | 21390.625 | 0.001 | 0.977 | 1.82 | 1.3417 | -26.28 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.47 | 1.671 | 69 |  |  | 0.977 |  |  |  | MM | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.977 |  |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.43 | 3.952 | 157 |  |  | 0.977 |  |  |  | MM | 2020-11-24
Compound 7: PFPE-6
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.67 | 7.649 | 267 |  |  | 0.997 |  |  |  | MM | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.64 | 20.492 | 490 |  |  | 0.997 |  |  |  | MM | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.71 | 14.427 | 319 |  |  | 0.997 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.65 | 15.691 | 361 |  |  | 0.997 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.66 | 40.602 | 772 | 31803.217 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.61 | 5.261 | 129 |  |  | 0.997 |  |  |  | MM | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.65 | 921.642 | 27945 | 32447.529 | 0 | 0.997 | 0.17 | 0.10827 | -36.31 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1892.717 | 55740 | 32267.967 | 0.001 | 0.997 | 0.28 | 0.28866 | 3.09 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3094.552 | 90538 | 33209.082 | 0.001 | 0.997 | 0.44 | 0.49462 | 12.41 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4415.33 | 131556 | 33202.188 | 0.001 | 0.997 | 0.71 | 0.73212 | 3.12 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.65 | 6632.528 | 197162 | 32022.859 | 0.002 | 0.997 | 1.14 | 1.1748 | 3.05 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.65 | 10836.208 | 318304 | 31705.016 | 0.003 | 0.997 | 1.82 | 1.9797 | 8.77 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.67 | 6.266 | 160 |  |  | 0.997 |  |  |  | db | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.65 | 15626.126 | 461167 | 34305.039 | 0.005 | 0.997 | 2.91 | 2.6603 | -8.58 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.65 | 25204.344 | 744851 | 33209.145 | 0.008 | 0.997 | 4.66 | 4.4804 | -3.85 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.65 | 39542.523 | 1165600 | 30390.211 | 0.013 | 0.997 | 7.45 | 7.7468 | 3.98 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.65 | 70232.273 | 2073122 | 32738.229 | 0.021 | 0.997 | 11.92 | 12.8693 | 7.96 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.65 | 125018.797 | 3684048 | 34820.98 | 0.036 | 0.997 | 19.07 | 21.7474 | 14.04 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.65 | 181213.188 | 5323364 | 35522.934 | 0.051 | 0.997 | 30.52 | 31.1834 | 2.17 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.65 | 63.036 | 1855 | 32830.207 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.65 | 272351.188 | 7932075 | 34067.25 | 0.08 | 0.997 | 48.83 | 49.7182 | 1.82 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.65 | 416492.813 | 12135428 | 34067.969 | 0.122 | 0.997 | 78.13 | 78.0479 | -0.11 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.65 | 584937.688 | 16983996 | 32972.277 | 0.177 | 0.997 | 125 | 117.5629 | -5.95 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.65 | 749184.313 | 21513114 | 33862.289 | 0.221 | 0.997 | 156.25 | 151.556 | -3 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.65 | 1160101.125 | 32674774 | 34733.418 | 0.334 | 0.997 | 250 | 254.8213 | 1.93 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.65 | 139.596 | 3672 |  |  | 0.997 |  |  |  | bb | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.65 | 3654.858 | 107228 | 32111.395 | 0.001 | 0.997 | 0.63 | 0.61774 | -1.95 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.65 | 17056.02 | 495244 | 35946.789 | 0.005 | 0.997 | 2.5 | 2.774 | 10.96 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.65 | 39337.004 | 1152457 | 34698.945 | 0.011 | 0.997 | 6.25 | 6.7358 | 7.77 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.65 | 150432.547 | 4406268 | 35005.285 | 0.043 | 0.997 | 25 | 26.1433 | 4.57 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.65 | 23.374 | 439 |  |  | 0.997 |  |  |  | MM | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1010.817 | 29652 | 32869.164 | 0 | 0.997 | 0.17 | 0.12227 | -28.08 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1967.408 | 59010 | 33731.746 | 0.001 | 0.997 | 0.28 | 0.28669 | 2.39 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3335.631 | 98158 | 35506.813 | 0.001 | 0.997 | 0.44 | 0.49915 | 13.44 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4474.458 | 129596 | 35507.883 | 0.001 | 0.997 | 0.71 | 0.69052 | -2.74 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.65 | 6970.117 | 207434 | 33292.348 | 0.002 | 0.997 | 1.14 | 1.1882 | 4.22 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11246.253 | 331176 | 33909.344 | 0.003 | 0.997 | 1.82 | 1.9191 | 5.45 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.65 | 65.394 | 1925 | 33975.277 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1031.095 | 29137 | 34714.773 | 0 | 0.997 | 0.17 | 0.116 | -31.76 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2055.06 | 59931 | 36595.254 | 0.001 | 0.997 | 0.28 | 0.27375 | -2.23 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3398.498 | 97760 | 34187.922 | 0.001 | 0.997 | 0.44 | 0.53175 | 20.85 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4496.663 | 133073 | 34225.648 | 0.001 | 0.997 | 0.71 | 0.72257 | 1.77 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7280.397 | 213942 | 33805.305 | 0.002 | 0.997 | 1.14 | 1.224 | 7.37 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11218.586 | 341307 | 33640.234 | 0.003 | 0.997 | 1.82 | 1.9301 | 6.05 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.64 | 2.933 | 160 |  |  | 0.997 |  |  |  | MM | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.65 | 20.206 | 751 | 35815.477 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.65 | 48.589 | 954 | 38437.145 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.65 | 36.054 | 960 | 36556.488 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.65 | 3861.719 | 116503 | 37541.629 | 0.001 | 0.997 |  | 0.55238 |  | MM | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.65 | 3670.254 | 108089 | 35548.031 | 0.001 | 0.997 |  | 0.55467 |  | MM | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.65 | 4291.371 | 125681 | 35730.238 | 0.001 | 0.997 |  | 0.65526 |  | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.65 | 59.23 | 1431 | 35277.617 | 0 | 0.997 |  |  |  | MMI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.66 | 46.087 | 738 | 37072.477 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.66 | 22.878 | 332 | 38052.293 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.65 | 35.243 | 1118 | 34841.289 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.65 | 37.985 | 1123 | 33680.504 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.66 | 20.323 | 481 | 33313.039 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.65 | 75.389 | 1965 | 35218.301 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.65 | 7.274 | 181 |  |  | 0.997 |  |  |  | bb | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.65 | 637827.313 | 18381228 | 31820.521 | 0.2 | 0.997 |  | 135.1066 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.65 | 689029.875 | 19863128 | 35913.09 | 0.192 | 0.997 |  | 128.4911 |  | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.65 | 696401.188 | 19752984 | 34425.738 | 0.202 | 0.997 |  | 136.5417 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.65 | 272.514 | 7109 | 33026.402 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.65 | 84.635 | 2296 | 33342.656 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.65 | 103.472 | 2750 | 35495.684 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.67 | 15.721 | 454 |  |  | 0.997 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.65 | 70.968 | 1289 | 41065.688 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.65 | 99.183 | 2854 | 36334.883 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.65 | 40.615 | 1160 | 33962.336 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.65 | 662656.688 | 18908058 | 33873.305 | 0.196 | 0.997 |  | 131.3825 |  | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.65 | 686166 | 19567960 | 34683.414 | 0.198 | 0.997 |  | 133.0863 |  | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.65 | 685958.188 | 19722810 | 34392.777 | 0.199 | 0.997 |  | 134.3334 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.65 | 212.124 | 6384 | 35834.434 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.65 | 243.215 | 6252 | 34104.559 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.65 | 56.921 | 1696 | 33587.531 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.65 | 104.988 | 2933 | 34354.273 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.65 | 54.949 | 1671 | 36040.789 | 0 | 0.997 |  |  |  | bdI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.75 | 29.395 | 394 |  |  | 0.997 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1015.728 | 29053 | 33756.629 | 0 | 0.997 | 0.17 | 0.11832 | -30.4 | bb | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1956.784 | 56560 | 36770.707 | 0.001 | 0.997 | 0.28 | 0.25622 | -8.49 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3448.573 | 102044 | 36297.711 | 0.001 | 0.997 | 0.44 | 0.5055 | 14.89 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4881.735 | 145085 | 36058.262 | 0.001 | 0.997 | 0.71 | 0.74645 | 5.13 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7141.836 | 210311 | 33890.82 | 0.002 | 0.997 | 1.14 | 1.1963 | 4.94 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11494.638 | 336124 | 34649.969 | 0.003 | 0.997 | 1.82 | 1.9196 | 5.47 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.68 | 9.299 | 203 |  |  | 0.997 |  |  |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.65 | 17015.816 | 503700 | 36403.238 | 0.005 | 0.997 | 2.91 | 2.7317 | -6.13 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.65 | 26487.211 | 777943 | 35900.168 | 0.007 | 0.997 | 4.66 | 4.3533 | -6.58 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.65 | 43067.457 | 1273534 | 33082.625 | 0.013 | 0.997 | 7.45 | 7.7507 | 4.04 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.65 | 75099.82 | 2225947 | 34432.25 | 0.022 | 0.997 | 11.92 | 13.0877 | 9.8 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.65 | 133696.859 | 3914060 | 36870.191 | 0.036 | 0.997 | 19.07 | 21.9692 | 15.2 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.65 | 189827.453 | 5621227 | 37336.027 | 0.051 | 0.997 | 30.52 | 31.0763 | 1.82 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.65 | 122.007 | 3292 | 37023.516 | 0 | 0.997 |  |  |  | bbI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.65 | 289306.313 | 8436817 | 36169.313 | 0.08 | 0.997 | 48.83 | 49.7453 | 1.87 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.65 | 433477.563 | 12554088 | 34693.082 | 0.125 | 0.997 | 78.13 | 79.9052 | 2.27 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.65 | 621589.5 | 17811714 | 36845.703 | 0.169 | 0.997 | 125 | 111.1073 | -11.11 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.65 | 799396.875 | 22652376 | 36388.051 | 0.22 | 0.997 | 156.25 | 150.3022 | -3.81 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.65 | 1213751.25 | 33633596 | 35195.254 | 0.345 | 0.997 | 250 | 266.5475 | 6.62 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.65 | 190.765 | 5700 |  |  | 0.997 |  |  |  | bb | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.65 | 3979.43 | 116367 | 35430.559 | 0.001 | 0.997 | 0.63 | 0.60878 | -3.37 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.65 | 19007.379 | 562991 | 39598.043 | 0.005 | 0.997 | 2.5 | 2.8071 | 12.28 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.65 | 42555.238 | 1260814 | 38674.035 | 0.011 | 0.997 | 6.25 | 6.5349 | 4.56 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.65 | 156770.422 | 4610806 | 35777.035 | 0.044 | 0.997 | 25 | 26.6706 | 6.68 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.66 | 48.535 | 646 |  |  | 0.997 |  |  |  | bb | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.65 | 66.616 | 1643 | 37009.66 | 0 | 0.997 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1045.793 | 32103 | 36119.102 | 0 | 0.997 | 0.17 | 0.11154 | -34.39 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2080.034 | 60858 | 37791.871 | 0.001 | 0.997 | 0.28 | 0.26709 | -4.61 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3597.861 | 106337 | 37101.148 | 0.001 | 0.997 | 0.44 | 0.51724 | 17.55 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4702.09 | 135939 | 36060.398 | 0.001 | 0.997 | 0.71 | 0.71667 | 0.94 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7578.791 | 222935 | 35058.344 | 0.002 | 0.997 | 1.14 | 1.2289 | 7.8 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.65 | 12025.298 | 351434 | 35763.523 | 0.003 | 0.997 | 1.82 | 1.9466 | 6.95 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.65 | 22.222 | 484 | 12664.324 | 0 | 0.997 |  |  |  | bbI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.65 | 983.628 | 29409 | 35462.613 | 0 | 0.997 | 0.17 | 0.10429 | -38.65 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2147.191 | 62832 | 37730.301 | 0.001 | 0.997 | 0.28 | 0.27824 | -0.63 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3533.678 | 107018 | 37717.949 | 0.001 | 0.997 | 0.44 | 0.49762 | 13.1 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4677.835 | 139603 | 36561.789 | 0.001 | 0.997 | 0.71 | 0.70204 | -1.12 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7474.03 | 220792 | 34807.703 | 0.002 | 0.997 | 1.14 | 1.2202 | 7.03 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11785.349 | 348689 | 35477.816 | 0.003 | 0.997 | 1.82 | 1.9223 | 5.62 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [5 empty] | 0.997 |  |  |  |  | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1105.723 | 32191 | 35970.531 | 0 | 0.997 | 0.17 | 0.12219 | -28.12 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2290.525 | 67744 | 37755.715 | 0.001 | 0.997 | 0.28 | 0.30065 | 7.38 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3538.376 | 103766 | 36696.27 | 0.001 | 0.997 | 0.44 | 0.51395 | 16.81 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4899.534 | 144733 | 36844.078 | 0.001 | 0.997 | 0.71 | 0.7321 | 3.11 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7780.425 | 229346 | 35534.512 | 0.002 | 0.997 | 1.14 | 1.2455 | 9.25 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.65 | 12088.213 | 356663 | 36225.75 | 0.003 | 0.997 | 1.82 | 1.9313 | 6.11 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.66 | 13.691 | 211 |  |  | 0.997 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 2.75 | 49.989 | 547 |  |  | 0.997 |  |  |  | bb | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.65 | 2.424 | 74 |  |  | 0.997 |  |  |  | bb | 2020-11-24
Compound 8: PFTrDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 3.24 | 64.051 | 641 |  |  | 0.994 |  |  |  | bb | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 3.24 | 66.092 | 715 |  |  | 0.994 |  |  |  | MM | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 3.2 | 50.479 | 528 |  |  | 0.994 |  |  |  | MM | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 3.3 | 43.052 | 544 |  |  | 0.994 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.25 | 18.166 | 557 | 10173.215 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 3.21 | 1.82 | 169 |  |  | 0.994 |  |  |  | MM | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.25 | 361.462 | 9725 | 9730.571 | 0 | 0.994 | 0.17 | 0.11196 | -34.14 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.25 | 566.751 | 15043 | 11120.537 | 0.001 | 0.994 | 0.28 | 0.18563 | -33.7 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1490.313 | 38142 | 10876.152 | 0.001 | 0.994 | 0.44 | 0.64466 | 46.51 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1231.507 | 33415 | 10360.407 | 0.001 | 0.994 | 0.71 | 0.54778 | -22.85 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2790.359 | 74985 | 10282.522 | 0.003 | 0.994 | 1.14 | 1.3617 | 19.45 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3754.326 | 101025 | 10162.122 | 0.004 | 0.994 | 1.82 | 1.8856 | 3.6 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 3.25 | 43.496 | 761 |  |  | 0.994 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.25 | 4767.658 | 128384 | 10490.147 | 0.005 | 0.994 | 2.91 | 2.3401 | -19.58 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.25 | 10289.728 | 279833 | 9708.633 | 0.011 | 0.994 | 4.66 | 5.5828 | 19.8 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.25 | 14118.135 | 384250 | 9783.861 | 0.014 | 0.994 | 7.45 | 7.6417 | 2.57 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.25 | 21835.18 | 595031 | 11453.623 | 0.019 | 0.994 | 11.92 | 10.1387 | -14.94 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.25 | 44740.945 | 1217712 | 10865.012 | 0.041 | 0.994 | 19.07 | 22.161 | 16.21 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.25 | 60740.434 | 1636760 | 11101.751 | 0.055 | 0.994 | 30.52 | 29.6069 | -2.99 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.25 | 58.275 | 1024 | 11173.664 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.25 | 92595.516 | 2486691 | 10508.111 | 0.088 | 0.994 | 48.83 | 48.2876 | -1.11 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.25 | 133000.719 | 3594597 | 9989.264 | 0.133 | 0.994 | 78.13 | 74.1923 | -5.04 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.25 | 208385.203 | 5623033 | 10623.775 | 0.196 | 0.994 | 125 | 111.9995 | -10.4 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.25 | 278826.625 | 7470329 | 10923.09 | 0.255 | 0.994 | 156.25 | 149.3655 | -4.41 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.25 | 386971.281 | 10065610 | 9730.678 | 0.398 | 0.994 | 250 | 249.1041 | -0.36 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 3.25 | 87.035 | 1492 |  |  | 0.994 |  |  |  | MM | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.25 | 1071.004 | 29043 | 10475.948 | 0.001 | 0.994 | 0.63 | 0.45906 | -27.13 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.25 | 7202.33 | 195015 | 11635.099 | 0.006 | 0.994 | 2.5 | 3.2202 | 28.81 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.25 | 13334.586 | 368465 | 12150.41 | 0.011 | 0.994 | 6.25 | 5.7846 | -7.45 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.25 | 46874.168 | 1286494 | 10343.043 | 0.045 | 0.994 | 25 | 24.4318 | -2.27 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [5 empty] | 0.994 |  |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.25 | 305.135 | 7545 | 11255 | 0 | 0.994 | 0.17 | 0.05845 | -65.62 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.25 | 629.304 | 16713 | 12190.335 | 0.001 | 0.994 | 0.28 | 0.18914 | -32.45 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1584.214 | 41103 | 11718.064 | 0.001 | 0.994 | 0.44 | 0.63489 | 44.29 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1250.094 | 32864 | 11317.867 | 0.001 | 0.994 | 0.71 | 0.5029 | -29.17 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2686.308 | 72115 | 10528.761 | 0.003 | 0.994 | 1.14 | 1.2751 | 11.85 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3925.793 | 105378 | 11242.785 | 0.003 | 0.994 | 1.82 | 1.7773 | -2.34 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.24 | 8.5 | 344 | 12048.315 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.25 | 332.151 | 8955 | 11228.034 | 0 | 0.994 | 0.17 | 0.07162 | -57.87 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.25 | 518.194 | 14739 | 11829.166 | 0 | 0.994 | 0.28 | 0.14746 | -47.33 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1553.966 | 40896 | 11542.465 | 0.001 | 0.994 | 0.44 | 0.63188 | 43.61 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1403.119 | 37945 | 11854.473 | 0.001 | 0.994 | 0.71 | 0.54509 | -23.23 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2752.255 | 76585 | 10775.113 | 0.003 | 0.994 | 1.14 | 1.2766 | 11.98 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3868.403 | 104825 | 10874.418 | 0.004 | 0.994 | 1.82 | 1.8123 | -0.42 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 3.28 | 56.314 | 578 |  |  | 0.994 |  |  |  | bb | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.21 | 84.168 | 702 | 11440.534 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.25 | 133.335 | 1209 | 13256.198 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte |  |  |  | 11719.4 |  | 0.994 |  |  |  |  | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.25 | 33506.41 | 915674 | 11316.626 | 0.03 | 0.994 |  | 15.8488 |  | MM | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.25 | 29692.799 | 819239 | 11121.704 | 0.027 | 0.994 |  | 14.2689 |  | MM | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.25 | 31607.668 | 866983 | 11384.287 | 0.028 | 0.994 |  | 14.8474 |  | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.25 | 30.159 | 885 | 11975.843 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.25 | 54.334 | 622 | 11676.331 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.23 | 42.634 | 734 | 11894.549 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.24 | 28.432 | 575 | 11048.28 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.25 | 159.608 | 2999 | 9862.112 | 0 | 0.994 |  | 0.00019 |  | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.25 | 131.481 | 3142 | 9161.524 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.25 | 134.293 | 3308 | 9049.197 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [5 empty] | 0.994 |  |  |  |  | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.25 | 208258.063 | 5472016 | 8897.753 | 0.234 | 0.994 |  | 135.7308 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.25 | 231957.547 | 6248865 | 10825.349 | 0.214 | 0.994 |  | 123.2462 |  | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.25 | 236618.672 | 6307770 | 9994.433 | 0.237 | 0.994 |  | 137.4478 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.25 | 147.506 | 3836 | 8190.83 | 0 | 0.994 |  | 0.00992 |  | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.25 | 84.121 | 1180 | 9179.411 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.25 | 62.151 | 1381 | 10046.605 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 3.21 | 6.538 | 147 |  |  | 0.994 |  |  |  | db | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.25 | 106.976 | 2961 | 11574.24 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.25 | 149.675 | 2844 | 8454.886 | 0 | 0.994 |  | 0.00829 |  | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.25 | 128.604 | 2937 | 8997.701 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.25 | 214740.047 | 5692729 | 9464.119 | 0.227 | 0.994 |  | 131.1884 |  | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.25 | 216414.156 | 5761683 | 8344.308 | 0.259 | 0.994 |  | 152.0277 |  | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.25 | 209803.172 | 5591388 | 8994.774 | 0.233 | 0.994 |  | 135.2173 |  | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.25 | 81.038 | 2420 | 10773.584 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.24 | 65.747 | 1251 | 9773.682 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.24 | 40.735 | 798 | 9471.696 | 0 | 0.994 |  |  |  | MMI | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.25 | 164.247 | 1293 | 9756.347 | 0 | 0.994 |  | 0.00366 |  | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank |  |  |  | 10459.238 |  | 0.994 |  |  |  |  | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 3.19 | 91.734 | 574 |  |  | 0.994 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.25 | 334.842 | 8353 | 10426.968 | 0 | 0.994 | 0.17 | 0.08512 | -49.93 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.25 | 582.102 | 14804 | 11531.981 | 0.001 | 0.994 | 0.28 | 0.18303 | -34.63 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1453.744 | 41948 | 10908.972 | 0.001 | 0.994 | 0.44 | 0.62458 | 41.95 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1271.26 | 34273 | 10369.82 | 0.001 | 0.994 | 0.71 | 0.56766 | -20.05 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2877.709 | 74907 | 9679.873 | 0.003 | 0.994 | 1.14 | 1.5001 | 31.59 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.25 | 4052.725 | 110355 | 10265.066 | 0.004 | 0.994 | 1.82 | 2.0211 | 11.05 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank |  |  |  | 3.524 |  | 0.994 |  |  |  |  | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.25 | 4873.014 | 136983 | 11516.313 | 0.004 | 0.994 | 2.91 | 2.1725 | -25.34 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.25 | 11089.063 | 302811 | 9848.446 | 0.011 | 0.994 | 4.66 | 5.9377 | 27.42 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.25 | 15005.965 | 402023 | 10084.969 | 0.015 | 0.994 | 7.45 | 7.8836 | 5.82 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.25 | 22863.998 | 623795 | 10773.109 | 0.021 | 0.994 | 11.92 | 11.3048 | -5.16 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.25 | 45678.59 | 1249665 | 11209.78 | 0.041 | 0.994 | 19.07 | 21.9256 | 14.97 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.25 | 63296.691 | 1727377 | 10679.77 | 0.059 | 0.994 | 30.52 | 32.1286 | 5.27 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.31 | 12.677 | 214 | 10166.878 | 0 | 0.994 |  |  |  | bbI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.25 | 92582.195 | 2506564 | 10697.722 | 0.087 | 0.994 | 48.83 | 47.3973 | -2.93 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.25 | 143420.453 | 3838232 | 10158.705 | 0.141 | 0.994 | 78.13 | 78.9092 | 1 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.25 | 218987.391 | 5837530 | 10158.639 | 0.216 | 0.994 | 125 | 124.0568 | -0.75 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.25 | 278772.781 | 7373114 | 9865.15 | 0.283 | 0.994 | 156.25 | 167.3413 | 7.1 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.25 | 401981.25 | 10638402 | 9832.22 | 0.409 | 0.994 | 250 | 257.6472 | 3.06 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 3.25 | 191.709 | 1927 |  |  | 0.994 |  |  |  | bb | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.25 | 1148.222 | 30281 | 10532.542 | 0.001 | 0.994 | 0.63 | 0.49524 | -21.39 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.25 | 5734.268 | 154740 | 6178.086 | 0.009 | 0.994 | 2.5 | 4.8763 | 95.05 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.25 | 12606.851 | 342443 | 10931.824 | 0.012 | 0.994 | 6.25 | 6.084 | -2.66 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.25 | 47936.488 | 1305538 | 9113.911 | 0.053 | 0.994 | 25 | 28.4385 | 13.75 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 3.26 | 89.381 | 786 |  |  | 0.994 |  |  |  | bb | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.24 | 116.384 | 1307 | 9387.197 | 0 | 0.994 |  |  |  | bbI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.25 | 390.738 | 9523 | 9798.542 | 0 | 0.994 | 0.17 | 0.12651 | -25.58 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.25 | 564.932 | 15445 | 10690.253 | 0.001 | 0.994 | 0.28 | 0.19566 | -30.12 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1654.718 | 43523 | 10546.275 | 0.002 | 0.994 | 0.44 | 0.7507 | 70.61 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1374.593 | 37593 | 10312.553 | 0.001 | 0.994 | 0.71 | 0.62475 | -12.01 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2912.181 | 77185 | 9162.68 | 0.003 | 0.994 | 1.14 | 1.6099 | 41.22 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3771.753 | 100105 | 9531.285 | 0.004 | 0.994 | 1.82 | 2.026 | 11.32 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.26 | 32.978 | 469 | 4318.83 | 0 | 0.994 |  |  |  | MMI | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.25 | 291.109 | 7256 | 9861.838 | 0 | 0.994 | 0.17 | 0.07129 | -58.07 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.25 | 589.088 | 15622 | 10913.814 | 0.001 | 0.994 | 0.28 | 0.20169 | -27.97 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1572.219 | 38941 | 9692.394 | 0.002 | 0.994 | 0.44 | 0.77904 | 77.06 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1355.567 | 35995 | 9741.055 | 0.001 | 0.994 | 0.71 | 0.65605 | -7.6 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2749.095 | 72310 | 8925.449 | 0.003 | 0.994 | 1.14 | 1.5574 | 36.61 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3863.057 | 105763 | 9877.485 | 0.004 | 0.994 | 1.82 | 2.0013 | 9.96 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 3.2 | 49.8 | 661 |  |  | 0.994 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.25 | 369.467 | 9542 | 8900.187 | 0 | 0.994 | 0.17 | 0.13523 | -20.45 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.25 | 530.259 | 14582 | 9899.175 | 0.001 | 0.994 | 0.28 | 0.1995 | -28.75 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1511.87 | 41215 | 9065.441 | 0.002 | 0.994 | 0.44 | 0.80338 | 82.59 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1248.689 | 32255 | 9599.27 | 0.001 | 0.994 | 0.71 | 0.60761 | -14.42 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2588.253 | 67985 | 8059.229 | 0.003 | 0.994 | 1.14 | 1.6276 | 42.77 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3943.559 | 106325 | 8519.921 | 0.005 | 0.994 | 1.82 | 2.3849 | 31.04 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 3.21 | 59.636 | 545 |  |  | 0.994 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 3.15 | 83.196 | 909 |  |  | 0.994 |  |  |  | bb | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 3.16 | 66.668 | 687 |  |  | 0.994 |  |  |  | bb | 2020-11-24
Compound 9: PFTeDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 3.48 | 0.826 | 51 |  |  | 0.995 |  |  |  | bb | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 3.54 | 9.21 | 111 |  |  | 0.995 |  |  |  | bb | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 3.4 | 2.286 | 77 |  |  | 0.995 |  |  |  | bb | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 3.6 | 15.792 | 173 |  |  | 0.995 |  |  |  | bb | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.48 | 19.907 | 704 | 10173.215 | 0 | 0.995 |  |  |  | MMI | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 3.49 | 2.49 | 101 |  |  | 0.995 |  |  |  | bb | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.48 | 76.018 | 2541 | 9730.571 | 0 | 0.995 | 0.17 | 0.09251 | -45.58 | bb | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.48 | 150.396 | 5053 | 11120.537 | 0 | 0.995 | 0.28 | 0.23707 | -15.33 | bb | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.48 | 400.484 | 13499 | 10876.152 | 0 | 0.995 | 0.44 | 0.82675 | 87.9 | bb | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.48 | 269.788 | 9343 | 10360.407 | 0 | 0.995 | 0.71 | 0.55385 | -21.99 | bb | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.48 | 527.814 | 18224 | 10282.522 | 0.001 | 0.995 | 1.14 | 1.194 | 4.74 | bb | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.735 | 24808 | 10162.122 | 0.001 | 0.995 | 1.82 | 1.6949 | -6.87 | bb | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 3.53 | 5.236 | 227 |  |  | 0.995 |  |  |  | bb | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.48 | 961.143 | 32731 | 10490.147 | 0.001 | 0.995 | 2.91 | 2.214 | -23.92 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.48 | 2133.045 | 71385 | 9708.633 | 0.002 | 0.995 | 4.66 | 5.4574 | 17.11 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.48 | 2800.844 | 93515 | 9783.861 | 0.003 | 0.995 | 7.45 | 7.1437 | -4.11 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.48 | 4996.432 | 166370 | 11453.623 | 0.004 | 0.995 | 11.92 | 10.944 | -8.19 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.48 | 9734.145 | 318777 | 10865.012 | 0.009 | 0.995 | 19.07 | 22.607 | 18.55 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.48 | 13404.201 | 440392 | 11101.751 | 0.012 | 0.995 | 30.52 | 30.5213 | 0 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.48 | 33.321 | 1070 | 11173.664 | 0 | 0.995 |  |  |  | MMI | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.48 | 19948.268 | 645349 | 10508.111 | 0.019 | 0.995 | 48.83 | 48.1116 | -1.47 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.48 | 30695.457 | 997521 | 9989.264 | 0.031 | 0.995 | 78.13 | 78.1173 | -0.02 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.48 | 49368.879 | 1625311 | 10623.775 | 0.046 | 0.995 | 125 | 118.5485 | -5.16 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.48 | 67692.227 | 2185550 | 10923.09 | 0.062 | 0.995 | 156.25 | 158.6059 | 1.51 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.48 | 97595.797 | 3123183 | 9730.678 | 0.1 | 0.995 | 250 | 258.7078 | 3.48 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 3.48 | 20.559 | 696 |  |  | 0.995 |  |  |  | bb | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.48 | 234.138 | 7814 | 10475.948 | 0 | 0.995 | 0.63 | 0.46045 | -26.91 | bb | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.48 | 1541.29 | 52945 | 11635.099 | 0.001 | 0.995 | 2.5 | 3.2481 | 29.93 | bb | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.48 | 2821.402 | 93332 | 12150.41 | 0.002 | 0.995 | 6.25 | 5.774 | -7.62 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.48 | 10033.756 | 330538 | 10343.043 | 0.01 | 0.995 | 25 | 24.491 | -2.04 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 3.48 | 2.372 | 126 |  |  | 0.995 |  |  |  | bb | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.48 | 96.457 | 3267 | 11255 | 0 | 0.995 | 0.17 | 0.11169 | -34.3 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.48 | 160.949 | 5460 | 12190.335 | 0 | 0.995 | 0.28 | 0.22894 | -18.24 | bb | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.48 | 419.594 | 14218 | 11718.064 | 0 | 0.995 | 0.44 | 0.80107 | 82.06 | bb | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.48 | 286.428 | 9994 | 11317.867 | 0 | 0.995 | 0.71 | 0.53531 | -24.6 | bb | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.48 | 541.634 | 18738 | 10528.761 | 0.001 | 0.995 | 1.14 | 1.1968 | 4.99 | bb | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.48 | 827.981 | 28603 | 11242.785 | 0.001 | 0.995 | 1.82 | 1.7589 | -3.36 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.48 | 32.465 | 1024 | 12048.315 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.48 | 88.639 | 3102 | 11228.034 | 0 | 0.995 | 0.17 | 0.09459 | -44.36 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.48 | 149.478 | 5371 | 11829.166 | 0 | 0.995 | 0.28 | 0.2146 | -23.36 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.48 | 397.977 | 13732 | 11542.465 | 0 | 0.995 | 0.44 | 0.76745 | 74.42 | bb | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.48 | 280.302 | 9542 | 11854.473 | 0 | 0.995 | 0.71 | 0.49323 | -30.53 | bb | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.48 | 572.384 | 19665 | 10775.113 | 0.001 | 0.995 | 1.14 | 1.2393 | 8.71 | bb | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.48 | 759.402 | 25597 | 10874.418 | 0.001 | 0.995 | 1.82 | 1.6624 | -8.66 | bb | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 3.43 | 10.672 | 237 |  |  | 0.995 |  |  |  | bb | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.48 | 3711.082 | 121322 | 11440.534 | 0.003 | 0.995 |  | 8.1092 |  | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.48 | 4811.606 | 159244 | 13256.198 | 0.004 | 0.995 |  | 9.0871 |  | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 3.48 | 3409.649 | 110851 | 11719.4 | 0.003 | 0.995 |  | 7.2619 |  | MM | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.48 | 40.495 | 1302 | 11316.626 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.48 | 29.886 | 940 | 11121.704 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.48 | 33.682 | 1221 | 11384.287 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.48 | 28.652 | 771 | 11975.843 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.48 | 24.934 | 939 | 11676.331 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.48 | 26.242 | 767 | 11894.549 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.48 | 24.269 | 516 | 11048.28 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.48 | 43379.383 | 1400961 | 9862.112 | 0.044 | 0.995 |  | 112.1514 |  | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.48 | 43472.313 | 1406561 | 9161.524 | 0.047 | 0.995 |  | 121.0757 |  | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.48 | 42736.262 | 1384160 | 9049.197 | 0.047 | 0.995 |  | 120.4974 |  | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 3.48 | 19.075 | 513 |  |  | 0.995 |  |  |  | bb | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.48 | 77.155 | 1940 | 8897.753 | 0 | 0.995 |  | 0.11425 |  | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.48 | 73.03 | 2627 | 10825.349 | 0 | 0.995 |  | 0.06553 |  | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.48 | 58.565 | 1986 | 9994.433 | 0 | 0.995 |  | 0.04309 |  | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.48 | 26.115 | 857 | 8190.83 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.48 | 23.96 | 691 | 9179.411 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.48 | 25.058 | 756 | 10046.605 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 3.53 | 1.747 | 81 |  |  | 0.995 |  |  |  | bb | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.48 | 50968.484 | 1646832 | 11574.24 | 0.044 | 0.995 |  | 112.2807 |  | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.48 | 41411.648 | 1337742 | 8454.886 | 0.049 | 0.995 |  | 125.0164 |  | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.48 | 44135.258 | 1418305 | 8997.701 | 0.049 | 0.995 |  | 125.2025 |  | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.48 | 81.135 | 2930 | 9464.119 | 0 | 0.995 |  | 0.11176 |  | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.48 | 50.627 | 1825 | 8344.308 | 0 | 0.995 |  | 0.04834 |  | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.48 | 52.271 | 1924 | 8994.774 | 0 | 0.995 |  | 0.04186 |  | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.48 | 24.97 | 652 | 10773.584 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.48 | 22.398 | 811 | 9773.682 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.48 | 30.817 | 960 | 9471.696 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.48 | 24.207 | 564 | 9756.347 | 0 | 0.995 |  |  |  | bdI | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 3.48 | 26.431 | 815 | 10459.238 | 0 | 0.995 |  |  |  | bdI | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 3.49 | 5.293 | 146 |  |  | 0.995 |  |  |  | bb | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.48 | 93.003 | 2974 | 10426.968 | 0 | 0.995 | 0.17 | 0.12053 | -29.1 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.48 | 132.244 | 4290 | 11531.981 | 0 | 0.995 | 0.28 | 0.18502 | -33.92 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.48 | 397.124 | 13495 | 10908.972 | 0 | 0.995 | 0.44 | 0.81615 | 85.49 | bb | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.48 | 269.659 | 9183 | 10369.82 | 0 | 0.995 | 0.71 | 0.55294 | -22.12 | bb | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.48 | 548.507 | 18901 | 9679.873 | 0.001 | 0.995 | 1.14 | 1.329 | 16.58 | bb | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.994 | 24870 | 10265.066 | 0.001 | 0.995 | 1.82 | 1.6775 | -7.83 | bb | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 3.55 | 1.865 | 73 | 3.524 | 0.005 | 0.995 |  | 13.3019 |  | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.48 | 1026.99 | 34671 | 11516.313 | 0.001 | 0.995 | 2.91 | 2.1521 | -26.05 | bb | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.48 | 2207.919 | 71839 | 9848.446 | 0.002 | 0.995 | 4.66 | 5.571 | 19.55 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.48 | 2886.225 | 93590 | 10084.969 | 0.003 | 0.995 | 7.45 | 7.1416 | -4.14 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.48 | 4827.986 | 156305 | 10773.109 | 0.004 | 0.995 | 11.92 | 11.2462 | -5.65 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.48 | 10000.793 | 325065 | 11209.78 | 0.009 | 0.995 | 19.07 | 22.5113 | 18.05 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.48 | 13072.131 | 428255 | 10679.77 | 0.012 | 0.995 | 30.52 | 30.9437 | 1.39 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.48 | 32.773 | 1001 | 10166.878 | 0 | 0.995 |  |  |  | bbI | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.48 | 19642.855 | 633581 | 10697.722 | 0.018 | 0.995 | 48.83 | 46.5263 | -4.72 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.48 | 31198.275 | 1010198 | 10158.705 | 0.031 | 0.995 | 78.13 | 78.0724 | -0.07 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.48 | 47129.316 | 1509683 | 10158.639 | 0.046 | 0.995 | 125 | 118.3505 | -5.32 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.48 | 63054.848 | 2014088 | 9865.15 | 0.064 | 0.995 | 156.25 | 163.6495 | 4.74 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.48 | 92332.867 | 2901464 | 9832.22 | 0.094 | 0.995 | 250 | 241.9142 | -3.23 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 3.48 | 40.174 | 694 |  |  | 0.995 |  |  |  | bb | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.48 | 226.375 | 7665 | 10532.542 | 0 | 0.995 | 0.63 | 0.43876 | -30.36 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.48 | 824.583 | 27864 | 6178.086 | 0.001 | 0.995 | 2.5 | 3.2735 | 30.94 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.48 | 2458.629 | 82647 | 10931.824 | 0.002 | 0.995 | 6.25 | 5.5891 | -10.57 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.48 | 9092.92 | 297083 | 9113.911 | 0.01 | 0.995 | 25 | 25.1921 | 0.77 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 3.44 | 5.467 | 68 |  |  | 0.995 |  |  |  | bd | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.48 | 24.036 | 656 | 9387.197 | 0 | 0.995 |  |  |  | MMI | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.48 | 89.666 | 2956 | 9798.542 | 0 | 0.995 | 0.17 | 0.12639 | -25.65 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.48 | 126.176 | 4428 | 10690.253 | 0 | 0.995 | 0.28 | 0.19351 | -30.89 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.48 | 359.254 | 12358 | 10546.275 | 0 | 0.995 | 0.44 | 0.75695 | 72.03 | bb | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.48 | 270.067 | 9229 | 10312.553 | 0 | 0.995 | 0.71 | 0.55759 | -21.47 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.48 | 498.669 | 17122 | 9162.68 | 0.001 | 0.995 | 1.14 | 1.2723 | 11.6 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.48 | 691.849 | 23128 | 9531.285 | 0.001 | 0.995 | 1.82 | 1.7321 | -4.83 | bb | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.48 | 20.635 | 411 | 4318.83 | 0 | 0.995 |  | 0.01571 |  | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.48 | 87.609 | 2972 | 9861.838 | 0 | 0.995 | 0.17 | 0.11962 | -29.63 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.48 | 133.241 | 4606 | 10913.814 | 0 | 0.995 | 0.28 | 0.20377 | -27.22 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.48 | 346.95 | 11689 | 9692.394 | 0 | 0.995 | 0.44 | 0.80078 | 82 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.48 | 241.777 | 8413 | 9741.055 | 0 | 0.995 | 0.71 | 0.52298 | -26.34 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.48 | 464.422 | 16113 | 8925.449 | 0.001 | 0.995 | 1.14 | 1.2118 | 6.3 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.285 | 24130 | 9877.485 | 0.001 | 0.995 | 1.82 | 1.7457 | -4.08 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 3.57 | 2.565 | 101 |  |  | 0.995 |  |  |  | bb | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.48 | 66.422 | 2167 | 8900.187 | 0 | 0.995 | 0.17 | 0.08367 | -50.78 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.48 | 134.751 | 4468 | 9899.175 | 0 | 0.995 | 0.28 | 0.2393 | -14.54 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.48 | 338.576 | 11860 | 9065.441 | 0 | 0.995 | 0.44 | 0.84006 | 90.92 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.48 | 252.29 | 8791 | 9599.27 | 0 | 0.995 | 0.71 | 0.55997 | -21.13 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.48 | 430.803 | 15089 | 8059.229 | 0.001 | 0.995 | 1.14 | 1.2478 | 9.45 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.48 | 631.31 | 21509 | 8519.921 | 0.001 | 0.995 | 1.82 | 1.7703 | -2.73 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 3.53 | 2.172 | 67 |  |  | 0.995 |  |  |  | bb | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 3.53 | 10.459 | 145 |  |  | 0.995 |  |  |  | bb | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 3.55 | 1.004 | 55 |  |  | 0.995 |  |  |  | bb | 2020-11-24
Compound 10: MPFBA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.06 | 50185.297 | 1079128 |  | 50185.297 |  | 0.01 | 0.00956 | -4.39 | bb | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.06 | 49927.879 | 1093793 |  | 49927.879 |  | 0.01 | 0.00951 | -4.88 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.06 | 49278.141 | 1055807 |  | 49278.141 |  | 0.01 | 0.00939 | -6.12 | bb | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.06 | 50032.969 | 1093526 |  | 50032.969 |  | 0.01 | 0.00953 | -4.68 | bb | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.06 | 48869.313 | 1083320 |  | 48869.313 |  | 0.01 | 0.00931 | -6.9 | bb | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.06 | 49424.641 | 1122354 |  | 49424.641 |  | 0.01 | 0.00942 | -5.84 | bb | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.06 | 48781.555 | 1097833 |  | 48781.555 |  | 0.01 | 0.00929 | -7.07 | bb | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.06 | 50327.848 | 1119790 |  | 50327.848 |  | 0.01 | 0.00959 | -4.12 | bb | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.06 | 49203.574 | 1136669 |  | 49203.574 |  | 0.01 | 0.00937 | -6.26 | bb | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.06 | 47733.922 | 1082948 |  | 47733.922 |  | 0.01 | 0.00909 | -9.06 | bb | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.06 | 50520.066 | 1152278 |  | 50520.066 |  | 0.01 | 0.00962 | -3.76 | bb | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.06 | 49274.195 | 1120503 |  | 49274.195 |  | 0.01 | 0.00939 | -6.13 | bb | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.06 | 49934.551 | 1152862 |  | 49934.551 |  | 0.01 | 0.00951 | -4.87 | bb | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.06 | 50805.66 | 1121952 |  | 50805.66 |  | 0.01 | 0.00968 | -3.21 | bb | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.06 | 50913.094 | 1176729 |  | 50913.094 |  | 0.01 | 0.0097 | -3.01 | bb | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.06 | 48568.758 | 1127876 |  | 48568.758 |  | 0.01 | 0.00925 | -7.47 | bb | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.06 | 51406.383 | 1220473 |  | 51406.383 |  | 0.01 | 0.00979 | -2.07 | bb | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.06 | 51234.66 | 1192786 |  | 51234.66 |  | 0.01 | 0.00976 | -2.39 | bb | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.06 | 51287.395 | 1203006 |  | 51287.395 |  | 0.01 | 0.00977 | -2.29 | bb | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.06 | 49652.594 | 1169542 |  | 49652.594 |  | 0.01 | 0.00946 | -5.41 | bb | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.06 | 52099.977 | 1294191 |  | 52099.977 |  | 0.01 | 0.00993 | -0.75 | bb | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.06 | 51356.805 | 1175106 |  | 51356.805 |  | 0.01 | 0.00978 | -2.16 | bb | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.06 | 51164.73 | 1208782 |  | 51164.73 |  | 0.01 | 0.00975 | -2.53 | bb | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.06 | 50434.191 | 1144254 |  | 50434.191 |  | 0.01 | 0.00961 | -3.92 | bb | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.06 | 51354.762 | 1178770 |  | 51354.762 |  | 0.01 | 0.00978 | -2.17 | bb | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.06 | 51327.613 | 1155420 |  | 51327.613 |  | 0.01 | 0.00978 | -2.22 | bb | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.06 | 49470.234 | 1141155 |  | 49470.234 |  | 0.01 | 0.00942 | -5.76 | bb | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.06 | 49787.16 | 1161960 |  | 49787.16 |  | 0.01 | 0.00948 | -5.15 | bb | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.06 | 50241.621 | 1172509 |  | 50241.621 |  | 0.01 | 0.00957 | -4.29 | bb | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.06 | 53139.266 | 1208079 |  | 53139.266 |  | 0.01 | 0.01012 | 1.23 | bb | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.06 | 51236.457 | 1180196 |  | 51236.457 |  | 0.01 | 0.00976 | -2.39 | bb | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.06 | 51647.742 | 1197739 |  | 51647.742 |  | 0.01 | 0.00984 | -1.61 | bb | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.06 | 51511.32 | 1185758 |  | 51511.32 |  | 0.01 | 0.00981 | -1.87 | bb | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.06 | 50949.309 | 1197432 |  | 50949.309 |  | 0.01 | 0.00971 | -2.94 | bb | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.06 | 50855.781 | 1213615 |  | 50855.781 |  | 0.01 | 0.00969 | -3.12 | bb | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.06 | 50099.988 | 1170203 |  | 50099.988 |  | 0.01 | 0.00954 | -4.56 | bb | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.06 | 55329.66 | 1289571 |  | 55329.66 |  | 0.01 | 0.01054 | 5.41 | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.06 | 55421.43 | 1277857 |  | 55421.43 |  | 0.01 | 0.01056 | 5.58 | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.06 | 52817.883 | 1207031 |  | 52817.883 |  | 0.01 | 0.01006 | 0.62 | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.06 | 53182.844 | 1285004 |  | 53182.844 |  | 0.01 | 0.01013 | 1.32 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.06 | 53391.262 | 1296578 |  | 53391.262 |  | 0.01 | 0.01017 | 1.71 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.06 | 53434.508 | 1288257 |  | 53434.508 |  | 0.01 | 0.01018 | 1.8 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.06 | 53001.285 | 1240747 |  | 53001.285 |  | 0.01 | 0.0101 | 0.97 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.06 | 54932.02 | 1311721 |  | 54932.02 |  | 0.01 | 0.01046 | 4.65 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.06 | 54176.371 | 1285553 |  | 54176.371 |  | 0.01 | 0.01032 | 3.21 | bb | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.06 | 52035.281 | 1237776 |  | 52035.281 |  | 0.01 | 0.00991 | -0.87 | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.06 | 54616.355 | 1345293 |  | 54616.355 |  | 0.01 | 0.0104 | 4.05 | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.06 | 53656.375 | 1364620 |  | 53656.375 |  | 0.01 | 0.01022 | 2.22 | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.06 | 51797.668 | 1288013 |  | 51797.668 |  | 0.01 | 0.00987 | -1.32 | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 1.08 | 4.41 | 155 |  | 4.41 |  | 0.01 | 0 | -99.99 | bb | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.06 | 52322.02 | 1310806 |  | 52322.02 |  | 0.01 | 0.00997 | -0.32 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.06 | 53899.488 | 1352836 |  | 53899.488 |  | 0.01 | 0.01027 | 2.68 | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.06 | 53863.215 | 1355801 |  | 53863.215 |  | 0.01 | 0.01026 | 2.61 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.06 | 52084.168 | 1324937 |  | 52084.168 |  | 0.01 | 0.00992 | -0.78 | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.06 | 52470.254 | 1353608 |  | 52470.254 |  | 0.01 | 0.01 | -0.04 | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.06 | 53591.719 | 1369970 |  | 53591.719 |  | 0.01 | 0.01021 | 2.1 | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.06 | 59098.656 | 1403313 |  | 59098.656 |  | 0.01 | 0.01126 | 12.59 | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.06 | 54825.996 | 1429172 |  | 54825.996 |  | 0.01 | 0.01044 | 4.45 | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.06 | 53523.332 | 1369929 |  | 53523.332 |  | 0.01 | 0.0102 | 1.97 | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.06 | 54352.184 | 1413246 |  | 54352.184 |  | 0.01 | 0.01035 | 3.54 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.06 | 55028.941 | 1441641 |  | 55028.941 |  | 0.01 | 0.01048 | 4.83 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.06 | 53986.629 | 1407332 |  | 53986.629 |  | 0.01 | 0.01028 | 2.85 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.06 | 53940.129 | 1314646 |  | 53940.129 |  | 0.01 | 0.01028 | 2.76 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.06 | 54814.41 | 1382624 |  | 54814.41 |  | 0.01 | 0.01044 | 4.42 | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.06 | 54693.875 | 1429813 |  | 54693.875 |  | 0.01 | 0.01042 | 4.2 | bb | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.06 | 55158.957 | 1416562 |  | 55158.957 |  | 0.01 | 0.01051 | 5.08 | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.06 | 54180.07 | 1338242 |  | 54180.07 |  | 0.01 | 0.01032 | 3.22 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.06 | 53647.801 | 1306898 |  | 53647.801 |  | 0.01 | 0.01022 | 2.2 | bb | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.06 | 55028.125 | 1360267 |  | 55028.125 |  | 0.01 | 0.01048 | 4.83 | bb | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.06 | 53959.531 | 1333779 |  | 53959.531 |  | 0.01 | 0.01028 | 2.8 | bb | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.06 | 52453.168 | 1314141 |  | 52453.168 |  | 0.01 | 0.00999 | -0.07 | bb | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.06 | 52932.809 | 1342364 |  | 52932.809 |  | 0.01 | 0.01008 | 0.84 | bb | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.06 | 53059.031 | 1324542 |  | 53059.031 |  | 0.01 | 0.01011 | 1.08 | bb | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.06 | 54954.105 | 1332580 |  | 54954.105 |  | 0.01 | 0.01047 | 4.69 | bb | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.06 | 53011.246 | 1295984 |  | 53011.246 |  | 0.01 | 0.0101 | 0.99 | bb | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.06 | 51497.086 | 1264257 |  | 51497.086 |  | 0.01 | 0.00981 | -1.89 | bb | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.06 | 54404.949 | 1359505 |  | 54404.949 |  | 0.01 | 0.01036 | 3.64 | bb | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.06 | 54264.176 | 1381855 |  | 54264.176 |  | 0.01 | 0.01034 | 3.38 | bb | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.06 | 53552.957 | 1318423 |  | 53552.957 |  | 0.01 | 0.0102 | 2.02 | bb | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.06 | 55176.629 | 1426925 |  | 55176.629 |  | 0.01 | 0.01051 | 5.11 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.06 | 53773.059 | 1374467 |  | 53773.059 |  | 0.01 | 0.01024 | 2.44 | bb | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.06 | 52683.031 | 1324650 |  | 52683.031 |  | 0.01 | 0.01004 | 0.36 | bb | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.06 | 55300.879 | 1439829 |  | 55300.879 |  | 0.01 | 0.01054 | 5.35 | bb | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.06 | 53606.293 | 1358384 |  | 53606.293 |  | 0.01 | 0.01021 | 2.12 | bb | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.06 | 54873.789 | 1419838 |  | 54873.789 |  | 0.01 | 0.01045 | 4.54 | bb | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.06 | 54029.367 | 1383832 |  | 54029.367 |  | 0.01 | 0.01029 | 2.93 | bb | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.07 | 61291.387 | 1905970 |  | 61291.387 |  | 0.01 | 0.01168 | 16.76 | bb | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.06 | 53206.5 | 1365779 |  | 53206.5 |  | 0.01 | 0.01014 | 1.36 | bb | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.06 | 54984.762 | 1446908 |  | 54984.762 |  | 0.01 | 0.01047 | 4.75 | bb | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.06 | 57833.516 | 1565582 |  | 57833.516 |  | 0.01 | 0.01102 | 10.18 | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.06 | 54130.621 | 1397893 |  | 54130.621 |  | 0.01 | 0.01031 | 3.12 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.06 | 54825.285 | 1426003 |  | 54825.285 |  | 0.01 | 0.01044 | 4.45 | bb | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.06 | 55636.879 | 1448618 |  | 55636.879 |  | 0.01 | 0.0106 | 5.99 | bb | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.06 | 54984.301 | 1456655 |  | 54984.301 |  | 0.01 | 0.01047 | 4.75 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54067.758 | 1437007 |  | 54067.758 |  | 0.01 | 0.0103 | 3 | bb | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.06 | 53643.699 | 1379321 |  | 53643.699 |  | 0.01 | 0.01022 | 2.19 | bb | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.06 | 19403.883 | 608476 |  | 19403.883 |  | 0.01 | 0.0037 | -63.03 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.06 | 55147.813 | 1457940 |  | 55147.813 |  | 0.01 | 0.01051 | 5.06 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.06 | 56832.938 | 1497038 |  | 56832.938 |  | 0.01 | 0.01083 | 8.27 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.06 | 55401.938 | 1475862 |  | 55401.938 |  | 0.01 | 0.01055 | 5.54 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.06 | 54257.391 | 1447254 |  | 54257.391 |  | 0.01 | 0.01034 | 3.36 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54348.875 | 1468223 |  | 54348.875 |  | 0.01 | 0.01035 | 3.54 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.06 | 54800.516 | 1443545 |  | 54800.516 |  | 0.01 | 0.01044 | 4.4 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.06 | 55044.574 | 1487081 |  | 55044.574 |  | 0.01 | 0.01049 | 4.86 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.06 | 55853.539 | 1506542 |  | 55853.539 |  | 0.01 | 0.01064 | 6.4 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.06 | 56901.824 | 1536143 |  | 56901.824 |  | 0.01 | 0.01084 | 8.4 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.06 | 55579.008 | 1518421 |  | 55579.008 |  | 0.01 | 0.01059 | 5.88 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54417.223 | 1506331 |  | 54417.223 |  | 0.01 | 0.01037 | 3.67 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.06 | 54960.121 | 1474615 |  | 54960.121 |  | 0.01 | 0.01047 | 4.7 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
Compound 11: M3PFBS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.69 | 21628.184 | 1107684 |  | 21628.184 |  | 0.01 | 0.0093 | -7 | bb | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.69 | 21861.801 | 1121154 |  | 21861.801 |  | 0.01 | 0.0094 | -6 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.69 | 22905.938 | 1157712 |  | 22905.938 |  | 0.01 | 0.00985 | -1.51 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.69 | 22565.848 | 1153684 |  | 22565.848 |  | 0.01 | 0.0097 | -2.97 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.69 | 22329.166 | 1134724 |  | 22329.166 |  | 0.01 | 0.0096 | -3.99 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.69 | 22016.998 | 1130137 |  | 22016.998 |  | 0.01 | 0.00947 | -5.33 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.69 | 22032.42 | 1128433 |  | 22032.42 |  | 0.01 | 0.00947 | -5.26 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.69 | 22609.912 | 1152922 |  | 22609.912 |  | 0.01 | 0.00972 | -2.78 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.69 | 21791.783 | 1105980 |  | 21791.783 |  | 0.01 | 0.00937 | -6.3 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.69 | 21652.168 | 1101526 |  | 21652.168 |  | 0.01 | 0.00931 | -6.9 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.69 | 21857.189 | 1100713 |  | 21857.189 |  | 0.01 | 0.0094 | -6.02 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.69 | 22171.375 | 1122924 |  | 22171.375 |  | 0.01 | 0.00953 | -4.67 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.69 | 21427.695 | 1074936 |  | 21427.695 |  | 0.01 | 0.00921 | -7.86 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.69 | 22818.307 | 1167930 |  | 22818.307 |  | 0.01 | 0.00981 | -1.89 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.69 | 20779.268 | 1021019 |  | 20779.268 |  | 0.01 | 0.00893 | -10.65 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.69 | 19147.74 | 938578 |  | 19147.74 |  | 0.01 | 0.00823 | -17.67 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.69 | 18008.693 | 859605 |  | 18008.693 |  | 0.01 | 0.00774 | -22.57 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.69 | 17279.416 | 819807 |  | 17279.416 |  | 0.01 | 0.00743 | -25.7 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.69 | 15521.709 | 719428 |  | 15521.709 |  | 0.01 | 0.00667 | -33.26 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.69 | 22865.168 | 1162819 |  | 22865.168 |  | 0.01 | 0.00983 | -1.68 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.69 | 24861.91 | 1255348 |  | 24861.91 |  | 0.01 | 0.01069 | 6.9 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.69 | 23299.66 | 1180502 |  | 23299.66 |  | 0.01 | 0.01002 | 0.18 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.69 | 21679.879 | 1091308 |  | 21679.879 |  | 0.01 | 0.00932 | -6.78 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.69 | 23188.197 | 1175156 |  | 23188.197 |  | 0.01 | 0.00997 | -0.29 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.69 | 23955.818 | 1215689 |  | 23955.818 |  | 0.01 | 0.0103 | 3.01 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.69 | 23133.342 | 1179881 |  | 23133.342 |  | 0.01 | 0.00995 | -0.53 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.69 | 22994.361 | 1176313 |  | 22994.361 |  | 0.01 | 0.00989 | -1.13 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.69 | 22848.482 | 1164140 |  | 22848.482 |  | 0.01 | 0.00982 | -1.76 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.69 | 22995.889 | 1178974 |  | 22995.889 |  | 0.01 | 0.00989 | -1.12 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.69 | 22658.945 | 1182864 |  | 22658.945 |  | 0.01 | 0.00974 | -2.57 | bb | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.69 | 22730.893 | 1171782 |  | 22730.893 |  | 0.01 | 0.00977 | -2.26 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.69 | 24510.729 | 1260704 |  | 24510.729 |  | 0.01 | 0.01054 | 5.39 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.69 | 24355.15 | 1247996 |  | 24355.15 |  | 0.01 | 0.01047 | 4.72 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.69 | 23756.479 | 1219422 |  | 23756.479 |  | 0.01 | 0.01021 | 2.15 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.69 | 23377.701 | 1195150 |  | 23377.701 |  | 0.01 | 0.01005 | 0.52 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.69 | 23030.449 | 1173826 |  | 23030.449 |  | 0.01 | 0.0099 | -0.97 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.69 | 25180.121 | 1308131 |  | 25180.121 |  | 0.01 | 0.01083 | 8.27 | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.69 | 25625.738 | 1327128 |  | 25625.738 |  | 0.01 | 0.01102 | 10.19 | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.69 | 24362.371 | 1265818 |  | 24362.371 |  | 0.01 | 0.01048 | 4.75 | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.69 | 24305.764 | 1252778 |  | 24305.764 |  | 0.01 | 0.01045 | 4.51 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.69 | 24109.816 | 1246263 |  | 24109.816 |  | 0.01 | 0.01037 | 3.67 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.69 | 24456.668 | 1274867 |  | 24456.668 |  | 0.01 | 0.01052 | 5.16 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.69 | 23242.805 | 1191134 |  | 23242.805 |  | 0.01 | 0.00999 | -0.06 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.69 | 24957.41 | 1282631 |  | 24957.41 |  | 0.01 | 0.01073 | 7.31 | MM | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.69 | 25152.914 | 1299701 |  | 25152.914 |  | 0.01 | 0.01082 | 8.15 | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.69 | 24200.555 | 1242784 |  | 24200.555 |  | 0.01 | 0.01041 | 4.06 | MM | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.69 | 22887.045 | 1178188 |  | 22887.045 |  | 0.01 | 0.00984 | -1.59 | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.69 | 22461.596 | 1154245 |  | 22461.596 |  | 0.01 | 0.00966 | -3.42 | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.69 | 23178.355 | 1180150 |  | 23178.355 |  | 0.01 | 0.00997 | -0.34 | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.69 | 22570.561 | 1168954 |  | 22570.561 |  | 0.01 | 0.0097 | -2.95 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.69 | 23917.525 | 1244149 |  | 23917.525 |  | 0.01 | 0.01028 | 2.84 | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.69 | 23604.629 | 1226601 |  | 23604.629 |  | 0.01 | 0.01015 | 1.5 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.69 | 17613.797 | 840636 |  | 17613.797 |  | 0.01 | 0.00757 | -24.26 | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.69 | 17847.734 | 849091 |  | 17847.734 |  | 0.01 | 0.00767 | -23.26 | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.69 | 18605.641 | 878553 |  | 18605.641 |  | 0.01 | 0.008 | -20 | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.69 | 26959.02 | 1406070 |  | 26959.02 |  | 0.01 | 0.01159 | 15.92 | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.69 | 23924.834 | 1249809 |  | 23924.834 |  | 0.01 | 0.01029 | 2.87 | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.69 | 24150.537 | 1252542 |  | 24150.537 |  | 0.01 | 0.01038 | 3.84 | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.69 | 23967.572 | 1236497 |  | 23967.572 |  | 0.01 | 0.01031 | 3.06 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.69 | 23285.992 | 1205143 |  | 23285.992 |  | 0.01 | 0.01001 | 0.13 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.69 | 24485.539 | 1255107 |  | 24485.539 |  | 0.01 | 0.01053 | 5.28 | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.69 | 24092.141 | 1241226 |  | 24092.141 |  | 0.01 | 0.01036 | 3.59 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.69 | 19292.301 | 915616 |  | 19292.301 |  | 0.01 | 0.0083 | -17.05 | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.69 | 18221.59 | 863347 |  | 18221.59 |  | 0.01 | 0.00783 | -21.65 | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.69 | 18441.566 | 862143 |  | 18441.566 |  | 0.01 | 0.00793 | -20.7 | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.69 | 24051.273 | 1252444 |  | 24051.273 |  | 0.01 | 0.01034 | 3.42 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.69 | 24736.178 | 1278515 |  | 24736.178 |  | 0.01 | 0.01064 | 6.36 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.69 | 25389.068 | 1294965 |  | 25389.068 |  | 0.01 | 0.01092 | 9.17 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.69 | 24953.592 | 1265246 |  | 24953.592 |  | 0.01 | 0.01073 | 7.3 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.69 | 24666.17 | 1268269 |  | 24666.17 |  | 0.01 | 0.01061 | 6.06 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.69 | 24743.717 | 1282216 |  | 24743.717 |  | 0.01 | 0.01064 | 6.39 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.69 | 24207.008 | 1245251 |  | 24207.008 |  | 0.01 | 0.01041 | 4.09 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.69 | 24554.58 | 1254301 |  | 24554.58 |  | 0.01 | 0.01056 | 5.58 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.69 | 23955.301 | 1231520 |  | 23955.301 |  | 0.01 | 0.0103 | 3 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.69 | 23626.941 | 1194339 |  | 23626.941 |  | 0.01 | 0.01016 | 1.59 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.69 | 24218.33 | 1223239 |  | 24218.33 |  | 0.01 | 0.01041 | 4.13 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.69 | 24038.738 | 1201898 |  | 24038.738 |  | 0.01 | 0.01034 | 3.36 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.69 | 23164.816 | 1156204 |  | 23164.816 |  | 0.01 | 0.00996 | -0.4 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.69 | 24598.893 | 1287087 |  | 24598.893 |  | 0.01 | 0.01058 | 5.77 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.69 | 22853.115 | 1129490 |  | 22853.115 |  | 0.01 | 0.00983 | -1.74 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.69 | 20442.584 | 987649 |  | 20442.584 |  | 0.01 | 0.00879 | -12.1 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.69 | 19474.186 | 935603 |  | 19474.186 |  | 0.01 | 0.00837 | -16.26 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.69 | 18660.311 | 879229 |  | 18660.311 |  | 0.01 | 0.00802 | -19.76 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.69 | 16898.207 | 769962 |  | 16898.207 |  | 0.01 | 0.00727 | -27.34 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.69 | 24217.922 | 1237288 |  | 24217.922 |  | 0.01 | 0.01041 | 4.13 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.69 | 29108.41 | 1484585 |  | 29108.41 |  | 0.01 | 0.01252 | 25.16 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.69 | 24923.5 | 1262333 |  | 24923.5 |  | 0.01 | 0.01072 | 7.17 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.69 | 24550.033 | 1239441 |  | 24550.033 |  | 0.01 | 0.01056 | 5.56 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.69 | 25048.125 | 1289327 |  | 25048.125 |  | 0.01 | 0.01077 | 7.7 | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.69 | 24669.846 | 1260177 |  | 24669.846 |  | 0.01 | 0.01061 | 6.08 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26169.432 | 1345344 |  | 26169.432 |  | 0.01 | 0.01125 | 12.52 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26166.074 | 1344524 |  | 26166.074 |  | 0.01 | 0.01125 | 12.51 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25000.137 | 1277774 |  | 25000.137 |  | 0.01 | 0.01075 | 7.5 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25457.678 | 1306186 |  | 25457.678 |  | 0.01 | 0.01095 | 9.46 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.69 | 24884.777 | 1263125 |  | 24884.777 |  | 0.01 | 0.0107 | 7 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.69 | 8914.928 | 470050 |  | 8914.928 |  | 0.01 | 0.00383 | -61.67 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.69 | 25334.619 | 1298374 |  | 25334.619 |  | 0.01 | 0.01089 | 8.93 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26174.143 | 1354444 |  | 26174.143 |  | 0.01 | 0.01125 | 12.54 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26063.229 | 1338491 |  | 26063.229 |  | 0.01 | 0.01121 | 12.07 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25239.68 | 1317405 |  | 25239.68 |  | 0.01 | 0.01085 | 8.53 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25575.846 | 1318077 |  | 25575.846 |  | 0.01 | 0.011 | 9.97 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.69 | 25309.125 | 1302560 |  | 25309.125 |  | 0.01 | 0.01088 | 8.82 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.69 | 25002.83 | 1282939 |  | 25002.83 |  | 0.01 | 0.01075 | 7.51 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26803.898 | 1364492 |  | 26803.898 |  | 0.01 | 0.01153 | 15.25 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26872.852 | 1385291 |  | 26872.852 |  | 0.01 | 0.01155 | 15.55 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25737.27 | 1325483 |  | 25737.27 |  | 0.01 | 0.01107 | 10.67 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25205.348 | 1300646 |  | 25205.348 |  | 0.01 | 0.01084 | 8.38 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.69 | 25344.141 | 1298578 |  | 25344.141 |  | 0.01 | 0.0109 | 8.98 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
Compound 12: M4PFHpA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.76 | 66453.484 | 3305070 |  | 66453.484 |  | 0.01 | 0.00947 | -5.33 | bb | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.76 | 66855.516 | 3320399 |  | 66855.516 |  | 0.01 | 0.00952 | -4.75 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.76 | 70175.805 | 3509487 |  | 70175.805 |  | 0.01 | 0.01 | -0.02 | bb | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.76 | 69660.461 | 3393695 |  | 69660.461 |  | 0.01 | 0.00992 | -0.76 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.76 | 69179.859 | 3379878 |  | 69179.859 |  | 0.01 | 0.00986 | -1.44 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.76 | 65949.148 | 3240065 |  | 65949.148 |  | 0.01 | 0.0094 | -6.04 | bb | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.76 | 65820.242 | 3263160 |  | 65820.242 |  | 0.01 | 0.00938 | -6.23 | bb | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.76 | 69313.656 | 3428823 |  | 69313.656 |  | 0.01 | 0.00987 | -1.25 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.76 | 67371.406 | 3263118 |  | 67371.406 |  | 0.01 | 0.0096 | -4.02 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.76 | 65924.273 | 3208166 |  | 65924.273 |  | 0.01 | 0.00939 | -6.08 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.76 | 66429.594 | 3218192 |  | 66429.594 |  | 0.01 | 0.00946 | -5.36 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.76 | 67007.094 | 3238182 |  | 67007.094 |  | 0.01 | 0.00955 | -4.54 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.76 | 65890.977 | 3142368 |  | 65890.977 |  | 0.01 | 0.00939 | -6.13 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.76 | 70428.734 | 3462973 |  | 70428.734 |  | 0.01 | 0.01003 | 0.34 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.76 | 64077.266 | 3046624 |  | 64077.266 |  | 0.01 | 0.00913 | -8.71 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.76 | 59706.527 | 2813015 |  | 59706.527 |  | 0.01 | 0.00851 | -14.94 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.76 | 56827.754 | 2640782 |  | 56827.754 |  | 0.01 | 0.0081 | -19.04 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.76 | 54128.734 | 2507089 |  | 54128.734 |  | 0.01 | 0.00771 | -22.88 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.76 | 50352.082 | 2284973 |  | 50352.082 |  | 0.01 | 0.00717 | -28.26 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.76 | 69882.695 | 3425608 |  | 69882.695 |  | 0.01 | 0.00996 | -0.44 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.76 | 74668.664 | 3631655 |  | 74668.664 |  | 0.01 | 0.01064 | 6.38 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.76 | 71009.594 | 3416943 |  | 71009.594 |  | 0.01 | 0.01012 | 1.17 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.76 | 66898.727 | 3226693 |  | 66898.727 |  | 0.01 | 0.00953 | -4.69 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.76 | 70140.469 | 3433732 |  | 70140.469 |  | 0.01 | 0.00999 | -0.07 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.76 | 73291.031 | 3595913 |  | 73291.031 |  | 0.01 | 0.01044 | 4.42 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.76 | 71199 | 3475520 |  | 71199 |  | 0.01 | 0.01014 | 1.44 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.76 | 70969.344 | 3452449 |  | 70969.344 |  | 0.01 | 0.01011 | 1.11 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.76 | 69821.609 | 3399067 |  | 69821.609 |  | 0.01 | 0.00995 | -0.53 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.76 | 70147.031 | 3440420 |  | 70147.031 |  | 0.01 | 0.00999 | -0.06 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.76 | 70267.711 | 3467658 |  | 70267.711 |  | 0.01 | 0.01001 | 0.11 | MM | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.76 | 69550.805 | 3433691 |  | 69550.805 |  | 0.01 | 0.00991 | -0.91 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.76 | 72982.797 | 3554466 |  | 72982.797 |  | 0.01 | 0.0104 | 3.98 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.76 | 73203.641 | 3581378 |  | 73203.641 |  | 0.01 | 0.01043 | 4.29 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.76 | 71708.977 | 3520012 |  | 71708.977 |  | 0.01 | 0.01022 | 2.16 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.76 | 69097.773 | 3396196 |  | 69097.773 |  | 0.01 | 0.00984 | -1.56 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.76 | 69563.125 | 3425429 |  | 69563.125 |  | 0.01 | 0.00991 | -0.9 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.76 | 75833.383 | 3743159 |  | 75833.383 |  | 0.01 | 0.0108 | 8.04 | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.76 | 78561.383 | 3845300 |  | 78561.383 |  | 0.01 | 0.01119 | 11.92 | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.76 | 75786.555 | 3701649 |  | 75786.555 |  | 0.01 | 0.0108 | 7.97 | MM | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.76 | 75532.336 | 3711094 |  | 75532.336 |  | 0.01 | 0.01076 | 7.61 | MM | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.76 | 73999.875 | 3664879 |  | 73999.875 |  | 0.01 | 0.01054 | 5.43 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.76 | 73473.383 | 3686369 |  | 73473.383 |  | 0.01 | 0.01047 | 4.68 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.76 | 71625.82 | 3505632 |  | 71625.82 |  | 0.01 | 0.0102 | 2.04 | MM | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.76 | 72956.234 | 3652750 |  | 72956.234 |  | 0.01 | 0.01039 | 3.94 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.76 | 75183.023 | 3684570 |  | 75183.023 |  | 0.01 | 0.01071 | 7.11 | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.76 | 72624.836 | 3574495 |  | 72624.836 |  | 0.01 | 0.01035 | 3.47 | MM | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.76 | 57842.676 | 2695074 |  | 57842.676 |  | 0.01 | 0.00824 | -17.59 | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.76 | 55906.832 | 2600066 |  | 55906.832 |  | 0.01 | 0.00796 | -20.35 | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.76 | 57936.676 | 2688977 |  | 57936.676 |  | 0.01 | 0.00825 | -17.46 | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.76 | 68748.156 | 3423723 |  | 68748.156 |  | 0.01 | 0.00979 | -2.06 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.76 | 73999.633 | 3635966 |  | 73999.633 |  | 0.01 | 0.01054 | 5.43 | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.76 | 71143.813 | 3516017 |  | 71143.813 |  | 0.01 | 0.01014 | 1.36 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.76 | 69815.391 | 3450382 |  | 69815.391 |  | 0.01 | 0.00995 | -0.54 | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.76 | 71290.555 | 3511303 |  | 71290.555 |  | 0.01 | 0.01016 | 1.57 | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.76 | 74547.953 | 3626257 |  | 74547.953 |  | 0.01 | 0.01062 | 6.21 | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.76 | 65037.824 | 3074126 |  | 65037.824 |  | 0.01 | 0.00927 | -7.34 | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.76 | 58343.555 | 2721553 |  | 58343.555 |  | 0.01 | 0.00831 | -16.88 | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.76 | 59147.715 | 2757025 |  | 59147.715 |  | 0.01 | 0.00843 | -15.73 | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.76 | 71500.539 | 3564321 |  | 71500.539 |  | 0.01 | 0.01019 | 1.87 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.76 | 70718.625 | 3511219 |  | 70718.625 |  | 0.01 | 0.01008 | 0.75 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.76 | 71977.289 | 3577291 |  | 71977.289 |  | 0.01 | 0.01025 | 2.54 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.76 | 73164.828 | 3636845 |  | 73164.828 |  | 0.01 | 0.01042 | 4.24 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.76 | 75227.086 | 3777262 |  | 75227.086 |  | 0.01 | 0.01072 | 7.17 | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.76 | 72306.594 | 3534427 |  | 72306.594 |  | 0.01 | 0.0103 | 3.01 | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.76 | 72174.523 | 3583457 |  | 72174.523 |  | 0.01 | 0.01028 | 2.83 | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.76 | 72886.563 | 3650300 |  | 72886.563 |  | 0.01 | 0.01038 | 3.84 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.76 | 73650.867 | 3635436 |  | 73650.867 |  | 0.01 | 0.01049 | 4.93 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.76 | 75305.336 | 3690926 |  | 75305.336 |  | 0.01 | 0.01073 | 7.29 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.76 | 75526.563 | 3681046 |  | 75526.563 |  | 0.01 | 0.01076 | 7.6 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.76 | 74249.516 | 3651160 |  | 74249.516 |  | 0.01 | 0.01058 | 5.78 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.76 | 74140.367 | 3651768 |  | 74140.367 |  | 0.01 | 0.01056 | 5.63 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72727.977 | 3560901 |  | 72727.977 |  | 0.01 | 0.01036 | 3.61 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.76 | 73489.727 | 3628375 |  | 73489.727 |  | 0.01 | 0.01047 | 4.7 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.76 | 72791.852 | 3591268 |  | 72791.852 |  | 0.01 | 0.01037 | 3.7 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.76 | 69266.383 | 3432012 |  | 69266.383 |  | 0.01 | 0.00987 | -1.32 | bb | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.76 | 72812.047 | 3563225 |  | 72812.047 |  | 0.01 | 0.01037 | 3.73 | bb | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.76 | 70597.922 | 3442672 |  | 70597.922 |  | 0.01 | 0.01006 | 0.58 | bb | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.76 | 69640.266 | 3376419 |  | 69640.266 |  | 0.01 | 0.00992 | -0.79 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.76 | 74817.703 | 3688547 |  | 74817.703 |  | 0.01 | 0.01066 | 6.59 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.76 | 67770.406 | 3271291 |  | 67770.406 |  | 0.01 | 0.00966 | -3.45 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.76 | 61003.473 | 2885865 |  | 61003.473 |  | 0.01 | 0.00869 | -13.09 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.76 | 59756.648 | 2784971 |  | 59756.648 |  | 0.01 | 0.00851 | -14.87 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.76 | 57883.109 | 2654882 |  | 57883.109 |  | 0.01 | 0.00825 | -17.54 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.76 | 53039.328 | 2395597 |  | 53039.328 |  | 0.01 | 0.00756 | -24.44 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.76 | 73898.461 | 3579202 |  | 73898.461 |  | 0.01 | 0.01053 | 5.28 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.76 | 86813.352 | 4249834 |  | 86813.352 |  | 0.01 | 0.01237 | 23.68 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.76 | 74901.398 | 3616913 |  | 74901.398 |  | 0.01 | 0.01067 | 6.71 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.76 | 73049.375 | 3523750 |  | 73049.375 |  | 0.01 | 0.01041 | 4.07 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.76 | 77838.516 | 3829249 |  | 77838.516 |  | 0.01 | 0.01109 | 10.89 | MM | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.76 | 75592.711 | 3724434 |  | 75592.711 |  | 0.01 | 0.01077 | 7.7 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.76 | 78736.781 | 3888135 |  | 78736.781 |  | 0.01 | 0.01122 | 12.17 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.76 | 78371.031 | 3866716 |  | 78371.031 |  | 0.01 | 0.01117 | 11.65 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.76 | 74719.867 | 3721994 |  | 74719.867 |  | 0.01 | 0.01065 | 6.45 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.76 | 75418.586 | 3694990 |  | 75418.586 |  | 0.01 | 0.01074 | 7.45 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72627.703 | 3593555 |  | 72627.703 |  | 0.01 | 0.01035 | 3.47 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.76 | 26529.941 | 1325094 |  | 26529.941 |  | 0.01 | 0.00378 | -62.2 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.76 | 75935.922 | 3752660 |  | 75935.922 |  | 0.01 | 0.01082 | 8.18 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.76 | 77973.492 | 3802600 |  | 77973.492 |  | 0.01 | 0.01111 | 11.09 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.76 | 77142.375 | 3761686 |  | 77142.375 |  | 0.01 | 0.01099 | 9.9 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.76 | 76099.219 | 3714412 |  | 76099.219 |  | 0.01 | 0.01084 | 8.42 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.76 | 76403.75 | 3709347 |  | 76403.75 |  | 0.01 | 0.01089 | 8.85 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72442.719 | 3600272 |  | 72442.719 |  | 0.01 | 0.01032 | 3.21 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.76 | 74628.094 | 3664734 |  | 74628.094 |  | 0.01 | 0.01063 | 6.32 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.76 | 79756.063 | 3861355 |  | 79756.063 |  | 0.01 | 0.01136 | 13.63 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.76 | 79699.555 | 3884271 |  | 79699.555 |  | 0.01 | 0.01135 | 13.55 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.76 | 76448.648 | 3750645 |  | 76448.648 |  | 0.01 | 0.01089 | 8.91 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.76 | 74293.961 | 3689391 |  | 74293.961 |  | 0.01 | 0.01058 | 5.84 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.76 | 75961.703 | 3739865 |  | 75961.703 |  | 0.01 | 0.01082 | 8.22 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
Compound 13: M6PFDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.29 | 37792.707 | 1217880 |  | 37792.707 |  | 0.01 | 0.00975 | -2.52 | bb | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.29 | 37339.707 | 1206869 |  | 37339.707 |  | 0.01 | 0.00963 | -3.68 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.29 | 37559.73 | 1208986 |  | 37559.73 |  | 0.01 | 0.00969 | -3.12 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.29 | 38314.09 | 1223482 |  | 38314.09 |  | 0.01 | 0.00988 | -1.17 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.29 | 38319.367 | 1218543 |  | 38319.367 |  | 0.01 | 0.00988 | -1.16 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.29 | 35923.289 | 1140870 |  | 35923.289 |  | 0.01 | 0.00927 | -7.34 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.29 | 36542.281 | 1160598 |  | 36542.281 |  | 0.01 | 0.00943 | -5.74 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.29 | 37978.066 | 1233818 |  | 37978.066 |  | 0.01 | 0.0098 | -2.04 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.29 | 37018.879 | 1176726 |  | 37018.879 |  | 0.01 | 0.00955 | -4.51 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.29 | 34951.309 | 1116667 |  | 34951.309 |  | 0.01 | 0.00902 | -9.85 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.29 | 36311.348 | 1143914 |  | 36311.348 |  | 0.01 | 0.00937 | -6.34 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.29 | 37384.508 | 1185239 |  | 37384.508 |  | 0.01 | 0.00964 | -3.57 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.29 | 37091.914 | 1176282 |  | 37091.914 |  | 0.01 | 0.00957 | -4.32 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.29 | 38336.848 | 1233551 |  | 38336.848 |  | 0.01 | 0.00989 | -1.11 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.29 | 35457.441 | 1116592 |  | 35457.441 |  | 0.01 | 0.00915 | -8.54 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.29 | 33010.672 | 1030582 |  | 33010.672 |  | 0.01 | 0.00851 | -14.85 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.29 | 31009.475 | 962542 |  | 31009.475 |  | 0.01 | 0.008 | -20.01 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.29 | 29707.164 | 916690 |  | 29707.164 |  | 0.01 | 0.00766 | -23.37 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.29 | 27123.66 | 810713 |  | 27123.66 |  | 0.01 | 0.007 | -30.04 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.29 | 38590.461 | 1230275 |  | 38590.461 |  | 0.01 | 0.00995 | -0.46 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.29 | 40285.297 | 1283915 |  | 40285.297 |  | 0.01 | 0.01039 | 3.91 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.29 | 39404.684 | 1244131 |  | 39404.684 |  | 0.01 | 0.01016 | 1.64 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.29 | 36878.289 | 1158022 |  | 36878.289 |  | 0.01 | 0.00951 | -4.87 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.29 | 37417.016 | 1189700 |  | 37417.016 |  | 0.01 | 0.00965 | -3.49 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.29 | 40856.688 | 1297828 |  | 40856.688 |  | 0.01 | 0.01054 | 5.39 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.29 | 39381.18 | 1256250 |  | 39381.18 |  | 0.01 | 0.01016 | 1.58 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.29 | 39899.059 | 1258870 |  | 39899.059 |  | 0.01 | 0.01029 | 2.92 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.29 | 39440.199 | 1263270 |  | 39440.199 |  | 0.01 | 0.01017 | 1.73 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.29 | 39243.359 | 1245149 |  | 39243.359 |  | 0.01 | 0.01012 | 1.23 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.29 | 38771.027 | 1238209 |  | 38771.027 |  | 0.01 | 0.01 | 0.01 | MM | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.29 | 38646.242 | 1237572 |  | 38646.242 |  | 0.01 | 0.00997 | -0.31 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.29 | 42048.52 | 1339339 |  | 42048.52 |  | 0.01 | 0.01085 | 8.46 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.29 | 39681.668 | 1272905 |  | 39681.668 |  | 0.01 | 0.01024 | 2.36 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.29 | 38738.613 | 1223307 |  | 38738.613 |  | 0.01 | 0.00999 | -0.08 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.29 | 37568.254 | 1184707 |  | 37568.254 |  | 0.01 | 0.00969 | -3.09 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.29 | 38877.977 | 1241009 |  | 38877.977 |  | 0.01 | 0.01003 | 0.28 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.29 | 42185.547 | 1327132 |  | 42185.547 |  | 0.01 | 0.01088 | 8.82 | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.29 | 43559.676 | 1400670 |  | 43559.676 |  | 0.01 | 0.01124 | 12.36 | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.29 | 41573.668 | 1324033 |  | 41573.668 |  | 0.01 | 0.01072 | 7.24 | MM | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.29 | 42123.012 | 1375573 |  | 42123.012 |  | 0.01 | 0.01087 | 8.65 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.29 | 41354.645 | 1322921 |  | 41354.645 |  | 0.01 | 0.01067 | 6.67 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.29 | 39860.156 | 1284787 |  | 39860.156 |  | 0.01 | 0.01028 | 2.82 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.29 | 38245.285 | 1242430 |  | 38245.285 |  | 0.01 | 0.00987 | -1.35 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.29 | 40662.246 | 1307754 |  | 40662.246 |  | 0.01 | 0.01049 | 4.89 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.29 | 41766.652 | 1350220 |  | 41766.652 |  | 0.01 | 0.01077 | 7.73 | bb | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.29 | 38594.965 | 1249997 |  | 38594.965 |  | 0.01 | 0.00996 | -0.45 | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.29 | 31312.432 | 979767 |  | 31312.432 |  | 0.01 | 0.00808 | -19.23 | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.29 | 31432.322 | 973478 |  | 31432.322 |  | 0.01 | 0.00811 | -18.92 | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.29 | 31703.865 | 984198 |  | 31703.865 |  | 0.01 | 0.00818 | -18.22 | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.29 | 36568.078 | 1199666 |  | 36568.078 |  | 0.01 | 0.00943 | -5.67 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.29 | 39947.965 | 1290917 |  | 39947.965 |  | 0.01 | 0.0103 | 3.04 | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.29 | 39242.902 | 1266126 |  | 39242.902 |  | 0.01 | 0.01012 | 1.22 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.29 | 37110.707 | 1210382 |  | 37110.707 |  | 0.01 | 0.00957 | -4.28 | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.29 | 38090.129 | 1220383 |  | 38090.129 |  | 0.01 | 0.00983 | -1.75 | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.29 | 39006.359 | 1255961 |  | 39006.359 |  | 0.01 | 0.01006 | 0.61 | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.29 | 36889.391 | 1128492 |  | 36889.391 |  | 0.01 | 0.00952 | -4.85 | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.29 | 33918.5 | 1031045 |  | 33918.5 |  | 0.01 | 0.00875 | -12.51 | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.29 | 33629.922 | 1016596 |  | 33629.922 |  | 0.01 | 0.00867 | -13.25 | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.29 | 39910.887 | 1278369 |  | 39910.887 |  | 0.01 | 0.01029 | 2.95 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.29 | 39164.621 | 1272815 |  | 39164.621 |  | 0.01 | 0.0101 | 1.02 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.29 | 39562.254 | 1298513 |  | 39562.254 |  | 0.01 | 0.0102 | 2.05 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.29 | 41662.426 | 1352565 |  | 41662.426 |  | 0.01 | 0.01075 | 7.47 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.29 | 39324.82 | 1259723 |  | 39324.82 |  | 0.01 | 0.01014 | 1.44 | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.29 | 37500.363 | 1223024 |  | 37500.363 |  | 0.01 | 0.00967 | -3.27 | bb | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.29 | 38443.734 | 1233450 |  | 38443.734 |  | 0.01 | 0.00992 | -0.84 | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.29 | 40001.32 | 1314374 |  | 40001.32 |  | 0.01 | 0.01032 | 3.18 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.29 | 39172.805 | 1248249 |  | 39172.805 |  | 0.01 | 0.0101 | 1.04 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.29 | 40422.34 | 1325262 |  | 40422.34 |  | 0.01 | 0.01043 | 4.27 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.29 | 41193.734 | 1318586 |  | 41193.734 |  | 0.01 | 0.01063 | 6.26 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.29 | 41931.879 | 1343595 |  | 41931.879 |  | 0.01 | 0.01082 | 8.16 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.29 | 40251.609 | 1303796 |  | 40251.609 |  | 0.01 | 0.01038 | 3.83 | bb | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.29 | 39630.223 | 1251803 |  | 39630.223 |  | 0.01 | 0.01022 | 2.22 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.29 | 41927.211 | 1325309 |  | 41927.211 |  | 0.01 | 0.01081 | 8.15 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.29 | 40052.781 | 1281326 |  | 40052.781 |  | 0.01 | 0.01033 | 3.31 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.29 | 40052.25 | 1280612 |  | 40052.25 |  | 0.01 | 0.01033 | 3.31 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.29 | 38973.023 | 1238290 |  | 38973.023 |  | 0.01 | 0.01005 | 0.53 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.29 | 40601.832 | 1289312 |  | 40601.832 |  | 0.01 | 0.01047 | 4.73 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.29 | 40061.527 | 1256952 |  | 40061.527 |  | 0.01 | 0.01033 | 3.34 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.29 | 42510.348 | 1391581 |  | 42510.348 |  | 0.01 | 0.01097 | 9.65 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.29 | 37898.871 | 1191283 |  | 37898.871 |  | 0.01 | 0.00978 | -2.24 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.29 | 33821.711 | 1061806 |  | 33821.711 |  | 0.01 | 0.00872 | -12.76 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.29 | 33647.129 | 1031146 |  | 33647.129 |  | 0.01 | 0.00868 | -13.21 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.29 | 32227.65 | 987816 |  | 32227.65 |  | 0.01 | 0.00831 | -16.87 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.29 | 29370.512 | 880441 |  | 29370.512 |  | 0.01 | 0.00758 | -24.24 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.29 | 40316.699 | 1282533 |  | 40316.699 |  | 0.01 | 0.0104 | 3.99 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.29 | 46858.434 | 1508552 |  | 46858.434 |  | 0.01 | 0.01209 | 20.87 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.29 | 43284.461 | 1398391 |  | 43284.461 |  | 0.01 | 0.01116 | 11.65 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.29 | 40310.613 | 1274189 |  | 40310.613 |  | 0.01 | 0.0104 | 3.98 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.29 | 42210.781 | 1343437 |  | 42210.781 |  | 0.01 | 0.01089 | 8.88 | MM | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.29 | 41901.957 | 1338794 |  | 41901.957 |  | 0.01 | 0.01081 | 8.08 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.29 | 44425.273 | 1445217 |  | 44425.273 |  | 0.01 | 0.01146 | 14.59 | bb | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.29 | 41899.441 | 1344911 |  | 41899.441 |  | 0.01 | 0.01081 | 8.08 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.29 | 42107.133 | 1360262 |  | 42107.133 |  | 0.01 | 0.01086 | 8.61 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.29 | 41405.266 | 1332128 |  | 41405.266 |  | 0.01 | 0.01068 | 6.8 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.29 | 41834.133 | 1337079 |  | 41834.133 |  | 0.01 | 0.01079 | 7.91 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.29 | 14514.265 | 480943 |  | 14514.265 |  | 0.01 | 0.00374 | -62.56 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.29 | 40732.211 | 1304799 |  | 40732.211 |  | 0.01 | 0.01051 | 5.07 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.29 | 43092.957 | 1418734 |  | 43092.957 |  | 0.01 | 0.01112 | 11.16 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.29 | 42587.418 | 1396713 |  | 42587.418 |  | 0.01 | 0.01099 | 9.85 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.29 | 43136.676 | 1396762 |  | 43136.676 |  | 0.01 | 0.01113 | 11.27 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.29 | 39118.82 | 1254442 |  | 39118.82 |  | 0.01 | 0.01009 | 0.9 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.29 | 40277.5 | 1296634 |  | 40277.5 |  | 0.01 | 0.01039 | 3.89 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.29 | 41387.645 | 1353108 |  | 41387.645 |  | 0.01 | 0.01068 | 6.76 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.29 | 45184.34 | 1445972 |  | 45184.34 |  | 0.01 | 0.01166 | 16.55 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.29 | 42725.34 | 1365175 |  | 42725.34 |  | 0.01 | 0.01102 | 10.21 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.29 | 41798.488 | 1364262 |  | 41798.488 |  | 0.01 | 0.01078 | 7.82 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.29 | 41191.141 | 1308029 |  | 41191.141 |  | 0.01 | 0.01063 | 6.25 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.29 | 42275.109 | 1357783 |  | 42275.109 |  | 0.01 | 0.0109 | 9.05 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
Compound 14: M8PFOS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.38 | 3612.967 | 119716 |  | 3612.967 |  | 0.01 | 0.00954 | -4.6 | MM | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3515.653 | 116548 |  | 3515.653 |  | 0.01 | 0.00928 | -7.17 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.39 | 3777.996 | 123463 |  | 3777.996 |  | 0.01 | 0.00998 | -0.24 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3699.702 | 122764 |  | 3699.702 |  | 0.01 | 0.00977 | -2.31 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3810.099 | 126252 |  | 3810.099 |  | 0.01 | 0.01006 | 0.61 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3518.784 | 116780 |  | 3518.784 |  | 0.01 | 0.00929 | -7.08 | bb | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3525.893 | 117373 |  | 3525.893 |  | 0.01 | 0.00931 | -6.9 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.38 | 3792.97 | 124818 |  | 3792.97 |  | 0.01 | 0.01002 | 0.16 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.38 | 3544.214 | 120506 |  | 3544.214 |  | 0.01 | 0.00936 | -6.41 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.39 | 3472.039 | 115185 |  | 3472.039 |  | 0.01 | 0.00917 | -8.32 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.38 | 3420.079 | 112099 |  | 3420.079 |  | 0.01 | 0.00903 | -9.69 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.38 | 3621.232 | 118412 |  | 3621.232 |  | 0.01 | 0.00956 | -4.38 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.38 | 3567.718 | 114506 |  | 3567.718 |  | 0.01 | 0.00942 | -5.79 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.38 | 3638.662 | 120770 |  | 3638.662 |  | 0.01 | 0.00961 | -3.92 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.39 | 3390.249 | 108665 |  | 3390.249 |  | 0.01 | 0.00895 | -10.48 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.39 | 3215.925 | 98917 |  | 3215.925 |  | 0.01 | 0.00849 | -15.08 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.38 | 2816.201 | 87491 |  | 2816.201 |  | 0.01 | 0.00744 | -25.64 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.38 | 2776.837 | 86453 |  | 2776.837 |  | 0.01 | 0.00733 | -26.68 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.38 | 2500.137 | 75944 |  | 2500.137 |  | 0.01 | 0.0066 | -33.98 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.39 | 3600.895 | 118613 |  | 3600.895 |  | 0.01 | 0.00951 | -4.92 | bb | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.38 | 3990.492 | 131836 |  | 3990.492 |  | 0.01 | 0.01054 | 5.37 | bb | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.38 | 3898.205 | 129684 |  | 3898.205 |  | 0.01 | 0.01029 | 2.94 | bb | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.38 | 3650.252 | 120728 |  | 3650.252 |  | 0.01 | 0.00964 | -3.61 | bb | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3625.022 | 121822 |  | 3625.022 |  | 0.01 | 0.00957 | -4.28 | bb | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.38 | 3695.023 | 122492 |  | 3695.023 |  | 0.01 | 0.00976 | -2.43 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3889.072 | 130869 |  | 3889.072 |  | 0.01 | 0.01027 | 2.69 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3767.749 | 123504 |  | 3767.749 |  | 0.01 | 0.00995 | -0.51 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3848.652 | 125059 |  | 3848.652 |  | 0.01 | 0.01016 | 1.63 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3778.806 | 126983 |  | 3778.806 |  | 0.01 | 0.00998 | -0.22 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.38 | 3744.692 | 124131 |  | 3744.692 |  | 0.01 | 0.00989 | -1.12 | MM | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3855.243 | 127878 |  | 3855.243 |  | 0.01 | 0.01018 | 1.8 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.38 | 4039.755 | 137260 |  | 4039.755 |  | 0.01 | 0.01067 | 6.67 | bb | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3796.014 | 125474 |  | 3796.014 |  | 0.01 | 0.01002 | 0.24 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3839.881 | 128310 |  | 3839.881 |  | 0.01 | 0.01014 | 1.4 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3659.798 | 123427 |  | 3659.798 |  | 0.01 | 0.00966 | -3.36 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3857.974 | 129229 |  | 3857.974 |  | 0.01 | 0.01019 | 1.87 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.38 | 3968.314 | 132324 |  | 3968.314 |  | 0.01 | 0.01048 | 4.79 | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.38 | 4331.905 | 143142 |  | 4331.905 |  | 0.01 | 0.01144 | 14.39 | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.38 | 3953.513 | 132694 |  | 3953.513 |  | 0.01 | 0.01044 | 4.4 | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.38 | 4149.038 | 136420 |  | 4149.038 |  | 0.01 | 0.01096 | 9.56 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.38 | 4039.311 | 135371 |  | 4039.311 |  | 0.01 | 0.01067 | 6.66 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.38 | 3948.045 | 134012 |  | 3948.045 |  | 0.01 | 0.01043 | 4.25 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.38 | 3766.95 | 124606 |  | 3766.95 |  | 0.01 | 0.00995 | -0.53 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.38 | 4110.244 | 134922 |  | 4110.244 |  | 0.01 | 0.01085 | 8.53 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.38 | 4134.708 | 137008 |  | 4134.708 |  | 0.01 | 0.01092 | 9.18 | bb | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.38 | 3806.312 | 126396 |  | 3806.312 |  | 0.01 | 0.01005 | 0.51 | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.39 | 3714.114 | 124654 |  | 3714.114 |  | 0.01 | 0.00981 | -1.93 | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.38 | 3808.625 | 124230 |  | 3808.625 |  | 0.01 | 0.01006 | 0.57 | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.38 | 3752.886 | 126474 |  | 3752.886 |  | 0.01 | 0.00991 | -0.9 | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.39 | 3651.146 | 121856 |  | 3651.146 |  | 0.01 | 0.00964 | -3.59 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.38 | 4014.833 | 133225 |  | 4014.833 |  | 0.01 | 0.0106 | 6.01 | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.38 | 3989.71 | 134886 |  | 3989.71 |  | 0.01 | 0.01054 | 5.35 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.38 | 2913.662 | 91176 |  | 2913.662 |  | 0.01 | 0.00769 | -23.06 | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.38 | 2855.056 | 90020 |  | 2855.056 |  | 0.01 | 0.00754 | -24.61 | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.38 | 2979.928 | 97678 |  | 2979.928 |  | 0.01 | 0.00787 | -21.31 | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.38 | 4423.004 | 146329 |  | 4423.004 |  | 0.01 | 0.01168 | 16.79 | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.38 | 3863.019 | 128981 |  | 3863.019 |  | 0.01 | 0.0102 | 2.01 | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.38 | 3884.237 | 128731 |  | 3884.237 |  | 0.01 | 0.01026 | 2.57 | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.38 | 3750.135 | 123730 |  | 3750.135 |  | 0.01 | 0.0099 | -0.97 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.38 | 3862.314 | 129248 |  | 3862.314 |  | 0.01 | 0.0102 | 1.99 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.38 | 3867.536 | 128089 |  | 3867.536 |  | 0.01 | 0.01021 | 2.13 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.39 | 4093.351 | 136601 |  | 4093.351 |  | 0.01 | 0.01081 | 8.09 | MM | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.38 | 3048.485 | 98100 |  | 3048.485 |  | 0.01 | 0.00805 | -19.5 | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.38 | 2956.842 | 92421 |  | 2956.842 |  | 0.01 | 0.00781 | -21.92 | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.38 | 3010.562 | 94271 |  | 3010.562 |  | 0.01 | 0.00795 | -20.5 | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.39 | 4037.25 | 135475 |  | 4037.25 |  | 0.01 | 0.01066 | 6.61 | MM | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3890.618 | 131933 |  | 3890.618 |  | 0.01 | 0.01027 | 2.73 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.38 | 4132.465 | 139020 |  | 4132.465 |  | 0.01 | 0.01091 | 9.12 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.38 | 4043.019 | 135627 |  | 4043.019 |  | 0.01 | 0.01068 | 6.76 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3970.605 | 133154 |  | 3970.605 |  | 0.01 | 0.01048 | 4.85 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3911.216 | 128306 |  | 3911.216 |  | 0.01 | 0.01033 | 3.28 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3887.929 | 128011 |  | 3887.929 |  | 0.01 | 0.01027 | 2.66 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.38 | 4020.333 | 132056 |  | 4020.333 |  | 0.01 | 0.01062 | 6.16 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.38 | 3883.1 | 128767 |  | 3883.1 |  | 0.01 | 0.01025 | 2.54 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.38 | 3891.241 | 125654 |  | 3891.241 |  | 0.01 | 0.01028 | 2.75 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.38 | 3876.859 | 129562 |  | 3876.859 |  | 0.01 | 0.01024 | 2.37 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.38 | 3949.588 | 127568 |  | 3949.588 |  | 0.01 | 0.01043 | 4.29 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.38 | 3972.074 | 130228 |  | 3972.074 |  | 0.01 | 0.01049 | 4.89 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.38 | 4145.597 | 135397 |  | 4145.597 |  | 0.01 | 0.01095 | 9.47 | MM | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.38 | 3703.594 | 118502 |  | 3703.594 |  | 0.01 | 0.00978 | -2.2 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.39 | 3471.925 | 109171 |  | 3471.925 |  | 0.01 | 0.00917 | -8.32 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.38 | 3256.591 | 101282 |  | 3256.591 |  | 0.01 | 0.0086 | -14.01 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.38 | 2948.698 | 91680 |  | 2948.698 |  | 0.01 | 0.00779 | -22.14 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.38 | 2836.917 | 86662 |  | 2836.917 |  | 0.01 | 0.00749 | -25.09 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.38 | 3799.566 | 128031 |  | 3799.566 |  | 0.01 | 0.01003 | 0.33 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.38 | 4658.844 | 155544 |  | 4658.844 |  | 0.01 | 0.0123 | 23.02 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.39 | 4262.813 | 140994 |  | 4262.813 |  | 0.01 | 0.01126 | 12.56 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.39 | 3777.862 | 121730 |  | 3777.862 |  | 0.01 | 0.00998 | -0.24 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.38 | 4047.1 | 136288 |  | 4047.1 |  | 0.01 | 0.01069 | 6.87 | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.39 | 4065.303 | 138492 |  | 4065.303 |  | 0.01 | 0.01073 | 7.35 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4359.24 | 146540 |  | 4359.24 |  | 0.01 | 0.01151 | 15.11 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.38 | 4362.105 | 140819 |  | 4362.105 |  | 0.01 | 0.01152 | 15.18 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4099.792 | 136853 |  | 4099.792 |  | 0.01 | 0.01083 | 8.26 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.38 | 4000.38 | 132484 |  | 4000.38 |  | 0.01 | 0.01056 | 5.63 | bb | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.38 | 4255.308 | 140860 |  | 4255.308 |  | 0.01 | 0.01124 | 12.36 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.39 | 1428.834 | 48845 |  | 1428.834 |  | 0.01 | 0.00377 | -62.27 | MM | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3958.613 | 130634 |  | 3958.613 |  | 0.01 | 0.01045 | 4.53 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4343.202 | 143908 |  | 4343.202 |  | 0.01 | 0.01147 | 14.69 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.39 | 4184.41 | 141733 |  | 4184.41 |  | 0.01 | 0.01105 | 10.49 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4219.208 | 141065 |  | 4219.208 |  | 0.01 | 0.01114 | 11.41 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.39 | 4015.649 | 133746 |  | 4015.649 |  | 0.01 | 0.0106 | 6.04 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3982.236 | 132684 |  | 3982.236 |  | 0.01 | 0.01052 | 5.15 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3983.047 | 131245 |  | 3983.047 |  | 0.01 | 0.01052 | 5.18 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4469.724 | 147677 |  | 4469.724 |  | 0.01 | 0.0118 | 18.03 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.39 | 4314.678 | 143586 |  | 4314.678 |  | 0.01 | 0.01139 | 13.93 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4151.266 | 138285 |  | 4151.266 |  | 0.01 | 0.01096 | 9.62 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.39 | 4116.965 | 134986 |  | 4116.965 |  | 0.01 | 0.01087 | 8.71 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.39 | 4224.377 | 140772 |  | 4224.377 |  | 0.01 | 0.01115 | 11.55 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
Compound 15: 13C6 n-Butylparaben
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.49 | 18857.293 | 505995 |  | 18857.293 |  | 0.01 | 0.00926 | -7.41 | MM | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19211.244 | 512622 |  | 19211.244 |  | 0.01 | 0.00943 | -5.67 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.49 | 19452.498 | 523086 |  | 19452.498 |  | 0.01 | 0.00955 | -4.49 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.49 | 19150.906 | 516256 |  | 19150.906 |  | 0.01 | 0.0094 | -5.97 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.49 | 19353 | 525200 |  | 19353 |  | 0.01 | 0.0095 | -4.97 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.49 | 18654.168 | 506819 |  | 18654.168 |  | 0.01 | 0.00916 | -8.41 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.49 | 18554.959 | 509030 |  | 18554.959 |  | 0.01 | 0.00911 | -8.89 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.49 | 19629.662 | 527429 |  | 19629.662 |  | 0.01 | 0.00964 | -3.62 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.49 | 18885.018 | 512800 |  | 18885.018 |  | 0.01 | 0.00927 | -7.27 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.49 | 18374.332 | 490746 |  | 18374.332 |  | 0.01 | 0.00902 | -9.78 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.49 | 19000.076 | 514764 |  | 19000.076 |  | 0.01 | 0.00933 | -6.71 | MM | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.49 | 20171.609 | 542902 |  | 20171.609 |  | 0.01 | 0.0099 | -0.95 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.49 | 19826.656 | 531010 |  | 19826.656 |  | 0.01 | 0.00974 | -2.65 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.49 | 19474.645 | 532376 |  | 19474.645 |  | 0.01 | 0.00956 | -4.38 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.49 | 19678.5 | 525681 |  | 19678.5 |  | 0.01 | 0.00966 | -3.38 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.49 | 18895.691 | 511302 |  | 18895.691 |  | 0.01 | 0.00928 | -7.22 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.49 | 18753.004 | 501592 |  | 18753.004 |  | 0.01 | 0.00921 | -7.92 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.49 | 19034.314 | 503536 |  | 19034.314 |  | 0.01 | 0.00935 | -6.54 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.49 | 19099.988 | 510694 |  | 19099.988 |  | 0.01 | 0.00938 | -6.22 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.49 | 19263.697 | 517174 |  | 19263.697 |  | 0.01 | 0.00946 | -5.41 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.49 | 20758.027 | 568945 |  | 20758.027 |  | 0.01 | 0.01019 | 1.92 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.49 | 20168.148 | 541891 |  | 20168.148 |  | 0.01 | 0.0099 | -0.97 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.49 | 19404.105 | 532543 |  | 19404.105 |  | 0.01 | 0.00953 | -4.72 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19356.43 | 523968 |  | 19356.43 |  | 0.01 | 0.0095 | -4.96 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.49 | 20514.73 | 557477 |  | 20514.73 |  | 0.01 | 0.01007 | 0.73 | MM | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.49 | 20179.088 | 541603 |  | 20179.088 |  | 0.01 | 0.00991 | -0.92 | MM | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.49 | 20475.318 | 551582 |  | 20475.318 |  | 0.01 | 0.01005 | 0.54 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.49 | 19498.076 | 523460 |  | 19498.076 |  | 0.01 | 0.00957 | -4.26 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.49 | 19607.979 | 524201 |  | 19607.979 |  | 0.01 | 0.00963 | -3.72 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.49 | 19650.477 | 527645 |  | 19650.477 |  | 0.01 | 0.00965 | -3.51 | MM | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19171.268 | 515644 |  | 19171.268 |  | 0.01 | 0.00941 | -5.87 | MM | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21051.086 | 566469 |  | 21051.086 |  | 0.01 | 0.01034 | 3.36 | MM | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.49 | 20352.199 | 553407 |  | 20352.199 |  | 0.01 | 0.00999 | -0.07 | MM | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.49 | 19914.854 | 535562 |  | 19914.854 |  | 0.01 | 0.00978 | -2.22 | MM | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.49 | 19468.383 | 523251 |  | 19468.383 |  | 0.01 | 0.00956 | -4.41 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20120.816 | 542770 |  | 20120.816 |  | 0.01 | 0.00988 | -1.2 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.49 | 21697.566 | 586996 |  | 21697.566 |  | 0.01 | 0.01065 | 6.54 | MM | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.49 | 21829.346 | 586073 |  | 21829.346 |  | 0.01 | 0.01072 | 7.19 | MM | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.49 | 20831.814 | 561981 |  | 20831.814 |  | 0.01 | 0.01023 | 2.29 | MM | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.49 | 21105.879 | 570134 |  | 21105.879 |  | 0.01 | 0.01036 | 3.63 | MM | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.49 | 20513.939 | 559055 |  | 20513.939 |  | 0.01 | 0.01007 | 0.73 | MM | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.49 | 20742.979 | 560428 |  | 20742.979 |  | 0.01 | 0.01019 | 1.85 | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.49 | 20304.375 | 543610 |  | 20304.375 |  | 0.01 | 0.00997 | -0.3 | MM | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.49 | 21071.273 | 570517 |  | 21071.273 |  | 0.01 | 0.01035 | 3.46 | MM | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.49 | 20743.402 | 573332 |  | 20743.402 |  | 0.01 | 0.01019 | 1.85 | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.49 | 20386.736 | 551776 |  | 20386.736 |  | 0.01 | 0.01001 | 0.1 | MM | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.49 | 19113.52 | 511307 |  | 19113.52 |  | 0.01 | 0.00939 | -6.15 | MM | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.49 | 19075.039 | 510117 |  | 19075.039 |  | 0.01 | 0.00937 | -6.34 | MM | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.49 | 19599.061 | 534441 |  | 19599.061 |  | 0.01 | 0.00962 | -3.77 | MM | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.49 | 18590.02 | 500460 |  | 18590.02 |  | 0.01 | 0.00913 | -8.72 | MM | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.49 | 20187.695 | 544215 |  | 20187.695 |  | 0.01 | 0.00991 | -0.88 | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.49 | 19809.779 | 537260 |  | 19809.779 |  | 0.01 | 0.00973 | -2.73 | MM | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.49 | 19047.59 | 510104 |  | 19047.59 |  | 0.01 | 0.00935 | -6.47 | MM | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.49 | 19612.295 | 525660 |  | 19612.295 |  | 0.01 | 0.00963 | -3.7 | MM | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.49 | 19736.887 | 531419 |  | 19736.887 |  | 0.01 | 0.00969 | -3.09 | MM | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.49 | 22790.889 | 620955 |  | 22790.889 |  | 0.01 | 0.01119 | 11.91 | MM | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.49 | 20082.863 | 540143 |  | 20082.863 |  | 0.01 | 0.00986 | -1.39 | MM | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.49 | 20098.73 | 550011 |  | 20098.73 |  | 0.01 | 0.00987 | -1.31 | MM | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.49 | 19749.709 | 542271 |  | 19749.709 |  | 0.01 | 0.0097 | -3.03 | MM | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.49 | 19920.66 | 541444 |  | 19920.66 |  | 0.01 | 0.00978 | -2.19 | MM | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.49 | 19372.84 | 527096 |  | 19372.84 |  | 0.01 | 0.00951 | -4.88 | MM | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.49 | 20612.568 | 553657 |  | 20612.568 |  | 0.01 | 0.01012 | 1.21 | MM | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.49 | 20799.916 | 564585 |  | 20799.916 |  | 0.01 | 0.01021 | 2.13 | MM | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.49 | 20129.732 | 544708 |  | 20129.732 |  | 0.01 | 0.00988 | -1.16 | MM | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.49 | 20311.908 | 546886 |  | 20311.908 |  | 0.01 | 0.00997 | -0.27 | MM | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.49 | 20241.559 | 560304 |  | 20241.559 |  | 0.01 | 0.00994 | -0.61 | MM | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20027.883 | 552934 |  | 20027.883 |  | 0.01 | 0.00983 | -1.66 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21530.92 | 588493 |  | 21530.92 |  | 0.01 | 0.01057 | 5.72 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.49 | 21003.221 | 577334 |  | 21003.221 |  | 0.01 | 0.01031 | 3.13 | MM | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21356.199 | 581177 |  | 21356.199 |  | 0.01 | 0.01049 | 4.86 | MM | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.49 | 20342.879 | 549242 |  | 20342.879 |  | 0.01 | 0.00999 | -0.11 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20303.764 | 551555 |  | 20303.764 |  | 0.01 | 0.00997 | -0.31 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.49 | 20294.021 | 556258 |  | 20294.021 |  | 0.01 | 0.00996 | -0.35 | MM | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.49 | 19831.742 | 541739 |  | 19831.742 |  | 0.01 | 0.00974 | -2.62 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.49 | 20010.529 | 544568 |  | 20010.529 |  | 0.01 | 0.00983 | -1.75 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.49 | 20213.203 | 557415 |  | 20213.203 |  | 0.01 | 0.00992 | -0.75 | MM | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.49 | 21640.061 | 585142 |  | 21640.061 |  | 0.01 | 0.01063 | 6.26 | MM | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.49 | 21196.979 | 574966 |  | 21196.979 |  | 0.01 | 0.01041 | 4.08 | MM | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.49 | 21380.4 | 582255 |  | 21380.4 |  | 0.01 | 0.0105 | 4.98 | MM | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.49 | 20941.443 | 568394 |  | 20941.443 |  | 0.01 | 0.01028 | 2.83 | MM | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.49 | 20097.592 | 543063 |  | 20097.592 |  | 0.01 | 0.00987 | -1.32 | MM | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.49 | 20623.277 | 568310 |  | 20623.277 |  | 0.01 | 0.01013 | 1.26 | MM | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.49 | 19677.9 | 548965 |  | 19677.9 |  | 0.01 | 0.00966 | -3.38 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.49 | 20306.045 | 546880 |  | 20306.045 |  | 0.01 | 0.00997 | -0.29 | MM | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.49 | 20563.285 | 569601 |  | 20563.285 |  | 0.01 | 0.0101 | 0.97 | MM | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.49 | 24810.557 | 681719 |  | 24810.557 |  | 0.01 | 0.01218 | 21.82 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.49 | 21870.688 | 595097 |  | 21870.688 |  | 0.01 | 0.01074 | 7.39 | MM | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.49 | 21461.814 | 589347 |  | 21461.814 |  | 0.01 | 0.01054 | 5.38 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.49 | 21765.705 | 595826 |  | 21765.705 |  | 0.01 | 0.01069 | 6.87 | MM | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.49 | 21216.219 | 576033 |  | 21216.219 |  | 0.01 | 0.01042 | 4.17 | MM | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21769.814 | 595623 |  | 21769.814 |  | 0.01 | 0.01069 | 6.89 | MM | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.49 | 21732.258 | 586430 |  | 21732.258 |  | 0.01 | 0.01067 | 6.71 | MM | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21173.01 | 577698 |  | 21173.01 |  | 0.01 | 0.0104 | 3.96 | MM | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21096.451 | 577658 |  | 21096.451 |  | 0.01 | 0.01036 | 3.59 | MM | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20774.008 | 565337 |  | 20774.008 |  | 0.01 | 0.0102 | 2 | MM | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.49 | 7133.3 | 196191 |  | 7133.3 |  | 0.01 | 0.0035 | -64.97 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20622.34 | 561597 |  | 20622.34 |  | 0.01 | 0.01013 | 1.26 | MM | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.49 | 22279.238 | 606721 |  | 22279.238 |  | 0.01 | 0.01094 | 9.39 | MM | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.49 | 22107.174 | 605748 |  | 22107.174 |  | 0.01 | 0.01085 | 8.55 | MM | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.49 | 22507.645 | 615348 |  | 22507.645 |  | 0.01 | 0.01105 | 10.52 | MM | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21428.309 | 584103 |  | 21428.309 |  | 0.01 | 0.01052 | 5.22 | MM | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.49 | 21171.861 | 573511 |  | 21171.861 |  | 0.01 | 0.0104 | 3.96 | MM | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20838.236 | 563891 |  | 20838.236 |  | 0.01 | 0.01023 | 2.32 | MM | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.49 | 22277.256 | 608110 |  | 22277.256 |  | 0.01 | 0.01094 | 9.38 | MM | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.49 | 22667.871 | 620313 |  | 22667.871 |  | 0.01 | 0.01113 | 11.3 | MM | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21603.74 | 592242 |  | 21603.74 |  | 0.01 | 0.01061 | 6.08 | MM | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21934.615 | 600256 |  | 21934.615 |  | 0.01 | 0.01077 | 7.7 | MM | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.49 | 21390.625 | 584158 |  | 21390.625 |  | 0.01 | 0.0105 | 5.03 | MM | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
Compound 16: M7PFUdA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.58 | 31803.217 | 947983 |  | 31803.217 |  | 0.01 | 0.0091 | -9.01 | MM | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.58 | 32447.529 | 974791 |  | 32447.529 |  | 0.01 | 0.00928 | -7.17 | MM | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.58 | 32267.967 | 986301 |  | 32267.967 |  | 0.01 | 0.00923 | -7.68 | MM | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.58 | 33209.082 | 973811 |  | 33209.082 |  | 0.01 | 0.0095 | -4.99 | MM | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.58 | 33202.188 | 1009091 |  | 33202.188 |  | 0.01 | 0.0095 | -5.01 | MM | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.58 | 32022.859 | 951242 |  | 32022.859 |  | 0.01 | 0.00916 | -8.38 | MM | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.58 | 31705.016 | 963116 |  | 31705.016 |  | 0.01 | 0.00907 | -9.29 | MM | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.58 | 34305.039 | 1017728 |  | 34305.039 |  | 0.01 | 0.00981 | -1.85 | MM | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.58 | 33209.145 | 983289 |  | 33209.145 |  | 0.01 | 0.0095 | -4.99 | MM | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.58 | 30390.211 | 927562 |  | 30390.211 |  | 0.01 | 0.00869 | -13.05 | MM | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.58 | 32738.229 | 977888 |  | 32738.229 |  | 0.01 | 0.00937 | -6.33 | bb | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.58 | 34820.98 | 1040287 |  | 34820.98 |  | 0.01 | 0.00996 | -0.38 | MM | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.58 | 35522.934 | 1047819 |  | 35522.934 |  | 0.01 | 0.01016 | 1.63 | MM | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.58 | 32830.207 | 975030 |  | 32830.207 |  | 0.01 | 0.00939 | -6.07 | MM | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.58 | 34067.25 | 1018919 |  | 34067.25 |  | 0.01 | 0.00975 | -2.53 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.58 | 34067.969 | 1023156 |  | 34067.969 |  | 0.01 | 0.00975 | -2.53 | MM | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.58 | 32972.277 | 988333 |  | 32972.277 |  | 0.01 | 0.00943 | -5.67 | MM | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.58 | 33862.289 | 1002519 |  | 33862.289 |  | 0.01 | 0.00969 | -3.12 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.58 | 34733.418 | 1034220 |  | 34733.418 |  | 0.01 | 0.00994 | -0.63 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.58 | 32111.395 | 978162 |  | 32111.395 |  | 0.01 | 0.00919 | -8.13 | MM | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.58 | 35946.789 | 1062522 |  | 35946.789 |  | 0.01 | 0.01028 | 2.85 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.58 | 34698.945 | 1020781 |  | 34698.945 |  | 0.01 | 0.00993 | -0.72 | MM | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.58 | 35005.285 | 1033252 |  | 35005.285 |  | 0.01 | 0.01002 | 0.15 | MM | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.58 | 32869.164 | 971478 |  | 32869.164 |  | 0.01 | 0.0094 | -5.96 | MM | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.58 | 33731.746 | 994147 |  | 33731.746 |  | 0.01 | 0.00965 | -3.49 | bb | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.58 | 35506.813 | 1053187 |  | 35506.813 |  | 0.01 | 0.01016 | 1.59 | bb | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.58 | 35507.883 | 1037444 |  | 35507.883 |  | 0.01 | 0.01016 | 1.59 | MM | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33292.348 | 989585 |  | 33292.348 |  | 0.01 | 0.00953 | -4.75 | MM | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.58 | 33909.344 | 1001427 |  | 33909.344 |  | 0.01 | 0.0097 | -2.98 | MM | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.58 | 33975.277 | 1013740 |  | 33975.277 |  | 0.01 | 0.00972 | -2.8 | bb | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.58 | 34714.773 | 1029571 |  | 34714.773 |  | 0.01 | 0.00993 | -0.68 | bb | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.58 | 36595.254 | 1075579 |  | 36595.254 |  | 0.01 | 0.01047 | 4.7 | bb | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.58 | 34187.922 | 1004147 |  | 34187.922 |  | 0.01 | 0.00978 | -2.19 | bb | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.58 | 34225.648 | 1002659 |  | 34225.648 |  | 0.01 | 0.00979 | -2.08 | bb | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33805.305 | 980715 |  | 33805.305 |  | 0.01 | 0.00967 | -3.28 | MM | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.58 | 33640.234 | 1009063 |  | 33640.234 |  | 0.01 | 0.00962 | -3.75 | MM | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.58 | 35815.477 | 1046068 |  | 35815.477 |  | 0.01 | 0.01025 | 2.47 | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.58 | 38437.145 | 1124788 |  | 38437.145 |  | 0.01 | 0.011 | 9.97 | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.58 | 36556.488 | 1079330 |  | 36556.488 |  | 0.01 | 0.01046 | 4.59 | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.58 | 37541.629 | 1112351 |  | 37541.629 |  | 0.01 | 0.01074 | 7.41 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.58 | 35548.031 | 1050255 |  | 35548.031 |  | 0.01 | 0.01017 | 1.7 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.58 | 35730.238 | 1047871 |  | 35730.238 |  | 0.01 | 0.01022 | 2.23 | MM | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.58 | 35277.617 | 1031353 |  | 35277.617 |  | 0.01 | 0.01009 | 0.93 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.58 | 37072.477 | 1098487 |  | 37072.477 |  | 0.01 | 0.01061 | 6.07 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.58 | 38052.293 | 1127983 |  | 38052.293 |  | 0.01 | 0.01089 | 8.87 | MM | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.58 | 34841.289 | 1025977 |  | 34841.289 |  | 0.01 | 0.00997 | -0.32 | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.58 | 33680.504 | 1011129 |  | 33680.504 |  | 0.01 | 0.00964 | -3.64 | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.58 | 33313.039 | 994208 |  | 33313.039 |  | 0.01 | 0.00953 | -4.69 | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.58 | 35218.301 | 1035739 |  | 35218.301 |  | 0.01 | 0.01008 | 0.76 | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.58 | 31820.521 | 953848 |  | 31820.521 |  | 0.01 | 0.0091 | -8.96 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.58 | 35913.09 | 1059917 |  | 35913.09 |  | 0.01 | 0.01027 | 2.75 | MM | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.58 | 34425.738 | 1012072 |  | 34425.738 |  | 0.01 | 0.00985 | -1.51 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.58 | 33026.402 | 977107 |  | 33026.402 |  | 0.01 | 0.00945 | -5.51 | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.58 | 33342.656 | 982047 |  | 33342.656 |  | 0.01 | 0.00954 | -4.61 | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.58 | 35495.684 | 1043413 |  | 35495.684 |  | 0.01 | 0.01016 | 1.55 | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.58 | 41065.688 | 1199133 |  | 41065.688 |  | 0.01 | 0.01175 | 17.49 | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.58 | 36334.883 | 1074823 |  | 36334.883 |  | 0.01 | 0.0104 | 3.96 | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.58 | 33962.336 | 1005645 |  | 33962.336 |  | 0.01 | 0.00972 | -2.83 | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.58 | 33873.305 | 1004033 |  | 33873.305 |  | 0.01 | 0.00969 | -3.09 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.58 | 34683.414 | 1029481 |  | 34683.414 |  | 0.01 | 0.00992 | -0.77 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.58 | 34392.777 | 1023789 |  | 34392.777 |  | 0.01 | 0.00984 | -1.6 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.58 | 35834.434 | 1074706 |  | 35834.434 |  | 0.01 | 0.01025 | 2.52 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.58 | 34104.559 | 997654 |  | 34104.559 |  | 0.01 | 0.00976 | -2.43 | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.58 | 33587.531 | 992023 |  | 33587.531 |  | 0.01 | 0.00961 | -3.9 | bb | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.58 | 34354.273 | 1023908 |  | 34354.273 |  | 0.01 | 0.00983 | -1.71 | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.58 | 36040.789 | 1079855 |  | 36040.789 |  | 0.01 | 0.01031 | 3.11 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.58 | 33756.629 | 983178 |  | 33756.629 |  | 0.01 | 0.00966 | -3.42 | MM | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.58 | 36770.707 | 1084336 |  | 36770.707 |  | 0.01 | 0.01052 | 5.2 | MM | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.58 | 36297.711 | 1072373 |  | 36297.711 |  | 0.01 | 0.01038 | 3.85 | bb | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36058.262 | 1061800 |  | 36058.262 |  | 0.01 | 0.01032 | 3.16 | bb | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33890.82 | 986507 |  | 33890.82 |  | 0.01 | 0.0097 | -3.04 | MM | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.58 | 34649.969 | 1022090 |  | 34649.969 |  | 0.01 | 0.00991 | -0.87 | MM | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.58 | 36403.238 | 1080267 |  | 36403.238 |  | 0.01 | 0.01042 | 4.15 | bb | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.58 | 35900.168 | 1066387 |  | 35900.168 |  | 0.01 | 0.01027 | 2.71 | MM | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.58 | 33082.625 | 981951 |  | 33082.625 |  | 0.01 | 0.00947 | -5.35 | MM | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.58 | 34432.25 | 1023852 |  | 34432.25 |  | 0.01 | 0.00985 | -1.49 | bb | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.58 | 36870.191 | 1095480 |  | 36870.191 |  | 0.01 | 0.01055 | 5.49 | bb | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.58 | 37336.027 | 1108892 |  | 37336.027 |  | 0.01 | 0.01068 | 6.82 | bb | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.58 | 37023.516 | 1096703 |  | 37023.516 |  | 0.01 | 0.01059 | 5.93 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.58 | 36169.313 | 1077802 |  | 36169.313 |  | 0.01 | 0.01035 | 3.48 | bb | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.58 | 34693.082 | 1046635 |  | 34693.082 |  | 0.01 | 0.00993 | -0.74 | bb | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.58 | 36845.703 | 1098532 |  | 36845.703 |  | 0.01 | 0.01054 | 5.42 | bb | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.58 | 36388.051 | 1073869 |  | 36388.051 |  | 0.01 | 0.01041 | 4.11 | MM | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.58 | 35195.254 | 1041674 |  | 35195.254 |  | 0.01 | 0.01007 | 0.69 | bb | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.58 | 35430.559 | 1047127 |  | 35430.559 |  | 0.01 | 0.01014 | 1.37 | bb | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.58 | 39598.043 | 1175237 |  | 39598.043 |  | 0.01 | 0.01133 | 13.29 | MM | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.58 | 38674.035 | 1156652 |  | 38674.035 |  | 0.01 | 0.01106 | 10.65 | bb | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.58 | 35777.035 | 1055289 |  | 35777.035 |  | 0.01 | 0.01024 | 2.36 | MM | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.58 | 37009.66 | 1102895 |  | 37009.66 |  | 0.01 | 0.01059 | 5.89 | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.58 | 36119.102 | 1081464 |  | 36119.102 |  | 0.01 | 0.01033 | 3.34 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37791.871 | 1131169 |  | 37791.871 |  | 0.01 | 0.01081 | 8.12 | bb | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.58 | 37101.148 | 1109658 |  | 37101.148 |  | 0.01 | 0.01061 | 6.15 | bb | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36060.398 | 1072029 |  | 36060.398 |  | 0.01 | 0.01032 | 3.17 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.58 | 35058.344 | 1046315 |  | 35058.344 |  | 0.01 | 0.01003 | 0.3 | bb | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.58 | 35763.523 | 1070814 |  | 35763.523 |  | 0.01 | 0.01023 | 2.32 | bb | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.58 | 12664.324 | 375482 |  | 12664.324 |  | 0.01 | 0.00362 | -63.77 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.58 | 35462.613 | 1070005 |  | 35462.613 |  | 0.01 | 0.01015 | 1.46 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37730.301 | 1124057 |  | 37730.301 |  | 0.01 | 0.01079 | 7.95 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.58 | 37717.949 | 1137986 |  | 37717.949 |  | 0.01 | 0.01079 | 7.91 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36561.789 | 1080162 |  | 36561.789 |  | 0.01 | 0.01046 | 4.6 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.58 | 34807.703 | 1039396 |  | 34807.703 |  | 0.01 | 0.00996 | -0.41 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.58 | 35477.816 | 1062062 |  | 35477.816 |  | 0.01 | 0.01015 | 1.5 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.58 | 35970.531 | 1074156 |  | 35970.531 |  | 0.01 | 0.01029 | 2.91 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37755.715 | 1135042 |  | 37755.715 |  | 0.01 | 0.0108 | 8.02 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.58 | 36696.27 | 1092910 |  | 36696.27 |  | 0.01 | 0.0105 | 4.99 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36844.078 | 1099248 |  | 36844.078 |  | 0.01 | 0.01054 | 5.41 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.58 | 35534.512 | 1064559 |  | 35534.512 |  | 0.01 | 0.01017 | 1.67 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.58 | 36225.75 | 1082349 |  | 36225.75 |  | 0.01 | 0.01036 | 3.64 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
Compound 17: M2PFTeDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
2 | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
3 | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
4 | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 |  |  |  | 2020-11-23
5 | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.48 | 10173.215 | 343416 |  | 10173.215 |  | 0.01 | 0.00979 | -2.12 | bb | 2020-11-23
6 | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
7 | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9730.571 | 329079 |  | 9730.571 |  | 0.01 | 0.00936 | -6.38 | bb | 2020-11-23
8 | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11120.537 | 373820 |  | 11120.537 |  | 0.01 | 0.0107 | 7 | bb | 2020-11-23
9 | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10876.152 | 369500 |  | 10876.152 |  | 0.01 | 0.01046 | 4.65 | bb | 2020-11-23
10 | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10360.407 | 348348 |  | 10360.407 |  | 0.01 | 0.00997 | -0.32 | bb | 2020-11-23
11 | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10282.522 | 349144 |  | 10282.522 |  | 0.01 | 0.00989 | -1.07 | bb | 2020-11-23
12 | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10162.122 | 345524 |  | 10162.122 |  | 0.01 | 0.00978 | -2.22 | bb | 2020-11-23
13 | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
14 | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.48 | 10490.147 | 354799 |  | 10490.147 |  | 0.01 | 0.01009 | 0.93 | bb | 2020-11-23
15 | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9708.633 | 329742 |  | 9708.633 |  | 0.01 | 0.00934 | -6.59 | bb | 2020-11-23
16 | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.48 | 9783.861 | 332235 |  | 9783.861 |  | 0.01 | 0.00941 | -5.86 | bb | 2020-11-23
17 | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.48 | 11453.623 | 387651 |  | 11453.623 |  | 0.01 | 0.01102 | 10.2 | bb | 2020-11-23
18 | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.48 | 10865.012 | 367049 |  | 10865.012 |  | 0.01 | 0.01045 | 4.54 | bb | 2020-11-23
19 | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.48 | 11101.751 | 374089 |  | 11101.751 |  | 0.01 | 0.01068 | 6.82 | bb | 2020-11-23
20 | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.48 | 11173.664 | 376859 |  | 11173.664 |  | 0.01 | 0.01075 | 7.51 | bb | 2020-11-23
21 | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10508.111 | 351061 |  | 10508.111 |  | 0.01 | 0.01011 | 1.1 | MM | 2020-11-23
22 | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.48 | 9989.264 | 334986 |  | 9989.264 |  | 0.01 | 0.00961 | -3.89 | bb | 2020-11-23
23 | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10623.775 | 354798 |  | 10623.775 |  | 0.01 | 0.01022 | 2.22 | bb | 2020-11-23
24 | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.48 | 10923.09 | 361685 |  | 10923.09 |  | 0.01 | 0.01051 | 5.1 | MM | 2020-11-23
25 | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9730.678 | 316707 |  | 9730.678 |  | 0.01 | 0.00936 | -6.38 | MM | 2020-11-23
26 | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
27 | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.48 | 10475.948 | 355039 |  | 10475.948 |  | 0.01 | 0.01008 | 0.8 | bb | 2020-11-23
28 | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.48 | 11635.099 | 388325 |  | 11635.099 |  | 0.01 | 0.01119 | 11.95 | MM | 2020-11-23
29 | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.48 | 12150.41 | 410208 |  | 12150.41 |  | 0.01 | 0.01169 | 16.91 | bb | 2020-11-23
30 | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.48 | 10343.043 | 345742 |  | 10343.043 |  | 0.01 | 0.00995 | -0.48 | bb | 2020-11-23
31 | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
32 | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11255 | 376966 |  | 11255 |  | 0.01 | 0.01083 | 8.29 | bb | 2020-11-23
33 | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.48 | 12190.335 | 409431 |  | 12190.335 |  | 0.01 | 0.01173 | 17.29 | bb | 2020-11-23
34 | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11718.064 | 392785 |  | 11718.064 |  | 0.01 | 0.01127 | 12.75 | bb | 2020-11-23
35 | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11317.867 | 377809 |  | 11317.867 |  | 0.01 | 0.01089 | 8.9 | bb | 2020-11-23
36 | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10528.761 | 357528 |  | 10528.761 |  | 0.01 | 0.01013 | 1.3 | bb | 2020-11-23
37 | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.48 | 11242.785 | 374717 |  | 11242.785 |  | 0.01 | 0.01082 | 8.17 | bb | 2020-11-23
38 | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.48 | 12048.315 | 404321 |  | 12048.315 |  | 0.01 | 0.01159 | 15.92 | bb | 2020-11-23
39 | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11228.034 | 376645 |  | 11228.034 |  | 0.01 | 0.0108 | 8.03 | bb | 2020-11-23
40 | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11829.166 | 397464 |  | 11829.166 |  | 0.01 | 0.01138 | 13.82 | bb | 2020-11-23
41 | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11542.465 | 386368 |  | 11542.465 |  | 0.01 | 0.01111 | 11.06 | bb | 2020-11-23
42 | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11854.473 | 394417 |  | 11854.473 |  | 0.01 | 0.01141 | 14.06 | bb | 2020-11-23
43 | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10775.113 | 363685 |  | 10775.113 |  | 0.01 | 0.01037 | 3.67 | bb | 2020-11-23
44 | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10874.418 | 363355 |  | 10874.418 |  | 0.01 | 0.01046 | 4.63 | bb | 2020-11-23
45 | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
46 | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.48 | 11440.534 | 382724 |  | 11440.534 |  | 0.01 | 0.01101 | 10.08 | bb | 2020-11-23
47 | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.48 | 13256.198 | 439493 |  | 13256.198 |  | 0.01 | 0.01275 | 27.55 | bb | 2020-11-23
48 | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 3.48 | 11719.4 | 391810 |  | 11719.4 |  | 0.01 | 0.01128 | 12.76 | bb | 2020-11-23
49 | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.48 | 11316.626 | 377070 |  | 11316.626 |  | 0.01 | 0.01089 | 8.88 | bb | 2020-11-23
50 | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.48 | 11121.704 | 373796 |  | 11121.704 |  | 0.01 | 0.0107 | 7.01 | bb | 2020-11-23
51 | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.48 | 11384.287 | 381386 |  | 11384.287 |  | 0.01 | 0.01095 | 9.53 | bb | 2020-11-23
52 | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.48 | 11975.843 | 401621 |  | 11975.843 |  | 0.01 | 0.01152 | 15.23 | bb | 2020-11-23
53 | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.48 | 11676.331 | 390146 |  | 11676.331 |  | 0.01 | 0.01123 | 12.34 | bb | 2020-11-23
54 | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.48 | 11894.549 | 401063 |  | 11894.549 |  | 0.01 | 0.01144 | 14.44 | bb | 2020-11-23
55 | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.48 | 11048.28 | 371884 |  | 11048.28 |  | 0.01 | 0.01063 | 6.3 | bb | 2020-11-23
56 | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.48 | 9862.112 | 325453 |  | 9862.112 |  | 0.01 | 0.00949 | -5.11 | bb | 2020-11-23
57 | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.48 | 9161.524 | 302491 |  | 9161.524 |  | 0.01 | 0.00881 | -11.85 | bb | 2020-11-23
58 | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.48 | 9049.197 | 303067 |  | 9049.197 |  | 0.01 | 0.00871 | -12.93 | bb | 2020-11-23
59 | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-23
60 | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.48 | 8897.753 | 300237 |  | 8897.753 |  | 0.01 | 0.00856 | -14.39 | bb | 2020-11-23
61 | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.48 | 10825.349 | 365981 |  | 10825.349 |  | 0.01 | 0.01042 | 4.16 | bb | 2020-11-23
62 | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.48 | 9994.433 | 339009 |  | 9994.433 |  | 0.01 | 0.00962 | -3.84 | bb | 2020-11-23
63 | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.48 | 8190.83 | 272286 |  | 8190.83 |  | 0.01 | 0.00788 | -21.19 | bb | 2020-11-23
64 | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.48 | 9179.411 | 308907 |  | 9179.411 |  | 0.01 | 0.00883 | -11.68 | bb | 2020-11-23
65 | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.48 | 10046.605 | 340119 |  | 10046.605 |  | 0.01 | 0.00967 | -3.34 | bb | 2020-11-23
66 | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
67 | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.48 | 11574.24 | 389337 |  | 11574.24 |  | 0.01 | 0.01114 | 11.36 | bb | 2020-11-23
68 | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.48 | 8454.886 | 279574 |  | 8454.886 |  | 0.01 | 0.00813 | -18.65 | bb | 2020-11-23
69 | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.48 | 8997.701 | 299688 |  | 8997.701 |  | 0.01 | 0.00866 | -13.43 | bb | 2020-11-23
70 | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.48 | 9464.119 | 317767 |  | 9464.119 |  | 0.01 | 0.00911 | -8.94 | bb | 2020-11-23
71 | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.48 | 8344.308 | 280834 |  | 8344.308 |  | 0.01 | 0.00803 | -19.71 | bb | 2020-11-23
72 | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.48 | 8994.774 | 300266 |  | 8994.774 |  | 0.01 | 0.00865 | -13.46 | bb | 2020-11-23
73 | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.48 | 10773.584 | 362156 |  | 10773.584 |  | 0.01 | 0.01037 | 3.66 | bb | 2020-11-23
74 | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.48 | 9773.682 | 330647 |  | 9773.682 |  | 0.01 | 0.0094 | -5.96 | bb | 2020-11-23
75 | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.48 | 9471.696 | 319932 |  | 9471.696 |  | 0.01 | 0.00911 | -8.87 | bb | 2020-11-23
76 | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.48 | 9756.347 | 327921 |  | 9756.347 |  | 0.01 | 0.00939 | -6.13 | bb | 2020-11-23
77 | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 3.48 | 10459.238 | 354868 |  | 10459.238 |  | 0.01 | 0.01006 | 0.63 | bb | 2020-11-23
78 | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-23
79 | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.48 | 10426.968 | 352207 |  | 10426.968 |  | 0.01 | 0.01003 | 0.32 | bb | 2020-11-23
80 | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11531.981 | 391568 |  | 11531.981 |  | 0.01 | 0.0111 | 10.96 | bb | 2020-11-23
81 | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10908.972 | 365733 |  | 10908.972 |  | 0.01 | 0.0105 | 4.96 | bb | 2020-11-23
82 | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10369.82 | 347172 |  | 10369.82 |  | 0.01 | 0.00998 | -0.23 | bb | 2020-11-23
83 | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9679.873 | 327049 |  | 9679.873 |  | 0.01 | 0.00931 | -6.86 | bb | 2020-11-23
84 | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10265.066 | 348695 |  | 10265.066 |  | 0.01 | 0.00988 | -1.23 | bb | 2020-11-23
85 | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 3.49 | 3.524 | 114 |  | 3.524 |  | 0.01 | 0 | -99.97 | bb | 2020-11-23
86 | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.48 | 11516.313 | 392960 |  | 11516.313 |  | 0.01 | 0.01108 | 10.81 | bb | 2020-11-23
87 | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9848.446 | 330014 |  | 9848.446 |  | 0.01 | 0.00948 | -5.24 | bb | 2020-11-23
88 | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.48 | 10084.969 | 342025 |  | 10084.969 |  | 0.01 | 0.0097 | -2.97 | bb | 2020-11-23
89 | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.48 | 10773.109 | 363168 |  | 10773.109 |  | 0.01 | 0.01037 | 3.65 | bb | 2020-11-23
90 | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.48 | 11209.78 | 378373 |  | 11209.78 |  | 0.01 | 0.01079 | 7.86 | bb | 2020-11-23
91 | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.48 | 10679.77 | 353296 |  | 10679.77 |  | 0.01 | 0.01028 | 2.76 | bb | 2020-11-23
92 | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.48 | 10166.878 | 344723 |  | 10166.878 |  | 0.01 | 0.00978 | -2.18 | bb | 2020-11-23
93 | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10697.722 | 357232 |  | 10697.722 |  | 0.01 | 0.01029 | 2.93 | bb | 2020-11-24
94 | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.48 | 10158.705 | 334343 |  | 10158.705 |  | 0.01 | 0.00977 | -2.26 | bb | 2020-11-24
95 | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10158.639 | 338366 |  | 10158.639 |  | 0.01 | 0.00977 | -2.26 | bb | 2020-11-24
96 | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.48 | 9865.15 | 326284 |  | 9865.15 |  | 0.01 | 0.00949 | -5.08 | bb | 2020-11-24
97 | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9832.22 | 321241 |  | 9832.22 |  | 0.01 | 0.00946 | -5.4 | bb | 2020-11-24
98 | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
99 | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.48 | 10532.542 | 355947 |  | 10532.542 |  | 0.01 | 0.01013 | 1.34 | bb | 2020-11-24
100 | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.48 | 6178.086 | 209790 |  | 6178.086 |  | 0.01 | 0.00594 | -40.56 | bb | 2020-11-24
101 | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.48 | 10931.824 | 367413 |  | 10931.824 |  | 0.01 | 0.01052 | 5.18 | bb | 2020-11-24
102 | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.48 | 9113.911 | 304569 |  | 9113.911 |  | 0.01 | 0.00877 | -12.31 | bb | 2020-11-24
103 | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
104 | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.48 | 9387.197 | 317571 |  | 9387.197 |  | 0.01 | 0.00903 | -9.68 | bb | 2020-11-24
105 | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9798.542 | 327493 |  | 9798.542 |  | 0.01 | 0.00943 | -5.72 | bb | 2020-11-24
106 | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10690.253 | 359731 |  | 10690.253 |  | 0.01 | 0.01029 | 2.86 | bb | 2020-11-24
107 | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10546.275 | 355639 |  | 10546.275 |  | 0.01 | 0.01015 | 1.47 | bb | 2020-11-24
108 | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10312.553 | 346986 |  | 10312.553 |  | 0.01 | 0.00992 | -0.78 | bb | 2020-11-24
109 | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9162.68 | 310580 |  | 9162.68 |  | 0.01 | 0.00882 | -11.84 | bb | 2020-11-24
110 | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9531.285 | 324174 |  | 9531.285 |  | 0.01 | 0.00917 | -8.29 | bb | 2020-11-24
111 | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.48 | 4318.83 | 148626 |  | 4318.83 |  | 0.01 | 0.00416 | -58.45 | bb | 2020-11-24
112 | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9861.838 | 334756 |  | 9861.838 |  | 0.01 | 0.00949 | -5.11 | bb | 2020-11-24
113 | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10913.814 | 370373 |  | 10913.814 |  | 0.01 | 0.0105 | 5.01 | bb | 2020-11-24
114 | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9692.394 | 329243 |  | 9692.394 |  | 0.01 | 0.00933 | -6.74 | bb | 2020-11-24
115 | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9741.055 | 331371 |  | 9741.055 |  | 0.01 | 0.00937 | -6.28 | bb | 2020-11-24
116 | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8925.449 | 297809 |  | 8925.449 |  | 0.01 | 0.00859 | -14.12 | bb | 2020-11-24
117 | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9877.485 | 335723 |  | 9877.485 |  | 0.01 | 0.0095 | -4.96 | bb | 2020-11-24
118 | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
119 | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.48 | 8900.187 | 298634 |  | 8900.187 |  | 0.01 | 0.00856 | -14.37 | bb | 2020-11-24
120 | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.48 | 9899.175 | 337103 |  | 9899.175 |  | 0.01 | 0.00952 | -4.75 | bb | 2020-11-24
121 | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9065.441 | 305364 |  | 9065.441 |  | 0.01 | 0.00872 | -12.78 | bb | 2020-11-24
122 | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9599.27 | 327251 |  | 9599.27 |  | 0.01 | 0.00924 | -7.64 | bb | 2020-11-24
123 | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8059.229 | 274308 |  | 8059.229 |  | 0.01 | 0.00775 | -22.46 | bb | 2020-11-24
124 | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.48 | 8519.921 | 288423 |  | 8519.921 |  | 0.01 | 0.0082 | -18.02 | bb | 2020-11-24
125 | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  |  | 2020-11-24
126 | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24
127 | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 |  |  | MM- | 2020-11-24

## Analytical Method Details
Analytical Analysis Date: | 2020-11-23
Analytical Technique: | UPLC-MS/MS
Instrument: | Waters Xevo TQ-S micro (QEB0036)
Analyst: | Marci G. Smeltz
Project: | C:\MassLynx\PFAS_PPB_UC.PRO
Sample Worklist: | 20201110_PFAS_PPB_UC_Repeats1_MGS.SPL
MS File: | 20201110_PFAS_PPB_UC_Repeat1_MGS.EXP
TargetLynx File: | 20201130_PFAS_PPB_UC_Repeats1.qld
MS Tune File: | 20201027_Waters_PFC_Tune_MGS.ipr
Inlet File: | 20201027_T3_0pt6mLmin_6pt5min_column2_MGS
Mobile Phase A1: | 2.5 mM ammonium acetate in 95:5 water:acetonitrile
Mobile Phase B1: | 2.5 mM ammonium acetate in 95:5 acetonitrile:water
Analytical Column: | CORTECS T3 Column, 2.7um, 3x100 mm; P/N 186008489 (installed with CORTECS T3 VanGuard Cartridge)
MS File Details |  |  |  |  | *ESI negative for all analytes*
Analyte | DTXSID | Function | Parent (Da) | Daughter (Da) | Cone (V) | Collision Energy (eV)
n-Butylparaben | DTXSID3020209 | 1 | 192.97 | 92.06 | 30 | 24
|  |  | 192.97 | 136.34 | 30 | 16
13C6 n-Butylparaben |  | 2 | 199.15 | 98.03 | 30 | 24
MPFBA |  | 3 | 216.8 | 171.8 | 15 | 10
Dipic Acid | DTXSID4059833 | 4 | 288.95 | 158.85 | 2 | 14
|  |  | 288.95 | 180.83 | 2 | 8
PFBS | DTXSID3037707 | 5 | 298.92 | 79.88 | 2 | 34
|  |  | 298.92 | 98.83 | 2 | 24
M3PFBS |  | 6 | 301.7 | 79.8 | 60 | 30
PFHpA | DTXSID1037303 | 7 | 362.87 | 168.82 | 4 | 14
|  |  | 362.87 | 318.85 | 4 | 6
M4PFHpA |  | 8 | 367 | 322 | 14 | 10
PFPE-6 | DTXSID50375114 | 9 | 466.97 | 235 | 4 | 24
|  |  | 560.92 | 235 | 4 | 28
|  |  | 560.92 | 466.97 | 4 | 16
PFOS | DTXSID80108992 | 10 | 498.97 | 79.87 | 2 | 48
|  |  | 498.97 | 98.82 | 2 | 36
M8PFOS |  | 11 | 507 | 9 | 60 | 48
PFDA | DTXSID3031860 | 12 | 512.86 | 218.87 | 4 | 16
|  |  | 512.86 | 468.84 | 4 | 8
M6PFDA |  | 13 | 519 | 474 | 16 | 12
M7PFUdA |  | 14 | 570 | 525 | 17 | 12
PFTrDA | DTXSID90868151 | 15 | 662.85 | 218.88 | 2 | 22
|  |  | 662.85 | 618.86 | 2 | 12
PFTeDA | DTXSID3059921 | 16 | 712.85 | 668.85 | 2 | 12
M2PFTeDA |  | 17 | 715 | 670 | 19 | 14

## DTXSID4059833
*Analyte in Mix 2 | [6 empty] | Perfluoroadipic Acid
[7 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [6 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 0.68 | 238.644 | 7371 | 49927.879 | 0 | 0.019 | 0.17
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 0.68 | 258.192 | 8212 | 50434.191 | 0 | 0.019 | 0.17
UC_AF_Mix2_A | 151.458 | 4846.6544 | 0.22892 | 0.23539 | 0.0062809 | 2.668% | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 0.68 | 266.135 | 8182 | 51236.457 | 0 | 0.019 | 0.17
UC_AF_Mix2_B | 130.4092 | 4173.0934 | 0.23579 |  |  |  | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 0.68 | 300.587 | 8760 | 53647.801 | 0 | 0.019 | 0.17 | 1.7211 | 912.44
UC_AF_Mix2_C | 122.8008 | 3929.6243 | 0.24147 |  |  |  | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 0.68 | 348.056 | 9903 | 54130.621 | 0 | 0.019 | 0.17 | 53.5733 | 31413.73
UC_T1hr_Mix2_A | 207.8939 | 16631.5128 |  |  |  |  | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 0.68 | 336.485 | 9306 | 55147.813 | 0 | 0.019 | 0.17 | 34.9526 | 20460.34
UC_T1hr_Mix2_B | 220.0781 | 17606.2472 |  |  |  |  | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 0.68 | 346.02 | 9443 | 55044.574 | 0 | 0.019 | 0.17 | 45.667 | 26762.94
UC_T1hr_Mix2_C | 228.8591 | 18308.7288 |  |  |  |  | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 0.68 | 242.337 | 7887 | 49278.141 | 0 | 0.019 | 0.28
UC_T5hr_Mix2_A | 264.6444 | 21171.5496 |  |  |  |  | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 0.68 | 281.091 | 8541 | 51354.762 | 0 | 0.019 | 0.28
UC_T5hr_Mix2_B | 221.2266 | 17698.1256 |  |  |  |  | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 0.68 | 262.542 | 7889 | 51647.742 | 0 | 0.019 | 0.28
UC_T5hr_Mix2_C | 203.4256 | 16274.0488 |  |  |  |  | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 0.68 | 303.198 | 8823 | 55028.125 | 0 | 0.019 | 0.28
[7 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 0.68 | 312.158 | 8578 | 54825.285 | 0 | 0.019 | 0.28 | 8.3666 | 2888.06
MDL Determination | [6 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 0.68 | 353.841 | 9756 | 56832.938 | 0 | 0.019 | 0.28 | 42.2481 | 14988.61
*not determined due to poor sensitivity, lack of data for calibration curve | [6 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 0.68 | 374.919 | 9863 | 55853.539 | 0 | 0.019 | 0.28 | 68.1963 | 24255.81
*recommended to repeat analysis with HILIC method as retention time comes out in column dead volume | [6 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 0.68 | 248.373 | 7951 | 50032.969 | 0 | 0.019 | 0.44
[7 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 0.68 | 268.349 | 8280 | 51327.613 | 0 | 0.019 | 0.44
[7 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 0.68 | 255.774 | 7779 | 51511.32 | 0 | 0.019 | 0.44
[7 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 0.68 | 300.952 | 9208 | 53959.531 | 0 | 0.019 | 0.44
[7 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 0.68 | 324.751 | 8953 | 55636.879 | 0 | 0.019 | 0.44 | 18.2544 | 4048.73
[7 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 0.68 | 341.82 | 9547 | 55401.938 | 0 | 0.019 | 0.44 | 38.9954 | 8762.59
[7 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 0.68 | 337.014 | 9210 | 56901.824 | 0 | 0.019 | 0.44 | 23.8669 | 5324.3
[7 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 0.68 | 230.388 | 7443 | 48869.313 | 0 | 0.019 | 0.71
[7 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 0.68 | 238.684 | 7443 | 49470.234 | 0 | 0.019 | 0.71
[7 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 0.68 | 248.871 | 7757 | 50949.309 | 0 | 0.019 | 0.71
[7 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 0.68 | 318.037 | 9626 | 52453.168 | 0 | 0.019 | 0.71 | 32.6444 | 4497.8
[7 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 0.68 | 311.763 | 9071 | 54984.301 | 0 | 0.019 | 0.71 | 6.6654 | 838.79
[7 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 0.68 | 331.014 | 9525 | 54257.391 | 0 | 0.019 | 0.71 | 34.9105 | 4816.97
[7 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 0.68 | 334.85 | 9637 | 55579.008 | 0 | 0.019 | 0.71 | 30.2877 | 4165.88
[7 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 0.68 | 221.738 | 7076 | 49424.641 | 0 | 0.019 | 1.14
[7 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 0.68 | 246.822 | 7753 | 49787.16 | 0 | 0.019 | 1.14
[7 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 0.68 | 247.588 | 7721 | 50855.781 | 0 | 0.019 | 1.14
[7 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 0.68 | 269.703 | 8030 | 52932.809 | 0 | 0.019 | 1.14
[7 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 0.68 | 320.574 | 9344 | 54067.758 | 0 | 0.019 | 1.14 | 24.2768 | 2029.55
QC FLAGS | [6 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 0.68 | 314.462 | 8687 | 54348.875 | 0 | 0.019 | 1.14 | 14.8139 | 1199.46
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 |  |  |  |  | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 0.68 | 321.177 | 8873 | 54417.223 | 0 | 0.019 | 1.14 | 22.5372 | 1876.94
Curve Precision Checks |  | <75 or >125% of expected |  |  |  |  | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 0.68 | 233.228 | 7486 | 48781.555 | 0 | 0.019 | 1.82
Blanks |  | >1/2 MDL of analytes |  |  |  |  | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 0.68 | 242.172 | 8026 | 50241.621 | 0 | 0.019 | 1.82
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean |  |  |  |  | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 0.68 | 253.325 | 8031 | 50099.988 | 0 | 0.019 | 1.82
[7 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 0.68 | 296.067 | 8883 | 53059.031 | 0 | 0.019 | 1.82
[7 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 0.68 | 302.173 | 8886 | 53643.699 | 0 | 0.019 | 1.82 | 3.955 | 117.31
[7 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 0.68 | 319.549 | 9162 | 54800.516 | 0 | 0.019 | 1.82 | 17.8642 | 881.55
[7 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 0.68 | 353.421 | 10046 | 54960.121 | 0 | 0.019 | 1.82 | 53.6041 | 2845.28
[7 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 0.68 | 294.299 | 8434 | 50327.848 | 0 | 0.019 | 2.91 | 18.9639 | 551.68
[7 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 0.68 | 347.083 | 9945 | 54954.105 | 0 | 0.019 | 2.91 | 47.3291 | 1526.43
[7 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 0.68 | 247.44 | 7224 | 49203.574 | 0 | 0.019 | 4.66
[7 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 0.68 | 302.522 | 9008 | 53011.246 | 0 | 0.019 | 4.66 | 9.2974 | 99.51
[7 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 0.68 | 223.968 | 7078 | 47733.922 | 0 | 0.019 | 7.45
[7 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 0.68 | 281.234 | 8115 | 51497.086 | 0 | 0.019 | 7.45
[7 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 0.68 | 296.6 | 8103 | 50520.066 | 0 | 0.019 | 11.92 | 20.5023 | 72
[7 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 0.68 | 345.506 | 9428 | 54404.949 | 0 | 0.019 | 11.92 | 49.2555 | 313.22
[7 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 0.68 | 280.624 | 7793 | 49274.195 | 0 | 0.019 | 19.07 | 8.4712 | -55.58
[7 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 0.68 | 379.932 | 9692 | 54264.176 | 0 | 0.019 | 19.07 | 82.1056 | 330.55
[7 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 0.68 | 300.632 | 8312 | 49934.551 | 0 | 0.019 | 30.52 | 30.0263 | -1.62
[7 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 0.68 | 355.709 | 8975 | 53552.957 | 0 | 0.019 | 30.52 | 64.6595 | 111.86
[7 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 0.68 | 355.055 | 7962 | 50913.094 | 0 | 0.019 | 48.83 | 80.8086 | 65.49
[7 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 0.68 | 329.829 | 8414 | 53773.059 | 0 | 0.019 | 48.83 | 36.8712 | -24.49
[7 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 0.68 | 342.088 | 7422 | 48568.758 | 0 | 0.019 | 78.13 | 84.0448 | 7.57
[7 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 0.68 | 406.014 | 8995 | 52683.031 | 0 | 0.019 | 78.13 | 112.7151 | 44.27
[7 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 0.68 | 452.098 | 8259 | 51406.383 | 0 | 0.019 | 125 | 153.5087 | 22.81
[7 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 0.68 | 507.918 | 9452 | 55300.879 | 0 | 0.019 | 125 | 166.7689 | 33.42
[7 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 0.65 | 452.462 | 8218 | 51234.66 | 0 | 0.019 | 156.25 | 154.7787 | -0.94
[7 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 0.68 | 511.885 | 9733 | 53606.293 | 0 | 0.019 | 156.25 | 178.6402 | 14.33
[7 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 0.65 | 563.38 | 11063 | 51287.395 | 0 | 0.019 | 250 | 221.5325 | -11.39
[7 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 0.65 | 616.254 | 11422 | 54873.789 | 0 | 0.019 | 250 | 228.3553 | -8.66
[7 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 0.66 | 1393.331 | 20801 |  |  | 0.019
[7 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 0.66 | 1502.766 | 25643 |  |  | 0.019
[7 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 0.65 | 1220.616 | 23424 |  |  | 0.019
[7 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 0.65 | 1005.532 | 19754 |  |  | 0.019
[7 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 0.68 | 250.861 | 7254 | 50185.297 | 0 | 0.019
[7 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 0.68 | 296.555 | 8557 | 50805.66 | 0 | 0.019 |  | 18.2593
[7 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 0.68 | 313.731 | 8913 | 53139.266 | 0 | 0.019 |  | 22.655
[7 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 0.68 | 318.624 | 8913 | 53001.285 | 0 | 0.019 |  | 29.4761
[7 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 0.68 | 335.093 | 9093 | 53940.129 | 0 | 0.019 |  | 41.4621
[7 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 0.68 | 362.04 | 9832 | 54180.07 | 0 | 0.019 |  | 66.6769
[7 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 0.68 | 359.962 | 9499 | 55176.629 | 0 | 0.019 |  | 58.5575
[7 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 0.68 | 388.666 | 9676 | 57833.516 | 0 | 0.019 |  | 68.589
[7 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 0.66 | 12.7 | 536 | 19403.883 | 0 | 0.019
[7 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 0.68 | 226.647 | 7164 | 49652.594 | 0 | 0.019 | 0.63
[7 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 0.68 | 317.799 | 9637 | 54029.367 | 0 | 0.019 | 0.63 | 21.2253 | 3269.1
[7 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 0.68 | 295.704 | 8762 | 52099.977 | 0 | 0.019 | 2.5 | 7.0749 | 183
[7 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 0.68 | 405.041 | 9916 | 61291.387 | 0 | 0.019 | 2.5 | 62.9399 | 2417.6
[7 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 0.68 | 262.242 | 7624 | 51356.805 | 0 | 0.019 | 6.25
[7 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 0.68 | 324.476 | 9044 | 53206.5 | 0 | 0.019 | 6.25 | 34.7677 | 456.28
[7 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 0.68 | 309.147 | 8904 | 51164.73 | 0 | 0.019 | 25 | 31.3591 | 25.44
[7 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 0.68 | 338.175 | 9132 | 54984.762 | 0 | 0.019 | 25 | 37.8525 | 51.41
[7 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 0.68 | 0.156 | 7 |  |  | 0.019
[7 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 0.68 | 0.243 | 12 |  |  | 0.019
[7 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 0.68 | 0.069 | 3 |  |  | 0.019
[7 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent |  |  |  | 4.41 |  | 0.019
[7 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 0.66 | 0.04 | 2 |  |  | 0.019
[7 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 0.68 | 0.487 | 12 |  |  | 0.019
[7 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 0.67 | 0.042 | 2 |  |  | 0.019
[7 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.019
[7 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 0.68 | 399.129 | 9404 | 55329.66 | 0 | 0.019 |  | 91.7593
[7 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 0.68 | 398.924 | 9809 | 55421.43 | 0 | 0.019 |  | 91.0618
[7 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 0.68 | 369.02 | 9798 | 52817.883 | 0 | 0.019 |  | 81.412
[7 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 0.68 | 464.598 | 9763 | 53182.844 | 0 | 0.019 |  | 151.458
[7 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 0.68 | 435.556 | 9527 | 53391.262 | 0 | 0.019 |  | 130.4092
[7 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 0.68 | 425.336 | 9729 | 53434.508 | 0 | 0.019 |  | 122.8008
[7 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 0.68 | 405.582 | 9958 | 54932.02 | 0 | 0.019 |  | 99.1882
[7 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 0.68 | 381.698 | 9809 | 54176.371 | 0 | 0.019 |  | 84.1414
[7 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 0.68 | 336.534 | 9306 | 52035.281 | 0 | 0.019 |  | 55.5798
[7 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 0.68 | 327.478 | 8181 | 54616.355 | 0 | 0.019 |  | 28.5026
[7 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 0.68 | 335.911 | 9067 | 53656.375 | 0 | 0.019 |  | 44.2118
[7 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 0.68 | 308.265 | 8142 | 51797.668 | 0 | 0.019 |  | 25.6937
[7 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 0.65 | 549.797 | 10329 | 52322.02 | 0 | 0.019 |  | 207.8939
[7 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 0.65 | 589.284 | 11168 | 53899.488 | 0 | 0.019 |  | 220.0781
[7 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 0.65 | 605.892 | 11355 | 53863.215 | 0 | 0.019 |  | 228.8591
[7 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 0.68 | 327.587 | 9297 | 52084.168 | 0 | 0.019 |  | 45.8576
[7 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 0.68 | 318.111 | 8677 | 52470.254 | 0 | 0.019 |  | 32.6101
[7 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 0.68 | 322.782 | 8480 | 53591.719 | 0 | 0.019 |  | 30.1783
[7 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 0.68 | 384.081 | 9198 | 59098.656 | 0 | 0.019 |  | 57.2521
[7 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 0.68 | 336.718 | 8595 | 54825.996 | 0 | 0.019 |  | 37.3357
[7 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 0.68 | 352.189 | 8996 | 53523.332 | 0 | 0.019 |  | 61.4833
[7 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 0.65 | 685.725 | 14944 | 54352.184 | 0 | 0.019 |  | 264.6444
[7 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 0.65 | 603.88 | 10458 | 55028.941 | 0 | 0.019 |  | 221.2266
[7 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 0.65 | 559.077 | 9523 | 53986.629 | 0 | 0.019 |  | 203.4256
[7 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 0.68 | 365.847 | 9218 | 54814.41 | 0 | 0.019 |  | 66.281
[7 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 0.68 | 336.965 | 8706 | 54693.875 | 0 | 0.019 |  | 38.4746
[7 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 0.68 | 359.952 | 9353 | 55158.957 | 0 | 0.019 |  | 58.6574
MPFBA
[7 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[7 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.06 | 49927.879 | 1093793 | 49927.879 |  |  | 0.01 | 0.00951 | -4.88
[7 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.06 | 50434.191 | 1144254 | 50434.191 |  |  | 0.01 | 0.00961 | -3.92
[7 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.06 | 51236.457 | 1180196 | 51236.457 |  |  | 0.01 | 0.00976 | -2.39
[7 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.06 | 53647.801 | 1306898 | 53647.801 |  |  | 0.01 | 0.01022 | 2.2
[7 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.06 | 54130.621 | 1397893 | 54130.621 |  |  | 0.01 | 0.01031 | 3.12
[7 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.06 | 55147.813 | 1457940 | 55147.813 |  |  | 0.01 | 0.01051 | 5.06
[7 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.06 | 55044.574 | 1487081 | 55044.574 |  |  | 0.01 | 0.01049 | 4.86
[7 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.06 | 49278.141 | 1055807 | 49278.141 |  |  | 0.01 | 0.00939 | -6.12
[7 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.06 | 51354.762 | 1178770 | 51354.762 |  |  | 0.01 | 0.00978 | -2.17
[7 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.06 | 51647.742 | 1197739 | 51647.742 |  |  | 0.01 | 0.00984 | -1.61
[7 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.06 | 55028.125 | 1360267 | 55028.125 |  |  | 0.01 | 0.01048 | 4.83
[7 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.06 | 54825.285 | 1426003 | 54825.285 |  |  | 0.01 | 0.01044 | 4.45
[7 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.06 | 56832.938 | 1497038 | 56832.938 |  |  | 0.01 | 0.01083 | 8.27
[7 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.06 | 55853.539 | 1506542 | 55853.539 |  |  | 0.01 | 0.01064 | 6.4
[7 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.06 | 50032.969 | 1093526 | 50032.969 |  |  | 0.01 | 0.00953 | -4.68
[7 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.06 | 51327.613 | 1155420 | 51327.613 |  |  | 0.01 | 0.00978 | -2.22
[7 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.06 | 51511.32 | 1185758 | 51511.32 |  |  | 0.01 | 0.00981 | -1.87
[7 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.06 | 53959.531 | 1333779 | 53959.531 |  |  | 0.01 | 0.01028 | 2.8
[7 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.06 | 55636.879 | 1448618 | 55636.879 |  |  | 0.01 | 0.0106 | 5.99
[7 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.06 | 55401.938 | 1475862 | 55401.938 |  |  | 0.01 | 0.01055 | 5.54
[7 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.06 | 56901.824 | 1536143 | 56901.824 |  |  | 0.01 | 0.01084 | 8.4
[7 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.06 | 48869.313 | 1083320 | 48869.313 |  |  | 0.01 | 0.00931 | -6.9
[7 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.06 | 49470.234 | 1141155 | 49470.234 |  |  | 0.01 | 0.00942 | -5.76
[7 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.06 | 50949.309 | 1197432 | 50949.309 |  |  | 0.01 | 0.00971 | -2.94
[7 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.06 | 52453.168 | 1314141 | 52453.168 |  |  | 0.01 | 0.00999 | -0.07
[7 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.06 | 54984.301 | 1456655 | 54984.301 |  |  | 0.01 | 0.01047 | 4.75
[7 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.06 | 54257.391 | 1447254 | 54257.391 |  |  | 0.01 | 0.01034 | 3.36
[7 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.06 | 55579.008 | 1518421 | 55579.008 |  |  | 0.01 | 0.01059 | 5.88
[7 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.06 | 49424.641 | 1122354 | 49424.641 |  |  | 0.01 | 0.00942 | -5.84
[7 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.06 | 49787.16 | 1161960 | 49787.16 |  |  | 0.01 | 0.00948 | -5.15
[7 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.06 | 50855.781 | 1213615 | 50855.781 |  |  | 0.01 | 0.00969 | -3.12
[7 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.06 | 52932.809 | 1342364 | 52932.809 |  |  | 0.01 | 0.01008 | 0.84
[7 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54067.758 | 1437007 | 54067.758 |  |  | 0.01 | 0.0103 | 3
[7 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54348.875 | 1468223 | 54348.875 |  |  | 0.01 | 0.01035 | 3.54
[7 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.06 | 54417.223 | 1506331 | 54417.223 |  |  | 0.01 | 0.01037 | 3.67
[7 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.06 | 48781.555 | 1097833 | 48781.555 |  |  | 0.01 | 0.00929 | -7.07
[7 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.06 | 50241.621 | 1172509 | 50241.621 |  |  | 0.01 | 0.00957 | -4.29
[7 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.06 | 50099.988 | 1170203 | 50099.988 |  |  | 0.01 | 0.00954 | -4.56
[7 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.06 | 53059.031 | 1324542 | 53059.031 |  |  | 0.01 | 0.01011 | 1.08
[7 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.06 | 53643.699 | 1379321 | 53643.699 |  |  | 0.01 | 0.01022 | 2.19
[7 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.06 | 54800.516 | 1443545 | 54800.516 |  |  | 0.01 | 0.01044 | 4.4
[7 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.06 | 54960.121 | 1474615 | 54960.121 |  |  | 0.01 | 0.01047 | 4.7
[7 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.06 | 50327.848 | 1119790 | 50327.848 |  |  | 0.01 | 0.00959 | -4.12
[7 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.06 | 54954.105 | 1332580 | 54954.105 |  |  | 0.01 | 0.01047 | 4.69
[7 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.06 | 49203.574 | 1136669 | 49203.574 |  |  | 0.01 | 0.00937 | -6.26
[7 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.06 | 53011.246 | 1295984 | 53011.246 |  |  | 0.01 | 0.0101 | 0.99
[7 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.06 | 47733.922 | 1082948 | 47733.922 |  |  | 0.01 | 0.00909 | -9.06
[7 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.06 | 51497.086 | 1264257 | 51497.086 |  |  | 0.01 | 0.00981 | -1.89
[7 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.06 | 50520.066 | 1152278 | 50520.066 |  |  | 0.01 | 0.00962 | -3.76
[7 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.06 | 54404.949 | 1359505 | 54404.949 |  |  | 0.01 | 0.01036 | 3.64
[7 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.06 | 49274.195 | 1120503 | 49274.195 |  |  | 0.01 | 0.00939 | -6.13
[7 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.06 | 54264.176 | 1381855 | 54264.176 |  |  | 0.01 | 0.01034 | 3.38
[7 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.06 | 49934.551 | 1152862 | 49934.551 |  |  | 0.01 | 0.00951 | -4.87
[7 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.06 | 53552.957 | 1318423 | 53552.957 |  |  | 0.01 | 0.0102 | 2.02
[7 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.06 | 50913.094 | 1176729 | 50913.094 |  |  | 0.01 | 0.0097 | -3.01
[7 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.06 | 53773.059 | 1374467 | 53773.059 |  |  | 0.01 | 0.01024 | 2.44
[7 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.06 | 48568.758 | 1127876 | 48568.758 |  |  | 0.01 | 0.00925 | -7.47
[7 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.06 | 52683.031 | 1324650 | 52683.031 |  |  | 0.01 | 0.01004 | 0.36
[7 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.06 | 51406.383 | 1220473 | 51406.383 |  |  | 0.01 | 0.00979 | -2.07
[7 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.06 | 55300.879 | 1439829 | 55300.879 |  |  | 0.01 | 0.01054 | 5.35
[7 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.06 | 51234.66 | 1192786 | 51234.66 |  |  | 0.01 | 0.00976 | -2.39
[7 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.06 | 53606.293 | 1358384 | 53606.293 |  |  | 0.01 | 0.01021 | 2.12
[7 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.06 | 51287.395 | 1203006 | 51287.395 |  |  | 0.01 | 0.00977 | -2.29
[7 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.06 | 54873.789 | 1419838 | 54873.789 |  |  | 0.01 | 0.01045 | 4.54
[7 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[7 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[7 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[7 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[7 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.06 | 50185.297 | 1079128 | 50185.297 |  |  | 0.01 | 0.00956 | -4.39
[7 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.06 | 50805.66 | 1121952 | 50805.66 |  |  | 0.01 | 0.00968 | -3.21
[7 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.06 | 53139.266 | 1208079 | 53139.266 |  |  | 0.01 | 0.01012 | 1.23
[7 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.06 | 53001.285 | 1240747 | 53001.285 |  |  | 0.01 | 0.0101 | 0.97
[7 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.06 | 53940.129 | 1314646 | 53940.129 |  |  | 0.01 | 0.01028 | 2.76
[7 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.06 | 54180.07 | 1338242 | 54180.07 |  |  | 0.01 | 0.01032 | 3.22
[7 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.06 | 55176.629 | 1426925 | 55176.629 |  |  | 0.01 | 0.01051 | 5.11
[7 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.06 | 57833.516 | 1565582 | 57833.516 |  |  | 0.01 | 0.01102 | 10.18
[7 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.06 | 19403.883 | 608476 | 19403.883 |  |  | 0.01 | 0.0037 | -63.03
[7 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.06 | 49652.594 | 1169542 | 49652.594 |  |  | 0.01 | 0.00946 | -5.41
[7 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.06 | 54029.367 | 1383832 | 54029.367 |  |  | 0.01 | 0.01029 | 2.93
[7 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.06 | 52099.977 | 1294191 | 52099.977 |  |  | 0.01 | 0.00993 | -0.75
[7 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.07 | 61291.387 | 1905970 | 61291.387 |  |  | 0.01 | 0.01168 | 16.76
[7 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.06 | 51356.805 | 1175106 | 51356.805 |  |  | 0.01 | 0.00978 | -2.16
[7 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.06 | 53206.5 | 1365779 | 53206.5 |  |  | 0.01 | 0.01014 | 1.36
[7 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.06 | 51164.73 | 1208782 | 51164.73 |  |  | 0.01 | 0.00975 | -2.53
[7 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.06 | 54984.762 | 1446908 | 54984.762 |  |  | 0.01 | 0.01047 | 4.75
[7 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 1.08 | 4.41 | 155 | 4.41 |  |  | 0.01 | 0 | -99.99
[7 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[7 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.06 | 55329.66 | 1289571 | 55329.66 |  |  | 0.01 | 0.01054 | 5.41
[7 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.06 | 55421.43 | 1277857 | 55421.43 |  |  | 0.01 | 0.01056 | 5.58
[7 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.06 | 52817.883 | 1207031 | 52817.883 |  |  | 0.01 | 0.01006 | 0.62
[7 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.06 | 53182.844 | 1285004 | 53182.844 |  |  | 0.01 | 0.01013 | 1.32
[7 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.06 | 53391.262 | 1296578 | 53391.262 |  |  | 0.01 | 0.01017 | 1.71
[7 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.06 | 53434.508 | 1288257 | 53434.508 |  |  | 0.01 | 0.01018 | 1.8
[7 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.06 | 54932.02 | 1311721 | 54932.02 |  |  | 0.01 | 0.01046 | 4.65
[7 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.06 | 54176.371 | 1285553 | 54176.371 |  |  | 0.01 | 0.01032 | 3.21
[7 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.06 | 52035.281 | 1237776 | 52035.281 |  |  | 0.01 | 0.00991 | -0.87
[7 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.06 | 54616.355 | 1345293 | 54616.355 |  |  | 0.01 | 0.0104 | 4.05
[7 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.06 | 53656.375 | 1364620 | 53656.375 |  |  | 0.01 | 0.01022 | 2.22
[7 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.06 | 51797.668 | 1288013 | 51797.668 |  |  | 0.01 | 0.00987 | -1.32
[7 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.06 | 52322.02 | 1310806 | 52322.02 |  |  | 0.01 | 0.00997 | -0.32
[7 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.06 | 53899.488 | 1352836 | 53899.488 |  |  | 0.01 | 0.01027 | 2.68
[7 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.06 | 53863.215 | 1355801 | 53863.215 |  |  | 0.01 | 0.01026 | 2.61
[7 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.06 | 52084.168 | 1324937 | 52084.168 |  |  | 0.01 | 0.00992 | -0.78
[7 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.06 | 52470.254 | 1353608 | 52470.254 |  |  | 0.01 | 0.01 | -0.04
[7 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.06 | 53591.719 | 1369970 | 53591.719 |  |  | 0.01 | 0.01021 | 2.1
[7 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.06 | 59098.656 | 1403313 | 59098.656 |  |  | 0.01 | 0.01126 | 12.59
[7 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.06 | 54825.996 | 1429172 | 54825.996 |  |  | 0.01 | 0.01044 | 4.45
[7 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.06 | 53523.332 | 1369929 | 53523.332 |  |  | 0.01 | 0.0102 | 1.97
[7 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.06 | 54352.184 | 1413246 | 54352.184 |  |  | 0.01 | 0.01035 | 3.54
[7 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.06 | 55028.941 | 1441641 | 55028.941 |  |  | 0.01 | 0.01048 | 4.83
[7 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.06 | 53986.629 | 1407332 | 53986.629 |  |  | 0.01 | 0.01028 | 2.85
[7 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.06 | 54814.41 | 1382624 | 54814.41 |  |  | 0.01 | 0.01044 | 4.42
[7 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.06 | 54693.875 | 1429813 | 54693.875 |  |  | 0.01 | 0.01042 | 4.2
[7 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.06 | 55158.957 | 1416562 | 55158.957 |  |  | 0.01 | 0.01051 | 5.08

## DTXSID3037707
*Analyte in Mix 3 | [8 empty] | PFBS
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.69 | 270.846 | 13568 | 21861.801 | 0 | 0.997 | 0.17 | 0.04426 | -73.96
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.69 | 305.436 | 14419 | 23188.197 | 0 | 0.997 | 0.17 | 0.05627 | -66.9
UC_AF_Mix3_A | 1.4851 | 47.5235 | 0.004276 | 0.0046354 | 0.00042468 | 9.162% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.69 | 279.533 | 14261 | 22730.893 | 0 | 0.997 | 0.17 | 0.04286 | -74.79
UC_AF_Mix3_B | 1.6338 | 52.2803 | 0.0045262 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.69 | 284.038 | 12562 | 24736.178 | 0 | 0.997 | 0.17 | 0.03036 | -82.14
UC_AF_Mix3_C | 1.8296 | 58.5456 | 0.005104 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.69 | 310.445 | 16342 | 24669.846 | 0 | 0.997 | 0.17 | 0.04725 | -72.21
UC_T1hr_Mix3_A | 140.8451 | 11267.6064 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.69 | 332.996 | 16342 | 25334.619 | 0 | 0.997 | 0.17 | 0.05584 | -67.15
UC_T1hr_Mix3_B | 141.7737 | 11341.8976 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.69 | 298.9 | 15508 | 25002.83 | 0 | 0.997 | 0.17 | 0.0376 | -77.88
UC_T1hr_Mix3_C | 143.6215 | 11489.72 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.69 | 765.518 | 39083 | 22905.938 | 0 | 0.997 | 0.28 | 0.36684 | 31.01
UC_T5hr_Mix3_A | 138.9257 | 11114.0552 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.69 | 794.594 | 40541 | 23955.818 | 0 | 0.997 | 0.28 | 0.36299 | 29.64
UC_T5hr_Mix3_B | 144.3828 | 11550.624 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.69 | 779.21 | 39072 | 24510.729 | 0 | 0.997 | 0.28 | 0.34184 | 22.09
UC_T5hr_Mix3_C | 143.381 | 11470.4824 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.69 | 849.65 | 42095 | 25389.068 | 0 | 0.997 | 0.28 | 0.36753 | 31.26
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.69 | 856.988 | 42860 | 26169.432 | 0 | 0.997 | 0.28 | 0.35652 | 27.33
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.69 | 860.977 | 41948 | 26174.143 | 0 | 0.997 | 0.28 | 0.35877 | 28.13
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.69 | 847.982 | 42497 | 26803.898 | 0 | 0.997 | 0.28 | 0.33948 | 21.24
20201123_PFAS_PPB_UC_Sample010 | CC4 | 0.71 | 0.91581 | 0.89463 | 0.01875 | 2.096% | 0.058931 | 26% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1098.736 | 54638 | 22565.848 | 0 | 0.997 | 0.44 | 0.60105 | 36.6
20201123_PFAS_PPB_UC_Sample035 | CC4 | 0.71 | 0.91994 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1132.518 | 56657 | 23133.342 | 0 | 0.997 | 0.44 | 0.60512 | 37.53
20201123_PFAS_PPB_UC_Sample042 | CC4 | 0.71 | 0.87747 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1118.739 | 56886 | 24355.15 | 0 | 0.997 | 0.44 | 0.55878 | 27
20201123_PFAS_PPB_UC_Sample082 | CC4 | 0.71 | 0.86858 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1192.203 | 59300 | 24953.592 | 0 | 0.997 | 0.44 | 0.58704 | 33.42
20201123_PFAS_PPB_UC_Sample108 | CC4 | 0.71 | 0.89318 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1220.072 | 61537 | 26166.074 | 0 | 0.997 | 0.44 | 0.56942 | 29.41
20201123_PFAS_PPB_UC_Sample115 | CC4 | 0.71 | 0.89837 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1230.669 | 61669 | 26063.229 | 0 | 0.997 | 0.44 | 0.57848 | 31.47
20201123_PFAS_PPB_UC_Sample122 | CC4 | 0.71 | 0.88906 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.69 | 1225.297 | 61938 | 26872.852 | 0 | 0.997 | 0.44 | 0.55359 | 25.82
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1545.47 | 79762 | 22329.166 | 0.001 | 0.997 | 0.71 | 0.91581 | 28.99
Blanks should be less than… | 0.029466 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1597.693 | 81589 | 22994.361 | 0.001 | 0.997 | 0.71 | 0.91994 | 29.57
eLOQ | 0.71 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1584.861 | 79827 | 23756.479 | 0.001 | 0.997 | 0.71 | 0.87747 | 23.59
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1631.263 | 82380 | 24666.17 | 0.001 | 0.997 | 0.71 | 0.86858 | 22.34
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1693.45 | 87004 | 25000.137 | 0.001 | 0.997 | 0.71 | 0.89318 | 25.8
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1718.22 | 87933 | 25239.68 | 0.001 | 0.997 | 0.71 | 0.89837 | 26.53
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.69 | 1736.466 | 86665 | 25737.27 | 0.001 | 0.997 | 0.71 | 0.88906 | 25.22
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1717.796 | 85594 | 22016.998 | 0.001 | 0.997 | 1.14 | 1.0509 | -7.82
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1727.973 | 85843 | 22848.482 | 0.001 | 0.997 | 1.14 | 1.0142 | -11.04
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1822.375 | 93105 | 23377.701 | 0.001 | 0.997 | 1.14 | 1.0499 | -7.91
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1962.02 | 99118 | 24743.717 | 0.001 | 0.997 | 1.14 | 1.0704 | -6.1
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1954.207 | 99743 | 25457.678 | 0.001 | 0.997 | 1.14 | 1.0316 | -9.51
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.69 | 2008.324 | 103481 | 25575.846 | 0.001 | 0.997 | 1.14 | 1.0586 | -7.14
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.69 | 1967.181 | 101259 | 25205.348 | 0.001 | 0.997 | 1.14 | 1.0513 | -7.78
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2849.414 | 143965 | 22032.42 | 0.001 | 0.997 | 1.82 | 1.8378 | 0.98
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2846.754 | 143471 | 22995.889 | 0.001 | 0.997 | 1.82 | 1.7529 | -3.69
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2931.468 | 148720 | 23030.449 | 0.001 | 0.997 | 1.82 | 1.8064 | -0.74
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3003.684 | 151614 | 24207.008 | 0.001 | 0.997 | 1.82 | 1.7573 | -3.44
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.69 | 2986.899 | 153470 | 24884.777 | 0.001 | 0.997 | 1.82 | 1.6951 | -6.86
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3014.977 | 149427 | 25309.125 | 0.001 | 0.997 | 1.82 | 1.6813 | -7.62
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.69 | 3135.531 | 160453 | 25344.141 | 0.001 | 0.997 | 1.82 | 1.7517 | -3.75
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.69 | 4048.259 | 203877 | 22609.912 | 0.002 | 0.997 | 2.91 | 2.6002 | -10.65
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.69 | 4509.937 | 226581 | 24554.58 | 0.002 | 0.997 | 2.91 | 2.6711 | -8.21
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.69 | 6474.464 | 332965 | 21791.783 | 0.003 | 0.997 | 4.66 | 4.4104 | -5.36
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.69 | 7021.911 | 350201 | 23955.301 | 0.003 | 0.997 | 4.66 | 4.3494 | -6.67
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.69 | 10672.343 | 534297 | 21652.168 | 0.005 | 0.997 | 7.45 | 7.4119 | -0.51
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.69 | 11412.911 | 576123 | 23626.941 | 0.005 | 0.997 | 7.45 | 7.2608 | -2.54
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.69 | 18260.215 | 923656 | 21857.189 | 0.008 | 0.997 | 11.92 | 12.6607 | 6.21
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.69 | 19943.035 | 1004406 | 24218.33 | 0.008 | 0.997 | 11.92 | 12.4774 | 4.68
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.69 | 31758.162 | 1595521 | 22171.375 | 0.014 | 0.997 | 19.07 | 21.8017 | 14.32
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.69 | 34046.984 | 1690798 | 24038.738 | 0.014 | 0.997 | 19.07 | 21.556 | 13.04
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.69 | 44598.066 | 2214801 | 21427.695 | 0.021 | 0.997 | 30.52 | 31.7294 | 3.96
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.69 | 47851.457 | 2390021 | 23164.816 | 0.021 | 0.997 | 30.52 | 31.4904 | 3.18
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.69 | 66633.859 | 3282871 | 20779.268 | 0.032 | 0.997 | 48.83 | 48.9239 | 0.19
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.69 | 72477.875 | 3580599 | 22853.115 | 0.032 | 0.997 | 48.83 | 48.3853 | -0.91
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.69 | 101232.203 | 4914728 | 19147.74 | 0.053 | 0.997 | 78.13 | 80.6296 | 3.2
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.69 | 106142.273 | 5134447 | 20442.584 | 0.052 | 0.997 | 78.13 | 79.1885 | 1.35
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.69 | 140618.609 | 6703322 | 18008.693 | 0.078 | 0.997 | 125 | 118.9318 | -4.85
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.69 | 148861.125 | 7077424 | 19474.186 | 0.076 | 0.997 | 125 | 116.4396 | -6.85
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.69 | 174261.922 | 8197679 | 17279.416 | 0.101 | 0.997 | 156.25 | 153.3926 | -1.83
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.69 | 187885.969 | 8887854 | 18660.311 | 0.101 | 0.997 | 156.25 | 153.1479 | -1.99
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.69 | 263390.281 | 12053817 | 15521.709 | 0.17 | 0.997 | 250 | 256.9115 | 2.76
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.69 | 281528.156 | 12888545 | 16898.207 | 0.167 | 0.997 | 250 | 252.2882 | 0.92
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [5 empty] | 0.997
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 1.7 | 1.247 | 119 |  |  | 0.997
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 1.7 | 12.433 | 347 |  |  | 0.997
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 1.7 | 2.178 | 166 |  |  | 0.997
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.69 | 32.985 | 1541 | 21628.184 | 0 | 0.997
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.69 | 44.182 | 2081 | 22818.307 | 0 | 0.997
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.69 | 43.611 | 2163 | 22658.945 | 0 | 0.997
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.7 | 42.821 | 2080 | 23242.805 | 0 | 0.997
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.69 | 46.517 | 2451 | 24092.141 | 0 | 0.997
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.7 | 84.043 | 2658 | 24051.273 | 0 | 0.997
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.69 | 56.11 | 2585 | 24598.893 | 0 | 0.997
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.69 | 42.679 | 1983 | 25048.125 | 0 | 0.997
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.69 | 17.689 | 933 | 8914.928 | 0 | 0.997
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.69 | 960.737 | 45972 | 22865.168 | 0 | 0.997 | 0.63 | 0.4987 | -20.84
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.69 | 1062.019 | 52544 | 24217.922 | 0 | 0.997 | 0.63 | 0.52685 | -16.37
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.69 | 4319.657 | 220943 | 24861.91 | 0.002 | 0.997 | 2.5 | 2.5189 | 0.76
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.69 | 5109.564 | 255476 | 29108.41 | 0.002 | 0.997 | 2.5 | 2.5464 | 1.85
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.69 | 10008.617 | 505952 | 23299.66 | 0.004 | 0.997 | 6.25 | 6.441 | 3.06
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.69 | 10634.096 | 531629 | 24923.5 | 0.004 | 0.997 | 6.25 | 6.3967 | 2.35
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.69 | 37407.426 | 1871102 | 21679.879 | 0.017 | 0.997 | 25 | 26.2861 | 5.14
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.69 | 40760.316 | 2015584 | 24550.033 | 0.017 | 0.997 | 25 | 25.2894 | 1.16
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 1.67 | 0.492 | 47 |  |  | 0.997
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 1.68 | 0.123 | 12 |  |  | 0.997
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 1.73 | 3.023 | 134 |  |  | 0.997
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 1.71 | 0.381 | 37 |  |  | 0.997
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [5 empty] | 0.997
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 1.67 | 0.584 | 57 |  |  | 0.997
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 1.74 | 1.503 | 80 |  |  | 0.997
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [5 empty] | 0.997
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [5 empty] | 0.997
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 1.66 | 27.867 | 667 |  |  | 0.997
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 1.72 | 4.584 | 217 |  |  | 0.997
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 1.73 | 5.926 | 246 |  |  | 0.997
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 1.69 | 2.708 | 188 |  |  | 0.997
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 1.76 | 2.691 | 182 |  |  | 0.997
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.997
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.69 | 86.739 | 4274 | 25180.121 | 0 | 0.997
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.69 | 101.228 | 4709 | 25625.738 | 0 | 0.997
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.69 | 78.525 | 3265 | 24362.371 | 0 | 0.997
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.69 | 79.826 | 4069 | 24305.764 | 0 | 0.997
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.7 | 83.702 | 3666 | 24109.816 | 0 | 0.997
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.69 | 102.883 | 5027 | 24456.668 | 0 | 0.997
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.69 | 2653.8 | 134402 | 24957.41 | 0.001 | 0.997 |  | 1.4851
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.69 | 2918.392 | 146406 | 25152.914 | 0.001 | 0.997 |  | 1.6338
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.69 | 3116.855 | 156086 | 24200.555 | 0.001 | 0.997 |  | 1.8296
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.69 | 38.017 | 1873 | 22887.045 | 0 | 0.997
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.69 | 100.904 | 5253 | 22461.596 | 0 | 0.997
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.69 | 239.046 | 12328 | 23178.355 | 0 | 0.997 |  | 0.01242
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.7 | 104.639 | 5040 | 22570.561 | 0 | 0.997
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.69 | 67.257 | 3414 | 23917.525 | 0 | 0.997
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.69 | 54.455 | 2625 | 23604.629 | 0 | 0.997
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.69 | 163018.234 | 7720060 | 17613.797 | 0.093 | 0.997 |  | 140.8451
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.69 | 166278.828 | 7892064 | 17847.734 | 0.093 | 0.997 |  | 141.7737
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.69 | 175612.5 | 8300231 | 18605.641 | 0.094 | 0.997 |  | 143.6215
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.69 | 41.388 | 1913 | 26959.02 | 0 | 0.997
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.69 | 63.592 | 2794 | 23924.834 | 0 | 0.997
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.69 | 35.389 | 1826 | 24150.537 | 0 | 0.997
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.69 | 49.728 | 2225 | 23967.572 | 0 | 0.997
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.69 | 103.038 | 5567 | 23285.992 | 0 | 0.997
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.69 | 248.04 | 12986 | 24485.539 | 0 | 0.997 |  | 0.00961
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.69 | 176105.891 | 8342776 | 19292.301 | 0.091 | 0.997 |  | 138.9257
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.69 | 172904.688 | 8105040 | 18221.59 | 0.095 | 0.997 |  | 144.3828
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.69 | 173770.672 | 8266025 | 18441.566 | 0.094 | 0.997 |  | 143.381
M3PFBS
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.69 | 21861.801 | 1121154 | 21861.801 |  |  | 0.01 | 0.0094 | -6
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.69 | 23188.197 | 1175156 | 23188.197 |  |  | 0.01 | 0.00997 | -0.29
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.69 | 22730.893 | 1171782 | 22730.893 |  |  | 0.01 | 0.00977 | -2.26
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.69 | 24736.178 | 1278515 | 24736.178 |  |  | 0.01 | 0.01064 | 6.36
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.69 | 24669.846 | 1260177 | 24669.846 |  |  | 0.01 | 0.01061 | 6.08
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.69 | 25334.619 | 1298374 | 25334.619 |  |  | 0.01 | 0.01089 | 8.93
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.69 | 25002.83 | 1282939 | 25002.83 |  |  | 0.01 | 0.01075 | 7.51
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.69 | 22905.938 | 1157712 | 22905.938 |  |  | 0.01 | 0.00985 | -1.51
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.69 | 23955.818 | 1215689 | 23955.818 |  |  | 0.01 | 0.0103 | 3.01
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.69 | 24510.729 | 1260704 | 24510.729 |  |  | 0.01 | 0.01054 | 5.39
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.69 | 25389.068 | 1294965 | 25389.068 |  |  | 0.01 | 0.01092 | 9.17
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26169.432 | 1345344 | 26169.432 |  |  | 0.01 | 0.01125 | 12.52
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26174.143 | 1354444 | 26174.143 |  |  | 0.01 | 0.01125 | 12.54
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.69 | 26803.898 | 1364492 | 26803.898 |  |  | 0.01 | 0.01153 | 15.25
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.69 | 22565.848 | 1153684 | 22565.848 |  |  | 0.01 | 0.0097 | -2.97
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.69 | 23133.342 | 1179881 | 23133.342 |  |  | 0.01 | 0.00995 | -0.53
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.69 | 24355.15 | 1247996 | 24355.15 |  |  | 0.01 | 0.01047 | 4.72
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.69 | 24953.592 | 1265246 | 24953.592 |  |  | 0.01 | 0.01073 | 7.3
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26166.074 | 1344524 | 26166.074 |  |  | 0.01 | 0.01125 | 12.51
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26063.229 | 1338491 | 26063.229 |  |  | 0.01 | 0.01121 | 12.07
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.69 | 26872.852 | 1385291 | 26872.852 |  |  | 0.01 | 0.01155 | 15.55
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.69 | 22329.166 | 1134724 | 22329.166 |  |  | 0.01 | 0.0096 | -3.99
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.69 | 22994.361 | 1176313 | 22994.361 |  |  | 0.01 | 0.00989 | -1.13
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.69 | 23756.479 | 1219422 | 23756.479 |  |  | 0.01 | 0.01021 | 2.15
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.69 | 24666.17 | 1268269 | 24666.17 |  |  | 0.01 | 0.01061 | 6.06
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25000.137 | 1277774 | 25000.137 |  |  | 0.01 | 0.01075 | 7.5
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25239.68 | 1317405 | 25239.68 |  |  | 0.01 | 0.01085 | 8.53
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.69 | 25737.27 | 1325483 | 25737.27 |  |  | 0.01 | 0.01107 | 10.67
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.69 | 22016.998 | 1130137 | 22016.998 |  |  | 0.01 | 0.00947 | -5.33
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.69 | 22848.482 | 1164140 | 22848.482 |  |  | 0.01 | 0.00982 | -1.76
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.69 | 23377.701 | 1195150 | 23377.701 |  |  | 0.01 | 0.01005 | 0.52
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.69 | 24743.717 | 1282216 | 24743.717 |  |  | 0.01 | 0.01064 | 6.39
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25457.678 | 1306186 | 25457.678 |  |  | 0.01 | 0.01095 | 9.46
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25575.846 | 1318077 | 25575.846 |  |  | 0.01 | 0.011 | 9.97
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.69 | 25205.348 | 1300646 | 25205.348 |  |  | 0.01 | 0.01084 | 8.38
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.69 | 22032.42 | 1128433 | 22032.42 |  |  | 0.01 | 0.00947 | -5.26
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.69 | 22995.889 | 1178974 | 22995.889 |  |  | 0.01 | 0.00989 | -1.12
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.69 | 23030.449 | 1173826 | 23030.449 |  |  | 0.01 | 0.0099 | -0.97
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.69 | 24207.008 | 1245251 | 24207.008 |  |  | 0.01 | 0.01041 | 4.09
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.69 | 24884.777 | 1263125 | 24884.777 |  |  | 0.01 | 0.0107 | 7
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.69 | 25309.125 | 1302560 | 25309.125 |  |  | 0.01 | 0.01088 | 8.82
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.69 | 25344.141 | 1298578 | 25344.141 |  |  | 0.01 | 0.0109 | 8.98
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.69 | 22609.912 | 1152922 | 22609.912 |  |  | 0.01 | 0.00972 | -2.78
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.69 | 24554.58 | 1254301 | 24554.58 |  |  | 0.01 | 0.01056 | 5.58
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.69 | 21791.783 | 1105980 | 21791.783 |  |  | 0.01 | 0.00937 | -6.3
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.69 | 23955.301 | 1231520 | 23955.301 |  |  | 0.01 | 0.0103 | 3
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.69 | 21652.168 | 1101526 | 21652.168 |  |  | 0.01 | 0.00931 | -6.9
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.69 | 23626.941 | 1194339 | 23626.941 |  |  | 0.01 | 0.01016 | 1.59
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.69 | 21857.189 | 1100713 | 21857.189 |  |  | 0.01 | 0.0094 | -6.02
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.69 | 24218.33 | 1223239 | 24218.33 |  |  | 0.01 | 0.01041 | 4.13
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.69 | 22171.375 | 1122924 | 22171.375 |  |  | 0.01 | 0.00953 | -4.67
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.69 | 24038.738 | 1201898 | 24038.738 |  |  | 0.01 | 0.01034 | 3.36
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.69 | 21427.695 | 1074936 | 21427.695 |  |  | 0.01 | 0.00921 | -7.86
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.69 | 23164.816 | 1156204 | 23164.816 |  |  | 0.01 | 0.00996 | -0.4
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.69 | 20779.268 | 1021019 | 20779.268 |  |  | 0.01 | 0.00893 | -10.65
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.69 | 22853.115 | 1129490 | 22853.115 |  |  | 0.01 | 0.00983 | -1.74
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.69 | 19147.74 | 938578 | 19147.74 |  |  | 0.01 | 0.00823 | -17.67
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.69 | 20442.584 | 987649 | 20442.584 |  |  | 0.01 | 0.00879 | -12.1
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.69 | 18008.693 | 859605 | 18008.693 |  |  | 0.01 | 0.00774 | -22.57
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.69 | 19474.186 | 935603 | 19474.186 |  |  | 0.01 | 0.00837 | -16.26
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.69 | 17279.416 | 819807 | 17279.416 |  |  | 0.01 | 0.00743 | -25.7
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.69 | 18660.311 | 879229 | 18660.311 |  |  | 0.01 | 0.00802 | -19.76
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.69 | 15521.709 | 719428 | 15521.709 |  |  | 0.01 | 0.00667 | -33.26
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.69 | 16898.207 | 769962 | 16898.207 |  |  | 0.01 | 0.00727 | -27.34
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.69 | 21628.184 | 1107684 | 21628.184 |  |  | 0.01 | 0.0093 | -7
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.69 | 22818.307 | 1167930 | 22818.307 |  |  | 0.01 | 0.00981 | -1.89
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.69 | 22658.945 | 1182864 | 22658.945 |  |  | 0.01 | 0.00974 | -2.57
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.69 | 23242.805 | 1191134 | 23242.805 |  |  | 0.01 | 0.00999 | -0.06
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.69 | 24092.141 | 1241226 | 24092.141 |  |  | 0.01 | 0.01036 | 3.59
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.69 | 24051.273 | 1252444 | 24051.273 |  |  | 0.01 | 0.01034 | 3.42
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.69 | 24598.893 | 1287087 | 24598.893 |  |  | 0.01 | 0.01058 | 5.77
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.69 | 25048.125 | 1289327 | 25048.125 |  |  | 0.01 | 0.01077 | 7.7
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.69 | 8914.928 | 470050 | 8914.928 |  |  | 0.01 | 0.00383 | -61.67
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.69 | 22865.168 | 1162819 | 22865.168 |  |  | 0.01 | 0.00983 | -1.68
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.69 | 24217.922 | 1237288 | 24217.922 |  |  | 0.01 | 0.01041 | 4.13
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.69 | 24861.91 | 1255348 | 24861.91 |  |  | 0.01 | 0.01069 | 6.9
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.69 | 29108.41 | 1484585 | 29108.41 |  |  | 0.01 | 0.01252 | 25.16
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.69 | 23299.66 | 1180502 | 23299.66 |  |  | 0.01 | 0.01002 | 0.18
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.69 | 24923.5 | 1262333 | 24923.5 |  |  | 0.01 | 0.01072 | 7.17
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.69 | 21679.879 | 1091308 | 21679.879 |  |  | 0.01 | 0.00932 | -6.78
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.69 | 24550.033 | 1239441 | 24550.033 |  |  | 0.01 | 0.01056 | 5.56
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.69 | 25180.121 | 1308131 | 25180.121 |  |  | 0.01 | 0.01083 | 8.27
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.69 | 25625.738 | 1327128 | 25625.738 |  |  | 0.01 | 0.01102 | 10.19
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.69 | 24362.371 | 1265818 | 24362.371 |  |  | 0.01 | 0.01048 | 4.75
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.69 | 24305.764 | 1252778 | 24305.764 |  |  | 0.01 | 0.01045 | 4.51
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.69 | 24109.816 | 1246263 | 24109.816 |  |  | 0.01 | 0.01037 | 3.67
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.69 | 24456.668 | 1274867 | 24456.668 |  |  | 0.01 | 0.01052 | 5.16
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.69 | 24957.41 | 1282631 | 24957.41 |  |  | 0.01 | 0.01073 | 7.31
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.69 | 25152.914 | 1299701 | 25152.914 |  |  | 0.01 | 0.01082 | 8.15
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.69 | 24200.555 | 1242784 | 24200.555 |  |  | 0.01 | 0.01041 | 4.06
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.69 | 22887.045 | 1178188 | 22887.045 |  |  | 0.01 | 0.00984 | -1.59
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.69 | 22461.596 | 1154245 | 22461.596 |  |  | 0.01 | 0.00966 | -3.42
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.69 | 23178.355 | 1180150 | 23178.355 |  |  | 0.01 | 0.00997 | -0.34
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.69 | 22570.561 | 1168954 | 22570.561 |  |  | 0.01 | 0.0097 | -2.95
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.69 | 23917.525 | 1244149 | 23917.525 |  |  | 0.01 | 0.01028 | 2.84
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.69 | 23604.629 | 1226601 | 23604.629 |  |  | 0.01 | 0.01015 | 1.5
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.69 | 17613.797 | 840636 | 17613.797 |  |  | 0.01 | 0.00757 | -24.26
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.69 | 17847.734 | 849091 | 17847.734 |  |  | 0.01 | 0.00767 | -23.26
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.69 | 18605.641 | 878553 | 18605.641 |  |  | 0.01 | 0.008 | -20
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.69 | 26959.02 | 1406070 | 26959.02 |  |  | 0.01 | 0.01159 | 15.92
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.69 | 23924.834 | 1249809 | 23924.834 |  |  | 0.01 | 0.01029 | 2.87
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.69 | 24150.537 | 1252542 | 24150.537 |  |  | 0.01 | 0.01038 | 3.84
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.69 | 23967.572 | 1236497 | 23967.572 |  |  | 0.01 | 0.01031 | 3.06
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.69 | 23285.992 | 1205143 | 23285.992 |  |  | 0.01 | 0.01001 | 0.13
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.69 | 24485.539 | 1255107 | 24485.539 |  |  | 0.01 | 0.01053 | 5.28
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.69 | 19292.301 | 915616 | 19292.301 |  |  | 0.01 | 0.0083 | -17.05
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.69 | 18221.59 | 863347 | 18221.59 |  |  | 0.01 | 0.00783 | -21.65
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.69 | 18441.566 | 862143 | 18441.566 |  |  | 0.01 | 0.00793 | -20.7

## DTXSID1037303
*Analyte in Mix 1 | [8 empty] | PFHpA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.76 | 850.71 | 43330 | 66855.516 | 0 | 0.998 | 0.17 | 0.14193 | -16.51
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.76 | 878.514 | 46919 | 70140.469 | 0 | 0.998 | 0.17 | 0.13929 | -18.07
UC_AF_Mix1_A | 0.10463 | 3.3482 | 0.00033564 | 0.0003581 | 6.2998e-05 | 17.59% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.76 | 910.793 | 47778 | 69550.805 | 0 | 0.998 | 0.17 | 0.14684 | -13.62
UC_AF_Mix1_B | 0.10449 | 3.3437 | 0.00030941 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.76 | 965.351 | 49330 | 73650.867 | 0 | 0.998 | 0.17 | 0.147 | -13.53
UC_AF_Mix1_C | 0.14318 | 4.5818 | 0.00042925 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.76 | 914.414 | 47027 | 75592.711 | 0 | 0.998 | 0.17 | 0.13361 | -21.41
UC_T1hr_Mix1_A | 123.9977 | 9919.8128 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.76 | 931.982 | 45388 | 75935.922 | 0 | 0.998 | 0.17 | 0.13595 | -20.03
UC_T1hr_Mix1_B | 135.4343 | 10834.7448 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.76 | 1011.845 | 51110 | 74628.094 | 0 | 0.998 | 0.17 | 0.15298 | -10.01
UC_T1hr_Mix1_C | 131.4578 | 10516.6224 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1596.522 | 80606 | 70175.805 | 0 | 0.998 | 0.28 | 0.2748 | -1.86
UC_T5hr_Mix1_A | 124.6923 | 9975.384 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1666.76 | 81526 | 73291.031 | 0 | 0.998 | 0.28 | 0.27469 | -1.9
UC_T5hr_Mix1_B | 135.0841 | 10806.7272 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1679.931 | 85479 | 72982.797 | 0 | 0.998 | 0.28 | 0.27835 | -0.59
UC_T5hr_Mix1_C | 133.4238 | 10673.9024 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1788.739 | 90149 | 75305.336 | 0 | 0.998 | 0.28 | 0.28809 | 2.89
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1758.781 | 87451 | 78736.781 | 0 | 0.998 | 0.28 | 0.26933 | -3.81
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1817.852 | 89845 | 77973.492 | 0 | 0.998 | 0.28 | 0.28227 | 0.81
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.76 | 1778.642 | 85922 | 79756.063 | 0 | 0.998 | 0.28 | 0.26885 | -3.98
20201123_PFAS_PPB_UC_Sample008 | CC2 | 0.28 | 0.2748 | 0.27663 | 0.0069165 | 2.5% | 0.021739 | -1.205% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2553.536 | 127184 | 69660.461 | 0 | 0.998 | 0.44 | 0.4591 | 4.34
20201123_PFAS_PPB_UC_Sample033 | CC2 | 0.28 | 0.27469 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2730.885 | 138122 | 71199 | 0 | 0.998 | 0.44 | 0.48161 | 9.46
20201123_PFAS_PPB_UC_Sample040 | CC2 | 0.28 | 0.27835 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2804.495 | 142603 | 73203.641 | 0 | 0.998 | 0.44 | 0.48102 | 9.32
20201123_PFAS_PPB_UC_Sample080 | CC2 | 0.28 | 0.28809 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3039.829 | 147500 | 75526.563 | 0 | 0.998 | 0.44 | 0.5067 | 15.16
20201123_PFAS_PPB_UC_Sample106 | CC2 | 0.28 | 0.26933 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3006.335 | 153274 | 78371.031 | 0 | 0.998 | 0.44 | 0.48167 | 9.47
20201123_PFAS_PPB_UC_Sample113 | CC2 | 0.28 | 0.28227 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.76 | 2829.649 | 139779 | 77142.375 | 0 | 0.998 | 0.44 | 0.45942 | 4.41
20201123_PFAS_PPB_UC_Sample120 | CC2 | 0.28 | 0.26885 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.76 | 3093.993 | 153457 | 79699.555 | 0 | 0.998 | 0.44 | 0.48777 | 10.86
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.76 | 3685.775 | 190921 | 69179.859 | 0.001 | 0.998 | 0.71 | 0.67937 | -4.31
Blanks should be less than… | 0.010869 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4133.796 | 204342 | 70969.344 | 0.001 | 0.998 | 0.71 | 0.74522 | 4.96
eLOQ | 0.28 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.76 | 3977.837 | 198806 | 71708.977 | 0.001 | 0.998 | 0.71 | 0.70844 | -0.22
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4200.137 | 215512 | 74249.516 | 0.001 | 0.998 | 0.71 | 0.72296 | 1.83
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4212.83 | 209553 | 74719.867 | 0.001 | 0.998 | 0.71 | 0.7205 | 1.48
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4355.408 | 219745 | 76099.219 | 0.001 | 0.998 | 0.71 | 0.73178 | 3.07
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.76 | 4305.967 | 212324 | 76448.648 | 0.001 | 0.998 | 0.71 | 0.71974 | 1.37
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.76 | 5816.126 | 293972 | 65949.148 | 0.001 | 0.998 | 1.14 | 1.142 | 0.18
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6447.323 | 312789 | 69821.609 | 0.001 | 0.998 | 1.14 | 1.197 | 5
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6033.968 | 310688 | 69097.773 | 0.001 | 0.998 | 1.14 | 1.1305 | -0.83
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6841.671 | 335362 | 74140.367 | 0.001 | 0.998 | 1.14 | 1.1962 | 4.93
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6562.722 | 325605 | 75418.586 | 0.001 | 0.998 | 1.14 | 1.1265 | -1.19
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6592.64 | 331029 | 76403.75 | 0.001 | 0.998 | 1.14 | 1.1168 | -2.04
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.76 | 6798.62 | 334024 | 74293.961 | 0.001 | 0.998 | 1.14 | 1.186 | 4.04
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.76 | 9600.393 | 482368 | 65820.242 | 0.001 | 0.998 | 1.82 | 1.9061 | 4.73
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.76 | 9697.646 | 488360 | 70147.031 | 0.001 | 0.998 | 1.82 | 1.8053 | -0.81
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10209.131 | 501514 | 69563.125 | 0.001 | 0.998 | 1.82 | 1.9181 | 5.39
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10747.685 | 524962 | 72727.977 | 0.001 | 0.998 | 1.82 | 1.9316 | 6.13
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10626.602 | 533370 | 72627.703 | 0.001 | 0.998 | 1.82 | 1.9122 | 5.07
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10844.135 | 536457 | 72442.719 | 0.001 | 0.998 | 1.82 | 1.9569 | 7.52
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.76 | 10520.271 | 523345 | 75961.703 | 0.001 | 0.998 | 1.82 | 1.8086 | -0.63
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.76 | 13913.541 | 696557 | 69313.656 | 0.002 | 0.998 | 2.91 | 2.6332 | -9.51
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.76 | 15120.782 | 752900 | 73489.727 | 0.002 | 0.998 | 2.91 | 2.6997 | -7.23
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.76 | 22342.369 | 1126751 | 67371.406 | 0.003 | 0.998 | 4.66 | 4.3672 | -6.28
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.76 | 25136.83 | 1237218 | 72791.852 | 0.003 | 0.998 | 4.66 | 4.5486 | -2.39
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.76 | 36733.152 | 1808733 | 65924.273 | 0.006 | 0.998 | 7.45 | 7.3546 | -1.28
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.76 | 38859.742 | 1924695 | 69266.383 | 0.006 | 0.998 | 7.45 | 7.4051 | -0.6
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.76 | 64327.168 | 3108039 | 66429.594 | 0.01 | 0.998 | 11.92 | 12.7966 | 7.35
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.76 | 69992.758 | 3395560 | 72812.047 | 0.01 | 0.998 | 11.92 | 12.703 | 6.57
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.76 | 110352.383 | 5346011 | 67007.094 | 0.016 | 0.998 | 19.07 | 21.7694 | 14.16
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.76 | 116698.813 | 5599537 | 70597.922 | 0.017 | 0.998 | 19.07 | 21.8504 | 14.58
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.76 | 157745.297 | 7601439 | 65890.977 | 0.024 | 0.998 | 30.52 | 31.638 | 3.66
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.76 | 165116.125 | 7948942 | 69640.266 | 0.024 | 0.998 | 30.52 | 31.3337 | 2.67
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.76 | 234145.531 | 11241628 | 64077.266 | 0.037 | 0.998 | 48.83 | 48.2531 | -1.18
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.76 | 253357.641 | 12204068 | 67770.406 | 0.037 | 0.998 | 48.83 | 49.3642 | 1.09
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.76 | 359678.344 | 16933280 | 59706.527 | 0.06 | 0.998 | 78.13 | 79.4088 | 1.64
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.76 | 367405.875 | 17545514 | 61003.473 | 0.06 | 0.998 | 78.13 | 79.3904 | 1.61
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.76 | 507858.75 | 23705804 | 56827.754 | 0.089 | 0.998 | 125 | 117.5312 | -5.98
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.76 | 530252.313 | 24655996 | 59756.648 | 0.089 | 0.998 | 125 | 116.705 | -6.64
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.76 | 645238.188 | 29797752 | 54128.734 | 0.119 | 0.998 | 156.25 | 156.393 | 0.09
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.76 | 667760.563 | 30854474 | 57883.109 | 0.115 | 0.998 | 156.25 | 151.4011 | -3.1
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.76 | 984418 | 44298656 | 50352.082 | 0.196 | 0.998 | 250 | 254.9259 | 1.97
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.76 | 1034223.938 | 46612284 | 53039.328 | 0.195 | 0.998 | 250 | 254.2649 | 1.71
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 1.76 | 20.155 | 721 |  |  | 0.998
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 1.77 | 20.274 | 901 |  |  | 0.998
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 1.77 | 28.679 | 1180 |  |  | 0.998
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 1.76 | 10.256 | 686 |  |  | 0.998
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.76 | 113.106 | 5453 | 66453.484 | 0 | 0.998
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.76 | 121.095 | 5130 | 70428.734 | 0 | 0.998
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.76 | 78.945 | 3535 | 70267.711 | 0 | 0.998
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.76 | 104.814 | 5277 | 71625.82 | 0 | 0.998
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.76 | 142.571 | 5999 | 73164.828 | 0 | 0.998
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.76 | 122.398 | 6077 | 72886.563 | 0 | 0.998
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.76 | 135.832 | 6397 | 74817.703 | 0 | 0.998
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.76 | 131.084 | 6036 | 77838.516 | 0 | 0.998
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.76 | 43.671 | 1682 | 26529.941 | 0 | 0.998
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.76 | 3246.877 | 163527 | 69882.695 | 0 | 0.998 | 0.63 | 0.58904 | -6.5
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.76 | 3587.323 | 179450 | 73898.461 | 0 | 0.998 | 0.63 | 0.61663 | -2.12
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.76 | 15453.368 | 764193 | 74668.664 | 0.002 | 0.998 | 2.5 | 2.7157 | 8.63
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.76 | 17789.33 | 869332 | 86813.352 | 0.002 | 0.998 | 2.5 | 2.6886 | 7.54
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.76 | 34994.387 | 1718356 | 71009.594 | 0.005 | 0.998 | 6.25 | 6.5019 | 4.03
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.76 | 37541.043 | 1824743 | 74901.398 | 0.005 | 0.998 | 6.25 | 6.6131 | 5.81
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.76 | 131310.563 | 6424507 | 66898.727 | 0.02 | 0.998 | 25 | 25.944 | 3.78
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.76 | 141971.547 | 6814487 | 73049.375 | 0.019 | 0.998 | 25 | 25.6887 | 2.75
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 1.82 | 97.424 | 2181 |  |  | 0.998
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 1.83 | 143.086 | 2718 |  |  | 0.998
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 1.83 | 82.237 | 2157 |  |  | 0.998
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 1.76 | 19.311 | 622 |  |  | 0.998
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 1.75 | 124.244 | 4163 |  |  | 0.998
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 1.76 | 18.022 | 999 |  |  | 0.998
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 1.8 | 4.15 | 234 |  |  | 0.998
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 1.76 | 88.152 | 2940 |  |  | 0.998
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 1.76 | 17.823 | 784 |  |  | 0.998
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 1.75 | 151.183 | 4824 |  |  | 0.998
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 1.76 | 26.737 | 1220 |  |  | 0.998
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 1.81 | 127.509 | 2806 |  |  | 0.998
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 1.76 | 26.945 | 1365 |  |  | 0.998
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 1.77 | 17.491 | 895 |  |  | 0.998
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 1.78 | 9.957 | 400 |  |  | 0.998
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.76 | 751.544 | 39201 | 75833.383 | 0 | 0.998 |  | 0.10463
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.76 | 777.696 | 39627 | 78561.383 | 0 | 0.998 |  | 0.10449
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.76 | 971.488 | 49541 | 75786.555 | 0 | 0.998 |  | 0.14318
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.76 | 70.381 | 3697 | 75532.336 | 0 | 0.998
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.76 | 115.072 | 5839 | 73999.875 | 0 | 0.998
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.76 | 102.199 | 4557 | 73473.383 | 0 | 0.998
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.76 | 70.733 | 3571 | 72956.234 | 0 | 0.998
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.76 | 93.345 | 4359 | 75183.023 | 0 | 0.998
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.76 | 103.114 | 4817 | 72624.836 | 0 | 0.998
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.76 | 545587.125 | 25399062 | 57842.676 | 0.094 | 0.998 |  | 123.9977
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.76 | 576372.25 | 26675222 | 55906.832 | 0.103 | 0.998 |  | 135.4343
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.76 | 579618.688 | 27357730 | 57936.676 | 0.1 | 0.998 |  | 131.4578
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.76 | 140.486 | 7038 | 68748.156 | 0 | 0.998 |  | 0.00037
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.76 | 249.245 | 12045 | 73999.633 | 0 | 0.998 |  | 0.01793
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.76 | 181.4 | 9688 | 71143.813 | 0 | 0.998 |  | 0.00708
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.76 | 129.401 | 6930 | 69815.391 | 0 | 0.998
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.76 | 105.802 | 4441 | 71290.555 | 0 | 0.998
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.76 | 108.512 | 4980 | 74547.953 | 0 | 0.998
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.76 | 616916.688 | 28676648 | 65037.824 | 0.095 | 0.998 |  | 124.6923
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.76 | 599925.25 | 27823716 | 58343.555 | 0.103 | 0.998 |  | 135.0841
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.76 | 600657.063 | 28247480 | 59147.715 | 0.102 | 0.998 |  | 133.4238
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.76 | 254.678 | 11025 | 71500.539 | 0 | 0.998 |  | 0.0205
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.76 | 239.811 | 11703 | 70718.625 | 0 | 0.998 |  | 0.01823
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.76 | 172.22 | 8458 | 71977.289 | 0 | 0.998 |  | 0.005
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.76 | 108.144 | 4597 | 75227.086 | 0 | 0.998
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.76 | 92.977 | 4295 | 72306.594 | 0 | 0.998
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.76 | 107.86 | 5319 | 72174.523 | 0 | 0.998
M4PFHpA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 1.76 | 66855.516 | 3320399 | 66855.516 |  |  | 0.01 | 0.00952 | -4.75
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 1.76 | 70140.469 | 3433732 | 70140.469 |  |  | 0.01 | 0.00999 | -0.07
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 1.76 | 69550.805 | 3433691 | 69550.805 |  |  | 0.01 | 0.00991 | -0.91
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 1.76 | 73650.867 | 3635436 | 73650.867 |  |  | 0.01 | 0.01049 | 4.93
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 1.76 | 75592.711 | 3724434 | 75592.711 |  |  | 0.01 | 0.01077 | 7.7
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 1.76 | 75935.922 | 3752660 | 75935.922 |  |  | 0.01 | 0.01082 | 8.18
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 1.76 | 74628.094 | 3664734 | 74628.094 |  |  | 0.01 | 0.01063 | 6.32
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 1.76 | 70175.805 | 3509487 | 70175.805 |  |  | 0.01 | 0.01 | -0.02
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 1.76 | 73291.031 | 3595913 | 73291.031 |  |  | 0.01 | 0.01044 | 4.42
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 1.76 | 72982.797 | 3554466 | 72982.797 |  |  | 0.01 | 0.0104 | 3.98
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 1.76 | 75305.336 | 3690926 | 75305.336 |  |  | 0.01 | 0.01073 | 7.29
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 1.76 | 78736.781 | 3888135 | 78736.781 |  |  | 0.01 | 0.01122 | 12.17
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 1.76 | 77973.492 | 3802600 | 77973.492 |  |  | 0.01 | 0.01111 | 11.09
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 1.76 | 79756.063 | 3861355 | 79756.063 |  |  | 0.01 | 0.01136 | 13.63
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 1.76 | 69660.461 | 3393695 | 69660.461 |  |  | 0.01 | 0.00992 | -0.76
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 1.76 | 71199 | 3475520 | 71199 |  |  | 0.01 | 0.01014 | 1.44
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 1.76 | 73203.641 | 3581378 | 73203.641 |  |  | 0.01 | 0.01043 | 4.29
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 1.76 | 75526.563 | 3681046 | 75526.563 |  |  | 0.01 | 0.01076 | 7.6
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 1.76 | 78371.031 | 3866716 | 78371.031 |  |  | 0.01 | 0.01117 | 11.65
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 1.76 | 77142.375 | 3761686 | 77142.375 |  |  | 0.01 | 0.01099 | 9.9
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 1.76 | 79699.555 | 3884271 | 79699.555 |  |  | 0.01 | 0.01135 | 13.55
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 1.76 | 69179.859 | 3379878 | 69179.859 |  |  | 0.01 | 0.00986 | -1.44
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 1.76 | 70969.344 | 3452449 | 70969.344 |  |  | 0.01 | 0.01011 | 1.11
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 1.76 | 71708.977 | 3520012 | 71708.977 |  |  | 0.01 | 0.01022 | 2.16
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 1.76 | 74249.516 | 3651160 | 74249.516 |  |  | 0.01 | 0.01058 | 5.78
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 1.76 | 74719.867 | 3721994 | 74719.867 |  |  | 0.01 | 0.01065 | 6.45
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 1.76 | 76099.219 | 3714412 | 76099.219 |  |  | 0.01 | 0.01084 | 8.42
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 1.76 | 76448.648 | 3750645 | 76448.648 |  |  | 0.01 | 0.01089 | 8.91
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 1.76 | 65949.148 | 3240065 | 65949.148 |  |  | 0.01 | 0.0094 | -6.04
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 1.76 | 69821.609 | 3399067 | 69821.609 |  |  | 0.01 | 0.00995 | -0.53
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 1.76 | 69097.773 | 3396196 | 69097.773 |  |  | 0.01 | 0.00984 | -1.56
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 1.76 | 74140.367 | 3651768 | 74140.367 |  |  | 0.01 | 0.01056 | 5.63
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 1.76 | 75418.586 | 3694990 | 75418.586 |  |  | 0.01 | 0.01074 | 7.45
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 1.76 | 76403.75 | 3709347 | 76403.75 |  |  | 0.01 | 0.01089 | 8.85
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 1.76 | 74293.961 | 3689391 | 74293.961 |  |  | 0.01 | 0.01058 | 5.84
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 1.76 | 65820.242 | 3263160 | 65820.242 |  |  | 0.01 | 0.00938 | -6.23
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 1.76 | 70147.031 | 3440420 | 70147.031 |  |  | 0.01 | 0.00999 | -0.06
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 1.76 | 69563.125 | 3425429 | 69563.125 |  |  | 0.01 | 0.00991 | -0.9
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72727.977 | 3560901 | 72727.977 |  |  | 0.01 | 0.01036 | 3.61
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72627.703 | 3593555 | 72627.703 |  |  | 0.01 | 0.01035 | 3.47
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 1.76 | 72442.719 | 3600272 | 72442.719 |  |  | 0.01 | 0.01032 | 3.21
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 1.76 | 75961.703 | 3739865 | 75961.703 |  |  | 0.01 | 0.01082 | 8.22
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 1.76 | 69313.656 | 3428823 | 69313.656 |  |  | 0.01 | 0.00987 | -1.25
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 1.76 | 73489.727 | 3628375 | 73489.727 |  |  | 0.01 | 0.01047 | 4.7
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 1.76 | 67371.406 | 3263118 | 67371.406 |  |  | 0.01 | 0.0096 | -4.02
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 1.76 | 72791.852 | 3591268 | 72791.852 |  |  | 0.01 | 0.01037 | 3.7
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 1.76 | 65924.273 | 3208166 | 65924.273 |  |  | 0.01 | 0.00939 | -6.08
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 1.76 | 69266.383 | 3432012 | 69266.383 |  |  | 0.01 | 0.00987 | -1.32
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 1.76 | 66429.594 | 3218192 | 66429.594 |  |  | 0.01 | 0.00946 | -5.36
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 1.76 | 72812.047 | 3563225 | 72812.047 |  |  | 0.01 | 0.01037 | 3.73
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 1.76 | 67007.094 | 3238182 | 67007.094 |  |  | 0.01 | 0.00955 | -4.54
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 1.76 | 70597.922 | 3442672 | 70597.922 |  |  | 0.01 | 0.01006 | 0.58
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 1.76 | 65890.977 | 3142368 | 65890.977 |  |  | 0.01 | 0.00939 | -6.13
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 1.76 | 69640.266 | 3376419 | 69640.266 |  |  | 0.01 | 0.00992 | -0.79
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 1.76 | 64077.266 | 3046624 | 64077.266 |  |  | 0.01 | 0.00913 | -8.71
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 1.76 | 67770.406 | 3271291 | 67770.406 |  |  | 0.01 | 0.00966 | -3.45
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 1.76 | 59706.527 | 2813015 | 59706.527 |  |  | 0.01 | 0.00851 | -14.94
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 1.76 | 61003.473 | 2885865 | 61003.473 |  |  | 0.01 | 0.00869 | -13.09
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 1.76 | 56827.754 | 2640782 | 56827.754 |  |  | 0.01 | 0.0081 | -19.04
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 1.76 | 59756.648 | 2784971 | 59756.648 |  |  | 0.01 | 0.00851 | -14.87
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 1.76 | 54128.734 | 2507089 | 54128.734 |  |  | 0.01 | 0.00771 | -22.88
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 1.76 | 57883.109 | 2654882 | 57883.109 |  |  | 0.01 | 0.00825 | -17.54
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 1.76 | 50352.082 | 2284973 | 50352.082 |  |  | 0.01 | 0.00717 | -28.26
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 1.76 | 53039.328 | 2395597 | 53039.328 |  |  | 0.01 | 0.00756 | -24.44
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 1.76 | 66453.484 | 3305070 | 66453.484 |  |  | 0.01 | 0.00947 | -5.33
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 1.76 | 70428.734 | 3462973 | 70428.734 |  |  | 0.01 | 0.01003 | 0.34
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 1.76 | 70267.711 | 3467658 | 70267.711 |  |  | 0.01 | 0.01001 | 0.11
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 1.76 | 71625.82 | 3505632 | 71625.82 |  |  | 0.01 | 0.0102 | 2.04
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 1.76 | 73164.828 | 3636845 | 73164.828 |  |  | 0.01 | 0.01042 | 4.24
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 1.76 | 72886.563 | 3650300 | 72886.563 |  |  | 0.01 | 0.01038 | 3.84
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 1.76 | 74817.703 | 3688547 | 74817.703 |  |  | 0.01 | 0.01066 | 6.59
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 1.76 | 77838.516 | 3829249 | 77838.516 |  |  | 0.01 | 0.01109 | 10.89
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 1.76 | 26529.941 | 1325094 | 26529.941 |  |  | 0.01 | 0.00378 | -62.2
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 1.76 | 69882.695 | 3425608 | 69882.695 |  |  | 0.01 | 0.00996 | -0.44
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 1.76 | 73898.461 | 3579202 | 73898.461 |  |  | 0.01 | 0.01053 | 5.28
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 1.76 | 74668.664 | 3631655 | 74668.664 |  |  | 0.01 | 0.01064 | 6.38
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 1.76 | 86813.352 | 4249834 | 86813.352 |  |  | 0.01 | 0.01237 | 23.68
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 1.76 | 71009.594 | 3416943 | 71009.594 |  |  | 0.01 | 0.01012 | 1.17
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 1.76 | 74901.398 | 3616913 | 74901.398 |  |  | 0.01 | 0.01067 | 6.71
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 1.76 | 66898.727 | 3226693 | 66898.727 |  |  | 0.01 | 0.00953 | -4.69
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 1.76 | 73049.375 | 3523750 | 73049.375 |  |  | 0.01 | 0.01041 | 4.07
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 1.76 | 75833.383 | 3743159 | 75833.383 |  |  | 0.01 | 0.0108 | 8.04
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 1.76 | 78561.383 | 3845300 | 78561.383 |  |  | 0.01 | 0.01119 | 11.92
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 1.76 | 75786.555 | 3701649 | 75786.555 |  |  | 0.01 | 0.0108 | 7.97
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 1.76 | 75532.336 | 3711094 | 75532.336 |  |  | 0.01 | 0.01076 | 7.61
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 1.76 | 73999.875 | 3664879 | 73999.875 |  |  | 0.01 | 0.01054 | 5.43
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 1.76 | 73473.383 | 3686369 | 73473.383 |  |  | 0.01 | 0.01047 | 4.68
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 1.76 | 72956.234 | 3652750 | 72956.234 |  |  | 0.01 | 0.01039 | 3.94
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 1.76 | 75183.023 | 3684570 | 75183.023 |  |  | 0.01 | 0.01071 | 7.11
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 1.76 | 72624.836 | 3574495 | 72624.836 |  |  | 0.01 | 0.01035 | 3.47
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 1.76 | 57842.676 | 2695074 | 57842.676 |  |  | 0.01 | 0.00824 | -17.59
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 1.76 | 55906.832 | 2600066 | 55906.832 |  |  | 0.01 | 0.00796 | -20.35
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 1.76 | 57936.676 | 2688977 | 57936.676 |  |  | 0.01 | 0.00825 | -17.46
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 1.76 | 68748.156 | 3423723 | 68748.156 |  |  | 0.01 | 0.00979 | -2.06
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 1.76 | 73999.633 | 3635966 | 73999.633 |  |  | 0.01 | 0.01054 | 5.43
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 1.76 | 71143.813 | 3516017 | 71143.813 |  |  | 0.01 | 0.01014 | 1.36
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 1.76 | 69815.391 | 3450382 | 69815.391 |  |  | 0.01 | 0.00995 | -0.54
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 1.76 | 71290.555 | 3511303 | 71290.555 |  |  | 0.01 | 0.01016 | 1.57
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 1.76 | 74547.953 | 3626257 | 74547.953 |  |  | 0.01 | 0.01062 | 6.21
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 1.76 | 65037.824 | 3074126 | 65037.824 |  |  | 0.01 | 0.00927 | -7.34
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 1.76 | 58343.555 | 2721553 | 58343.555 |  |  | 0.01 | 0.00831 | -16.88
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 1.76 | 59147.715 | 2757025 | 59147.715 |  |  | 0.01 | 0.00843 | -15.73
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 1.76 | 71500.539 | 3564321 | 71500.539 |  |  | 0.01 | 0.01019 | 1.87
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 1.76 | 70718.625 | 3511219 | 70718.625 |  |  | 0.01 | 0.01008 | 0.75
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 1.76 | 71977.289 | 3577291 | 71977.289 |  |  | 0.01 | 0.01025 | 2.54
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 1.76 | 75227.086 | 3777262 | 75227.086 |  |  | 0.01 | 0.01072 | 7.17
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 1.76 | 72306.594 | 3534427 | 72306.594 |  |  | 0.01 | 0.0103 | 3.01
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 1.76 | 72174.523 | 3583457 | 72174.523 |  |  | 0.01 | 0.01028 | 2.83

## DTXSID3031860
*Analyte in Mix 1 | [8 empty] | PFDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.29 | 812.475 | 26234 | 37339.707 | 0 | 0.998 | 0.17 | 0.1442 | -15.17
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.29 | 715.83 | 22620 | 37417.016 | 0 | 0.998 | 0.17 | 0.12157 | -28.49
UC_AF_Mix1_A | 0.88289 | 28.2525 | 0.0028133 | 0.0027497 | 0.00023113 | 8.406% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.29 | 793.24 | 25147 | 38646.242 | 0 | 0.998 | 0.17 | 0.13358 | -21.42
UC_AF_Mix1_B | 0.98923 | 31.6554 | 0.0029423 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.29 | 775.351 | 24917 | 39172.805 | 0 | 0.998 | 0.17 | 0.12727 | -25.13
UC_AF_Mix1_C | 0.83289 | 26.6525 | 0.0024934 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.29 | 844.023 | 28432 | 41901.957 | 0 | 0.998 | 0.17 | 0.13029 | -23.36
UC_T1hr_Mix1_A | 122.7682 | 9821.4552 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.29 | 892.164 | 29991 | 40732.211 | 0 | 0.998 | 0.17 | 0.14545 | -14.44
UC_T1hr_Mix1_B | 131.3548 | 10508.388 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.29 | 729.499 | 23738 | 41387.645 | 0 | 0.998 | 0.17 | 0.10861 | -36.11
UC_T1hr_Mix1_C | 135.2349 | 10818.792 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1422.246 | 46842 | 37559.73 | 0 | 0.998 | 0.28 | 0.28291 | 1.04
UC_T5hr_Mix1_A | 125.5328 | 10042.6256 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1486.707 | 49406 | 40856.688 | 0 | 0.998 | 0.28 | 0.27018 | -3.51
UC_T5hr_Mix1_B | 134.4819 | 10758.5544 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1580.923 | 52199 | 42048.52 | 0 | 0.998 | 0.28 | 0.28059 | 0.21
UC_T5hr_Mix1_C | 133.6157 | 10689.2536 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1632.62 | 54329 | 40422.34 | 0 | 0.998 | 0.28 | 0.30463 | 8.8
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1627.655 | 53157 | 44425.273 | 0 | 0.998 | 0.28 | 0.27233 | -2.74
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1658.027 | 53910 | 43092.957 | 0 | 0.998 | 0.28 | 0.28815 | 2.91
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.29 | 1746.708 | 56488 | 45184.34 | 0 | 0.998 | 0.28 | 0.28972 | 3.47
20201123_PFAS_PPB_UC_Sample008 | CC2 | 0.28 | 0.28291 | 0.28407 | 0.011658 | 4.104% | 0.036642 | 1.455% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2425.374 | 80815 | 38314.09 | 0.001 | 0.998 | 0.44 | 0.50194 | 14.08
20201123_PFAS_PPB_UC_Sample033 | CC2 | 0.28 | 0.27018 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2629.677 | 86066 | 39381.18 | 0.001 | 0.998 | 0.44 | 0.53184 | 20.87
20201123_PFAS_PPB_UC_Sample040 | CC2 | 0.28 | 0.28059 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2479.736 | 80126 | 39681.668 | 0.001 | 0.998 | 0.44 | 0.49495 | 12.49
20201123_PFAS_PPB_UC_Sample080 | CC2 | 0.28 | 0.30463 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2642.395 | 87039 | 41193.734 | 0.001 | 0.998 | 0.44 | 0.5092 | 15.73
20201123_PFAS_PPB_UC_Sample106 | CC2 | 0.28 | 0.27233 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2656.241 | 86361 | 41899.441 | 0.001 | 0.998 | 0.44 | 0.50274 | 14.26
20201123_PFAS_PPB_UC_Sample113 | CC2 | 0.28 | 0.28815 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2750.117 | 90688 | 42587.418 | 0.001 | 0.998 | 0.44 | 0.51291 | 16.57
20201123_PFAS_PPB_UC_Sample120 | CC2 | 0.28 | 0.28972 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.29 | 2783.193 | 90359 | 42725.34 | 0.001 | 0.998 | 0.44 | 0.51778 | 17.68
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3344.134 | 108165 | 38319.367 | 0.001 | 0.998 | 0.71 | 0.70833 | -0.24
Blanks should be less than… | 0.018321 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3486.108 | 112718 | 39899.059 | 0.001 | 0.998 | 0.71 | 0.70921 | -0.11
eLOQ | 0.28 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3425.66 | 112990 | 38738.613 | 0.001 | 0.998 | 0.71 | 0.71832 | 1.17
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3584.333 | 119301 | 41931.879 | 0.001 | 0.998 | 0.71 | 0.69291 | -2.41
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3708.542 | 124160 | 42107.133 | 0.001 | 0.998 | 0.71 | 0.71525 | 0.74
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3710.15 | 123010 | 43136.676 | 0.001 | 0.998 | 0.71 | 0.69747 | -1.76
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.29 | 3690.172 | 122912 | 41798.488 | 0.001 | 0.998 | 0.71 | 0.71706 | 0.99
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5157.513 | 167565 | 35923.289 | 0.001 | 0.998 | 1.14 | 1.1931 | 4.66
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5235.55 | 174282 | 39440.199 | 0.001 | 0.998 | 1.14 | 1.0999 | -3.52
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5363 | 172030 | 37568.254 | 0.001 | 0.998 | 1.14 | 1.1861 | 4.04
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5677.545 | 185915 | 40251.609 | 0.001 | 0.998 | 1.14 | 1.1714 | 2.76
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5653.892 | 185675 | 41405.266 | 0.001 | 0.998 | 1.14 | 1.1327 | -0.64
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5641.129 | 182794 | 39118.82 | 0.001 | 0.998 | 1.14 | 1.1986 | 5.14
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.29 | 5829.492 | 183732 | 41191.141 | 0.001 | 0.998 | 1.14 | 1.1755 | 3.11
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.29 | 7982.595 | 263229 | 36542.281 | 0.002 | 0.998 | 1.82 | 1.8378 | 0.98
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8397.799 | 272370 | 39243.359 | 0.002 | 0.998 | 1.82 | 1.7995 | -1.13
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8747.275 | 279911 | 38877.977 | 0.002 | 0.998 | 1.82 | 1.8942 | 4.08
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.29 | 8584.184 | 278202 | 39630.223 | 0.002 | 0.998 | 1.82 | 1.822 | 0.11
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9366.021 | 295213 | 41834.133 | 0.002 | 0.998 | 1.82 | 1.8846 | 3.55
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9061.154 | 297179 | 40277.5 | 0.002 | 0.998 | 1.82 | 1.894 | 4.06
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.29 | 9407.359 | 295921 | 42275.109 | 0.002 | 0.998 | 1.82 | 1.873 | 2.91
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.29 | 12157.958 | 400542 | 37978.066 | 0.003 | 0.998 | 2.91 | 2.7133 | -6.76
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.29 | 13500.459 | 429025 | 41927.211 | 0.003 | 0.998 | 2.91 | 2.7294 | -6.21
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.29 | 19602.594 | 620168 | 37018.879 | 0.005 | 0.998 | 4.66 | 4.516 | -3.09
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.29 | 20670.557 | 664651 | 40052.781 | 0.005 | 0.998 | 4.66 | 4.4003 | -5.57
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.29 | 30781.998 | 977786 | 34951.309 | 0.009 | 0.998 | 7.45 | 7.5387 | 1.19
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.29 | 34151.543 | 1101646 | 40052.25 | 0.009 | 0.998 | 7.45 | 7.2975 | -2.05
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.29 | 53970.766 | 1724786 | 36311.348 | 0.015 | 0.998 | 11.92 | 12.7499 | 6.96
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.29 | 57101.934 | 1824724 | 38973.023 | 0.015 | 0.998 | 11.92 | 12.5677 | 5.43
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.29 | 93540.961 | 2970972 | 37384.508 | 0.025 | 0.998 | 19.07 | 21.4855 | 12.67
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.29 | 102659.703 | 3279173 | 40601.832 | 0.025 | 0.998 | 19.07 | 21.7118 | 13.85
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.29 | 136254.234 | 4337225 | 37091.914 | 0.037 | 0.998 | 30.52 | 31.5511 | 3.38
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.29 | 142710.609 | 4517513 | 40061.527 | 0.036 | 0.998 | 30.52 | 30.5964 | 0.25
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.29 | 203442.953 | 6492925 | 35457.441 | 0.057 | 0.998 | 48.83 | 49.2709 | 0.9
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.29 | 213439.125 | 6869003 | 37898.871 | 0.056 | 0.998 | 48.83 | 48.3628 | -0.96
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.29 | 307602.375 | 9752293 | 33010.672 | 0.093 | 0.998 | 78.13 | 79.9495 | 2.33
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.29 | 321368.563 | 10134493 | 33821.711 | 0.095 | 0.998 | 78.13 | 81.5203 | 4.34
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.29 | 429381.469 | 13474437 | 31009.475 | 0.138 | 0.998 | 125 | 118.645 | -5.08
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.29 | 456410.469 | 14321911 | 33647.129 | 0.136 | 0.998 | 125 | 116.2373 | -7.01
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.29 | 533271.25 | 16663163 | 29707.164 | 0.18 | 0.998 | 156.25 | 153.6141 | -1.69
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.29 | 570029.625 | 17655558 | 32227.65 | 0.177 | 0.998 | 156.25 | 151.3732 | -3.12
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.29 | 821408.875 | 25130938 | 27123.66 | 0.303 | 0.998 | 250 | 258.1306 | 3.25
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.29 | 867564.75 | 26367316 | 29370.512 | 0.295 | 0.998 | 250 | 251.839 | 0.74
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.31 | 44.99 | 710 |  |  | 0.998
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.34 | 19.874 | 743 |  |  | 0.998
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.33 | 49.65 | 989 |  |  | 0.998
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.31 | 26.628 | 1282 |  |  | 0.998
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.29 | 150.075 | 2767 | 37792.707 | 0 | 0.998
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.29 | 103.346 | 3479 | 38336.848 | 0 | 0.998
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.29 | 133.816 | 4106 | 38771.027 | 0 | 0.998
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.29 | 113.565 | 3097 | 38245.285 | 0 | 0.998
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.29 | 151.703 | 4265 | 41662.426 | 0 | 0.998
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.29 | 108.92 | 2283 | 40001.32 | 0 | 0.998
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.29 | 80.47 | 2644 | 42510.348 | 0 | 0.998
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.29 | 113.05 | 3512 | 42210.781 | 0 | 0.998
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.29 | 75.13 | 1928 | 14514.265 | 0 | 0.998 |  | 0.0014
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.29 | 2693.958 | 89714 | 38590.461 | 0.001 | 0.998 | 0.63 | 0.55797 | -11.43
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.29 | 3139.04 | 97125 | 40316.699 | 0.001 | 0.998 | 0.63 | 0.62729 | -0.43
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.29 | 12772.691 | 406187 | 40285.297 | 0.003 | 0.998 | 2.5 | 2.6868 | 7.47
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.29 | 14056.702 | 463083 | 46858.434 | 0.003 | 0.998 | 2.5 | 2.5398 | 1.59
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.29 | 30260.361 | 977413 | 39404.684 | 0.008 | 0.998 | 6.25 | 6.5682 | 5.09
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.29 | 32222.326 | 1026852 | 43284.461 | 0.007 | 0.998 | 6.25 | 6.3658 | 1.85
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.29 | 113406.469 | 3612227 | 36878.289 | 0.031 | 0.998 | 25 | 26.4108 | 5.64
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.29 | 119831.102 | 3818558 | 40310.613 | 0.03 | 0.998 | 25 | 25.5303 | 2.12
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.25 | 181.391 | 1809 |  |  | 0.998
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.28 | 46.525 | 988 |  |  | 0.998
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.3 | 0.717 | 60 |  |  | 0.998
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.28 | 43.526 | 1093 |  |  | 0.998
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.29 | 205.321 | 5267 |  |  | 0.998
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.32 | 45.597 | 1050 |  |  | 0.998
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.3 | 53.246 | 931 |  |  | 0.998
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.29 | 92.209 | 1708 |  |  | 0.998
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.33 | 40.784 | 1301 |  |  | 0.998
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.29 | 197.565 | 5385 |  |  | 0.998
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.29 | 54.728 | 1442 |  |  | 0.998
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.27 | 20.866 | 493 |  |  | 0.998
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.25 | 50.808 | 736 |  |  | 0.998
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 2.25 | 70.522 | 1242 |  |  | 0.998
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.36 | 73.566 | 965 |  |  | 0.998
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.29 | 4536.727 | 148347 | 42185.547 | 0.001 | 0.998 |  | 0.88289
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.29 | 5222.466 | 171073 | 43559.676 | 0.001 | 0.998 |  | 0.98923
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.29 | 4229.53 | 140333 | 41573.668 | 0.001 | 0.998 |  | 0.83289
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.29 | 178.196 | 4457 | 42123.012 | 0 | 0.998
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.29 | 167.629 | 4287 | 41354.645 | 0 | 0.998
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.29 | 146.175 | 3256 | 39860.156 | 0 | 0.998
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.29 | 159.547 | 3706 | 40662.246 | 0 | 0.998
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.29 | 185.232 | 4299 | 41766.652 | 0 | 0.998
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.29 | 128.842 | 3062 | 38594.965 | 0 | 0.998
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.29 | 448711.094 | 13890057 | 31312.432 | 0.143 | 0.998 |  | 122.7682
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.29 | 482083.438 | 14960000 | 31432.322 | 0.153 | 0.998 |  | 131.3548
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.29 | 500682.344 | 15483513 | 31703.865 | 0.158 | 0.998 |  | 135.2349
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.29 | 146.408 | 4941 | 36568.078 | 0 | 0.998
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.29 | 209.484 | 7209 | 39947.965 | 0 | 0.998 |  | 0.00199
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.29 | 201.569 | 6163 | 39242.902 | 0 | 0.998 |  | 0.00106
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.29 | 156.397 | 3979 | 37110.707 | 0 | 0.998
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.29 | 84.559 | 2439 | 38090.129 | 0 | 0.998
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.29 | 113.601 | 2361 | 39006.359 | 0 | 0.998
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.29 | 540588 | 16791414 | 36889.391 | 0.147 | 0.998 |  | 125.5328
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.29 | 532659.75 | 16598030 | 33918.5 | 0.157 | 0.998 |  | 134.4819
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.29 | 524709.313 | 16397762 | 33629.922 | 0.156 | 0.998 |  | 133.6157
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.29 | 175.618 | 5829 | 39910.887 | 0 | 0.998
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.29 | 190.796 | 6047 | 39164.621 | 0 | 0.998
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.29 | 110.206 | 2674 | 39562.254 | 0 | 0.998
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.29 | 94.143 | 2920 | 39324.82 | 0 | 0.998
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.29 | 87.277 | 2357 | 37500.363 | 0 | 0.998
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.29 | 73.969 | 2556 | 38443.734 | 0 | 0.998
M6PFDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.29 | 37339.707 | 1206869 | 37339.707 |  |  | 0.01 | 0.00963 | -3.68
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.29 | 37417.016 | 1189700 | 37417.016 |  |  | 0.01 | 0.00965 | -3.49
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.29 | 38646.242 | 1237572 | 38646.242 |  |  | 0.01 | 0.00997 | -0.31
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.29 | 39172.805 | 1248249 | 39172.805 |  |  | 0.01 | 0.0101 | 1.04
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.29 | 41901.957 | 1338794 | 41901.957 |  |  | 0.01 | 0.01081 | 8.08
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.29 | 40732.211 | 1304799 | 40732.211 |  |  | 0.01 | 0.01051 | 5.07
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.29 | 41387.645 | 1353108 | 41387.645 |  |  | 0.01 | 0.01068 | 6.76
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.29 | 37559.73 | 1208986 | 37559.73 |  |  | 0.01 | 0.00969 | -3.12
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.29 | 40856.688 | 1297828 | 40856.688 |  |  | 0.01 | 0.01054 | 5.39
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.29 | 42048.52 | 1339339 | 42048.52 |  |  | 0.01 | 0.01085 | 8.46
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.29 | 40422.34 | 1325262 | 40422.34 |  |  | 0.01 | 0.01043 | 4.27
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.29 | 44425.273 | 1445217 | 44425.273 |  |  | 0.01 | 0.01146 | 14.59
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.29 | 43092.957 | 1418734 | 43092.957 |  |  | 0.01 | 0.01112 | 11.16
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.29 | 45184.34 | 1445972 | 45184.34 |  |  | 0.01 | 0.01166 | 16.55
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.29 | 38314.09 | 1223482 | 38314.09 |  |  | 0.01 | 0.00988 | -1.17
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.29 | 39381.18 | 1256250 | 39381.18 |  |  | 0.01 | 0.01016 | 1.58
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.29 | 39681.668 | 1272905 | 39681.668 |  |  | 0.01 | 0.01024 | 2.36
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.29 | 41193.734 | 1318586 | 41193.734 |  |  | 0.01 | 0.01063 | 6.26
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.29 | 41899.441 | 1344911 | 41899.441 |  |  | 0.01 | 0.01081 | 8.08
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.29 | 42587.418 | 1396713 | 42587.418 |  |  | 0.01 | 0.01099 | 9.85
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.29 | 42725.34 | 1365175 | 42725.34 |  |  | 0.01 | 0.01102 | 10.21
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.29 | 38319.367 | 1218543 | 38319.367 |  |  | 0.01 | 0.00988 | -1.16
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.29 | 39899.059 | 1258870 | 39899.059 |  |  | 0.01 | 0.01029 | 2.92
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.29 | 38738.613 | 1223307 | 38738.613 |  |  | 0.01 | 0.00999 | -0.08
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.29 | 41931.879 | 1343595 | 41931.879 |  |  | 0.01 | 0.01082 | 8.16
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.29 | 42107.133 | 1360262 | 42107.133 |  |  | 0.01 | 0.01086 | 8.61
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.29 | 43136.676 | 1396762 | 43136.676 |  |  | 0.01 | 0.01113 | 11.27
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.29 | 41798.488 | 1364262 | 41798.488 |  |  | 0.01 | 0.01078 | 7.82
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.29 | 35923.289 | 1140870 | 35923.289 |  |  | 0.01 | 0.00927 | -7.34
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.29 | 39440.199 | 1263270 | 39440.199 |  |  | 0.01 | 0.01017 | 1.73
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.29 | 37568.254 | 1184707 | 37568.254 |  |  | 0.01 | 0.00969 | -3.09
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.29 | 40251.609 | 1303796 | 40251.609 |  |  | 0.01 | 0.01038 | 3.83
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.29 | 41405.266 | 1332128 | 41405.266 |  |  | 0.01 | 0.01068 | 6.8
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.29 | 39118.82 | 1254442 | 39118.82 |  |  | 0.01 | 0.01009 | 0.9
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.29 | 41191.141 | 1308029 | 41191.141 |  |  | 0.01 | 0.01063 | 6.25
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.29 | 36542.281 | 1160598 | 36542.281 |  |  | 0.01 | 0.00943 | -5.74
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.29 | 39243.359 | 1245149 | 39243.359 |  |  | 0.01 | 0.01012 | 1.23
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.29 | 38877.977 | 1241009 | 38877.977 |  |  | 0.01 | 0.01003 | 0.28
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.29 | 39630.223 | 1251803 | 39630.223 |  |  | 0.01 | 0.01022 | 2.22
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.29 | 41834.133 | 1337079 | 41834.133 |  |  | 0.01 | 0.01079 | 7.91
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.29 | 40277.5 | 1296634 | 40277.5 |  |  | 0.01 | 0.01039 | 3.89
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.29 | 42275.109 | 1357783 | 42275.109 |  |  | 0.01 | 0.0109 | 9.05
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.29 | 37978.066 | 1233818 | 37978.066 |  |  | 0.01 | 0.0098 | -2.04
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.29 | 41927.211 | 1325309 | 41927.211 |  |  | 0.01 | 0.01081 | 8.15
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.29 | 37018.879 | 1176726 | 37018.879 |  |  | 0.01 | 0.00955 | -4.51
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.29 | 40052.781 | 1281326 | 40052.781 |  |  | 0.01 | 0.01033 | 3.31
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.29 | 34951.309 | 1116667 | 34951.309 |  |  | 0.01 | 0.00902 | -9.85
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.29 | 40052.25 | 1280612 | 40052.25 |  |  | 0.01 | 0.01033 | 3.31
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.29 | 36311.348 | 1143914 | 36311.348 |  |  | 0.01 | 0.00937 | -6.34
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.29 | 38973.023 | 1238290 | 38973.023 |  |  | 0.01 | 0.01005 | 0.53
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.29 | 37384.508 | 1185239 | 37384.508 |  |  | 0.01 | 0.00964 | -3.57
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.29 | 40601.832 | 1289312 | 40601.832 |  |  | 0.01 | 0.01047 | 4.73
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.29 | 37091.914 | 1176282 | 37091.914 |  |  | 0.01 | 0.00957 | -4.32
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.29 | 40061.527 | 1256952 | 40061.527 |  |  | 0.01 | 0.01033 | 3.34
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.29 | 35457.441 | 1116592 | 35457.441 |  |  | 0.01 | 0.00915 | -8.54
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.29 | 37898.871 | 1191283 | 37898.871 |  |  | 0.01 | 0.00978 | -2.24
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.29 | 33010.672 | 1030582 | 33010.672 |  |  | 0.01 | 0.00851 | -14.85
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.29 | 33821.711 | 1061806 | 33821.711 |  |  | 0.01 | 0.00872 | -12.76
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.29 | 31009.475 | 962542 | 31009.475 |  |  | 0.01 | 0.008 | -20.01
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.29 | 33647.129 | 1031146 | 33647.129 |  |  | 0.01 | 0.00868 | -13.21
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.29 | 29707.164 | 916690 | 29707.164 |  |  | 0.01 | 0.00766 | -23.37
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.29 | 32227.65 | 987816 | 32227.65 |  |  | 0.01 | 0.00831 | -16.87
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.29 | 27123.66 | 810713 | 27123.66 |  |  | 0.01 | 0.007 | -30.04
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.29 | 29370.512 | 880441 | 29370.512 |  |  | 0.01 | 0.00758 | -24.24
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.29 | 37792.707 | 1217880 | 37792.707 |  |  | 0.01 | 0.00975 | -2.52
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.29 | 38336.848 | 1233551 | 38336.848 |  |  | 0.01 | 0.00989 | -1.11
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.29 | 38771.027 | 1238209 | 38771.027 |  |  | 0.01 | 0.01 | 0.01
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.29 | 38245.285 | 1242430 | 38245.285 |  |  | 0.01 | 0.00987 | -1.35
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.29 | 41662.426 | 1352565 | 41662.426 |  |  | 0.01 | 0.01075 | 7.47
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.29 | 40001.32 | 1314374 | 40001.32 |  |  | 0.01 | 0.01032 | 3.18
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.29 | 42510.348 | 1391581 | 42510.348 |  |  | 0.01 | 0.01097 | 9.65
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.29 | 42210.781 | 1343437 | 42210.781 |  |  | 0.01 | 0.01089 | 8.88
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.29 | 14514.265 | 480943 | 14514.265 |  |  | 0.01 | 0.00374 | -62.56
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.29 | 38590.461 | 1230275 | 38590.461 |  |  | 0.01 | 0.00995 | -0.46
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.29 | 40316.699 | 1282533 | 40316.699 |  |  | 0.01 | 0.0104 | 3.99
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.29 | 40285.297 | 1283915 | 40285.297 |  |  | 0.01 | 0.01039 | 3.91
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.29 | 46858.434 | 1508552 | 46858.434 |  |  | 0.01 | 0.01209 | 20.87
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.29 | 39404.684 | 1244131 | 39404.684 |  |  | 0.01 | 0.01016 | 1.64
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.29 | 43284.461 | 1398391 | 43284.461 |  |  | 0.01 | 0.01116 | 11.65
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.29 | 36878.289 | 1158022 | 36878.289 |  |  | 0.01 | 0.00951 | -4.87
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.29 | 40310.613 | 1274189 | 40310.613 |  |  | 0.01 | 0.0104 | 3.98
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.29 | 42185.547 | 1327132 | 42185.547 |  |  | 0.01 | 0.01088 | 8.82
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.29 | 43559.676 | 1400670 | 43559.676 |  |  | 0.01 | 0.01124 | 12.36
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.29 | 41573.668 | 1324033 | 41573.668 |  |  | 0.01 | 0.01072 | 7.24
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.29 | 42123.012 | 1375573 | 42123.012 |  |  | 0.01 | 0.01087 | 8.65
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.29 | 41354.645 | 1322921 | 41354.645 |  |  | 0.01 | 0.01067 | 6.67
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.29 | 39860.156 | 1284787 | 39860.156 |  |  | 0.01 | 0.01028 | 2.82
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.29 | 40662.246 | 1307754 | 40662.246 |  |  | 0.01 | 0.01049 | 4.89
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.29 | 41766.652 | 1350220 | 41766.652 |  |  | 0.01 | 0.01077 | 7.73
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.29 | 38594.965 | 1249997 | 38594.965 |  |  | 0.01 | 0.00996 | -0.45
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.29 | 31312.432 | 979767 | 31312.432 |  |  | 0.01 | 0.00808 | -19.23
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.29 | 31432.322 | 973478 | 31432.322 |  |  | 0.01 | 0.00811 | -18.92
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.29 | 31703.865 | 984198 | 31703.865 |  |  | 0.01 | 0.00818 | -18.22
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.29 | 36568.078 | 1199666 | 36568.078 |  |  | 0.01 | 0.00943 | -5.67
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.29 | 39947.965 | 1290917 | 39947.965 |  |  | 0.01 | 0.0103 | 3.04
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.29 | 39242.902 | 1266126 | 39242.902 |  |  | 0.01 | 0.01012 | 1.22
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.29 | 37110.707 | 1210382 | 37110.707 |  |  | 0.01 | 0.00957 | -4.28
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.29 | 38090.129 | 1220383 | 38090.129 |  |  | 0.01 | 0.00983 | -1.75
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.29 | 39006.359 | 1255961 | 39006.359 |  |  | 0.01 | 0.01006 | 0.61
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.29 | 36889.391 | 1128492 | 36889.391 |  |  | 0.01 | 0.00952 | -4.85
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.29 | 33918.5 | 1031045 | 33918.5 |  |  | 0.01 | 0.00875 | -12.51
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.29 | 33629.922 | 1016596 | 33629.922 |  |  | 0.01 | 0.00867 | -13.25
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.29 | 39910.887 | 1278369 | 39910.887 |  |  | 0.01 | 0.01029 | 2.95
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.29 | 39164.621 | 1272815 | 39164.621 |  |  | 0.01 | 0.0101 | 1.02
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.29 | 39562.254 | 1298513 | 39562.254 |  |  | 0.01 | 0.0102 | 2.05
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.29 | 39324.82 | 1259723 | 39324.82 |  |  | 0.01 | 0.01014 | 1.44
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.29 | 37500.363 | 1223024 | 37500.363 |  |  | 0.01 | 0.00967 | -3.27
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.29 | 38443.734 | 1233450 | 38443.734 |  |  | 0.01 | 0.00992 | -0.84

## DXTSID80108992
*Analyte in Mix 3 | [8 empty] | PFOS
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.39 | 540.821 | 10969 | 3515.653 | 0.002 | 0.998 | 0.17 | 0.12168 | -28.42
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.38 | 527.062 | 11033 | 3625.022 | 0.001 | 0.998 | 0.17 | 0.09814 | -42.27
UC_AF_Mix3_A | 1.5635 | 50.0317 | 0.0045917 | 0.0046763 | 0.0002145 | 4.587% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.38 | 567.657 | 13006 | 3855.243 | 0.001 | 0.998 | 0.17 | 0.1033 | -39.24
UC_AF_Mix3_B | 1.5723 | 50.3146 | 0.0045171 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.39 | 586.011 | 12529 | 3890.618 | 0.002 | 0.998 | 0.17 | 0.11273 | -33.69
UC_AF_Mix3_C | 1.6938 | 54.201 | 0.0049202 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.39 | 627.342 | 14206 | 4065.303 | 0.002 | 0.998 | 0.17 | 0.12304 | -27.63
UC_T1hr_Mix3_A | 134.8422 | 10787.3776 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.38 | 605.618 | 12945 | 3958.613 | 0.002 | 0.998 | 0.17 | 0.11933 | -29.81
UC_T1hr_Mix3_B | 141.3775 | 11310.2 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.38 | 588.431 | 13426 | 3983.047 | 0.001 | 0.998 | 0.17 | 0.10467 | -38.43
UC_T1hr_Mix3_C | 139.7575 | 11180.5984 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.39 | 684.029 | 14818 | 3777.996 | 0.002 | 0.998 | 0.28 | 0.19764 | -29.41
UC_T5hr_Mix3_A | 136.2024 | 10896.1912 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.39 | 827.368 | 19812 | 3695.023 | 0.002 | 0.998 | 0.28 | 0.31721 | 13.29
UC_T5hr_Mix3_B | 139.2339 | 11138.7096 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.38 | 839.677 | 20095 | 4039.755 | 0.002 | 0.998 | 0.28 | 0.27241 | -2.71
UC_T5hr_Mix3_C | 137.699 | 11015.9184 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.39 | 827.47 | 18625 | 4132.465 | 0.002 | 0.998 | 0.28 | 0.25116 | -10.3
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.39 | 954.839 | 21923 | 4359.24 | 0.002 | 0.998 | 0.28 | 0.30361 | 8.43
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.38 | 864.866 | 20423 | 4343.202 | 0.002 | 0.998 | 0.28 | 0.24807 | -11.4
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.39 | 948.439 | 22130 | 4469.724 | 0.002 | 0.998 | 0.28 | 0.28451 | 1.61
20201123_PFAS_PPB_UC_Sample010 | CC4 | 0.71 | 0.66358 | 0.70308 | 0.023203 | 3.3% | 0.072928 | -0.9744% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1153.145 | 30019 | 3699.702 | 0.003 | 0.998 | 0.44 | 0.56207 | 27.74
20201123_PFAS_PPB_UC_Sample035 | CC4 | 0.71 | 0.72889 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1271.47 | 31590 | 3889.072 | 0.003 | 0.998 | 0.44 | 0.6046 | 37.41
20201123_PFAS_PPB_UC_Sample042 | CC4 | 0.71 | 0.68288 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1328.937 | 33861 | 3796.014 | 0.004 | 0.998 | 0.44 | 0.66918 | 52.09
20201123_PFAS_PPB_UC_Sample082 | CC4 | 0.71 | 0.71415 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1365.324 | 33701 | 4043.019 | 0.003 | 0.998 | 0.44 | 0.63463 | 44.23
20201123_PFAS_PPB_UC_Sample108 | CC4 | 0.71 | 0.70111 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1292.338 | 29650 | 4362.105 | 0.003 | 0.998 | 0.44 | 0.51905 | 17.97
20201123_PFAS_PPB_UC_Sample115 | CC4 | 0.71 | 0.72448 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.38 | 1350.651 | 33993 | 4184.41 | 0.003 | 0.998 | 0.44 | 0.59302 | 34.78
20201123_PFAS_PPB_UC_Sample122 | CC4 | 0.71 | 0.70648 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.39 | 1399.109 | 35520 | 4314.678 | 0.003 | 0.998 | 0.44 | 0.59716 | 35.72
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1326.21 | 34770 | 3810.099 | 0.003 | 0.998 | 0.71 | 0.66358 | -6.54
Blanks should be less than… | 0.036464 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1399.694 | 34316 | 3767.749 | 0.004 | 0.998 | 0.71 | 0.72889 | 2.66
eLOQ | 0.71 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1363.15 | 33780 | 3839.881 | 0.004 | 0.998 | 0.71 | 0.68288 | -3.82
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.38 | 1454.075 | 35802 | 3970.605 | 0.004 | 0.998 | 0.71 | 0.71415 | 0.58
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1482.216 | 36689 | 4099.792 | 0.004 | 0.998 | 0.71 | 0.70111 | -1.25
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1560.742 | 39680 | 4219.208 | 0.004 | 0.998 | 0.71 | 0.72448 | 2.04
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.39 | 1508.809 | 36809 | 4151.266 | 0.004 | 0.998 | 0.71 | 0.70648 | -0.5
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.39 | 1883.786 | 47481 | 3518.784 | 0.005 | 0.998 | 1.14 | 1.1859 | 4.02
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2004.768 | 50060 | 3848.652 | 0.005 | 0.998 | 1.14 | 1.1456 | 0.49
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.38 | 1926.949 | 49041 | 3659.798 | 0.005 | 0.998 | 1.14 | 1.1612 | 1.86
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.38 | 2100.144 | 53892 | 3911.216 | 0.005 | 0.998 | 1.14 | 1.1903 | 4.41
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2028.367 | 51723 | 4000.38 | 0.005 | 0.998 | 1.14 | 1.1069 | -2.9
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2153.438 | 55100 | 4015.649 | 0.005 | 0.998 | 1.14 | 1.1884 | 4.25
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.39 | 2146.172 | 53294 | 4116.965 | 0.005 | 0.998 | 1.14 | 1.1467 | 0.59
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.39 | 2778.222 | 73282 | 3525.893 | 0.008 | 0.998 | 1.82 | 1.8901 | 3.85
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.38 | 2839.042 | 73645 | 3778.806 | 0.008 | 0.998 | 1.82 | 1.7879 | -1.76
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.38 | 2903.629 | 76605 | 3857.974 | 0.008 | 0.998 | 1.82 | 1.7916 | -1.56
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.39 | 2960.601 | 78339 | 3887.929 | 0.008 | 0.998 | 1.82 | 1.8163 | -0.2
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3126.272 | 82932 | 4255.308 | 0.007 | 0.998 | 1.82 | 1.7416 | -4.31
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3208.746 | 87627 | 3982.236 | 0.008 | 0.998 | 1.82 | 1.9397 | 6.58
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3324.506 | 86797 | 4224.377 | 0.008 | 0.998 | 1.82 | 1.8874 | 3.7
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.39 | 4068.736 | 109708 | 3792.97 | 0.011 | 0.998 | 2.91 | 2.6836 | -7.78
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.38 | 4287.631 | 110532 | 4020.333 | 0.011 | 0.998 | 2.91 | 2.6663 | -8.38
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.38 | 6271.539 | 167990 | 3544.214 | 0.018 | 0.998 | 4.66 | 4.6238 | -0.78
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.39 | 6823.404 | 186699 | 3883.1 | 0.018 | 0.998 | 4.66 | 4.5896 | -1.51
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.39 | 9385.291 | 253109 | 3472.039 | 0.027 | 0.998 | 7.45 | 7.22 | -3.09
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.38 | 10241.668 | 273649 | 3891.241 | 0.026 | 0.998 | 7.45 | 7.0223 | -5.74
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.39 | 16472.125 | 450029 | 3420.079 | 0.048 | 0.998 | 11.92 | 13.0821 | 9.75
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.38 | 17748.586 | 475577 | 3876.859 | 0.046 | 0.998 | 11.92 | 12.4223 | 4.21
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.39 | 28893.621 | 771561 | 3621.232 | 0.08 | 0.998 | 19.07 | 21.8188 | 14.41
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.39 | 30783.832 | 821268 | 3949.588 | 0.078 | 0.998 | 19.07 | 21.3096 | 11.74
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.39 | 40739.523 | 1081526 | 3567.718 | 0.114 | 0.998 | 30.52 | 31.2723 | 2.46
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.39 | 44354.648 | 1184102 | 3972.074 | 0.112 | 0.998 | 30.52 | 30.5806 | 0.2
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.38 | 60867.402 | 1597439 | 3390.249 | 0.18 | 0.998 | 48.83 | 49.092 | 0.54
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.39 | 66189.313 | 1736158 | 3703.594 | 0.179 | 0.998 | 48.83 | 48.8694 | 0.08
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.38 | 93646.594 | 2432739 | 3215.925 | 0.291 | 0.998 | 78.13 | 79.1339 | 1.28
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.39 | 97748.891 | 2529847 | 3471.925 | 0.282 | 0.998 | 78.13 | 76.5559 | -2.01
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.39 | 127253.023 | 3214305 | 2816.201 | 0.452 | 0.998 | 125 | 121.5041 | -2.8
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.39 | 139309.797 | 3475948 | 3256.591 | 0.428 | 0.998 | 125 | 115.2151 | -7.83
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.39 | 161479.188 | 4011924 | 2776.837 | 0.582 | 0.998 | 156.25 | 155.0028 | -0.8
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.39 | 176249.031 | 4387810 | 2948.698 | 0.598 | 0.998 | 156.25 | 159.1454 | 1.85
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.39 | 249333.906 | 5930930 | 2500.137 | 0.997 | 0.998 | 250 | 258.609 | 3.44
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.38 | 268494.438 | 6311834 | 2836.917 | 0.946 | 0.998 | 250 | 246.2291 | -1.51
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [5 empty] | 0.998
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [5 empty] | 0.998
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.38 | 9.846 | 305 |  |  | 0.998
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.4 | 0.946 | 45 |  |  | 0.998
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.39 | 318.799 | 5621 | 3612.967 | 0.001 | 0.998
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.38 | 263.402 | 4646 | 3638.662 | 0.001 | 0.998
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.38 | 278.284 | 4387 | 3744.692 | 0.001 | 0.998
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.39 | 316.148 | 5605 | 3766.95 | 0.001 | 0.998
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.39 | 357.354 | 6692 | 4093.351 | 0.001 | 0.998
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.39 | 302.362 | 6403 | 4037.25 | 0.001 | 0.998
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.39 | 342 | 6891 | 4145.597 | 0.001 | 0.998
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.38 | 319.602 | 5525 | 4047.1 | 0.001 | 0.998
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.39 | 103.959 | 2131 | 1428.834 | 0.001 | 0.998
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.39 | 1191.374 | 29767 | 3600.895 | 0.003 | 0.998 | 0.63 | 0.61554 | -2.3
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.38 | 1205.266 | 28393 | 3799.566 | 0.003 | 0.998 | 0.63 | 0.57748 | -8.34
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.39 | 4194.164 | 112327 | 3990.492 | 0.011 | 0.998 | 2.5 | 2.6232 | 4.93
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.38 | 4807.84 | 127955 | 4658.844 | 0.01 | 0.998 | 2.5 | 2.5701 | 2.81
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.38 | 9285.781 | 249606 | 3898.205 | 0.024 | 0.998 | 6.25 | 6.3277 | 1.24
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.38 | 9814.89 | 268958 | 4262.813 | 0.023 | 0.998 | 6.25 | 6.1063 | -2.3
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.38 | 33847.813 | 890702 | 3650.252 | 0.093 | 0.998 | 25 | 25.3803 | 1.52
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.38 | 36575.203 | 979482 | 3777.862 | 0.097 | 0.998 | 25 | 26.5039 | 6.02
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.4 | 1.81 | 69 |  |  | 0.998
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.38 | 1.317 | 60 |  |  | 0.998
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.38 | 1.662 | 81 |  |  | 0.998
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.38 | 1.939 | 95 |  |  | 0.998
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.38 | 19.47 | 536 |  |  | 0.998
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.39 | 3.406 | 132 |  |  | 0.998
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.38 | 0.68 | 36 |  |  | 0.998
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.38 | 0.677 | 36 |  |  | 0.998
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.3 | 1.455 | 61 |  |  | 0.998
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.39 | 35.317 | 1175 |  |  | 0.998
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.38 | 3.556 | 134 |  |  | 0.998
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.39 | 2.669 | 119 |  |  | 0.998
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.4 | 1.987 | 98 |  |  | 0.998
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.998
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [5 empty] | 0.998
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.38 | 295.277 | 4998 | 3968.314 | 0.001 | 0.998
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.39 | 297.621 | 5067 | 4331.905 | 0.001 | 0.998
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.39 | 316.359 | 6489 | 3953.513 | 0.001 | 0.998
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.39 | 337.693 | 6453 | 4149.038 | 0.001 | 0.998
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.39 | 340.582 | 5610 | 4039.311 | 0.001 | 0.998
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.38 | 330.39 | 6190 | 3948.045 | 0.001 | 0.998
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.39 | 2757.122 | 81393 | 4110.244 | 0.007 | 0.998 |  | 1.5635
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.38 | 2786.65 | 83571 | 4134.708 | 0.007 | 0.998 |  | 1.5723
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.38 | 2731.141 | 80144 | 3806.312 | 0.007 | 0.998 |  | 1.6938
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.39 | 123.017 | 2036 | 3714.114 | 0 | 0.998
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.31 | 91.615 | 1338 | 3808.625 | 0 | 0.998
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.38 | 135.134 | 2066 | 3752.886 | 0 | 0.998
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.39 | 148.575 | 2751 | 3651.146 | 0 | 0.998
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.39 | 168.705 | 2828 | 4014.833 | 0 | 0.998
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.39 | 161.723 | 2740 | 3989.71 | 0 | 0.998
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.39 | 146618.016 | 3627747 | 2913.662 | 0.503 | 0.998 |  | 134.8422
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.39 | 150890.422 | 3728096 | 2855.056 | 0.529 | 0.998 |  | 141.3775
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.38 | 155619.047 | 3811427 | 2979.928 | 0.522 | 0.998 |  | 139.7575
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.38 | 125.056 | 2465 | 4423.004 | 0 | 0.998
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.38 | 138.453 | 2327 | 3863.019 | 0 | 0.998
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.38 | 170 | 3330 | 3884.237 | 0 | 0.998
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.38 | 153.994 | 2656 | 3750.135 | 0 | 0.998
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.39 | 195.995 | 3559 | 3862.314 | 0.001 | 0.998
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.39 | 304.253 | 6793 | 3867.536 | 0.001 | 0.998
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.38 | 155005.031 | 3792690 | 3048.485 | 0.508 | 0.998 |  | 136.2024
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.38 | 153813.781 | 3738978 | 2956.842 | 0.52 | 0.998 |  | 139.2339
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.39 | 154819.5 | 3822974 | 3010.562 | 0.514 | 0.998 |  | 137.699
M8PFOS
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3515.653 | 116548 | 3515.653 |  |  | 0.01 | 0.00928 | -7.17
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3625.022 | 121822 | 3625.022 |  |  | 0.01 | 0.00957 | -4.28
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3855.243 | 127878 | 3855.243 |  |  | 0.01 | 0.01018 | 1.8
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.38 | 3890.618 | 131933 | 3890.618 |  |  | 0.01 | 0.01027 | 2.73
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.39 | 4065.303 | 138492 | 4065.303 |  |  | 0.01 | 0.01073 | 7.35
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3958.613 | 130634 | 3958.613 |  |  | 0.01 | 0.01045 | 4.53
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.39 | 3983.047 | 131245 | 3983.047 |  |  | 0.01 | 0.01052 | 5.18
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.39 | 3777.996 | 123463 | 3777.996 |  |  | 0.01 | 0.00998 | -0.24
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.38 | 3695.023 | 122492 | 3695.023 |  |  | 0.01 | 0.00976 | -2.43
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.38 | 4039.755 | 137260 | 4039.755 |  |  | 0.01 | 0.01067 | 6.67
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.38 | 4132.465 | 139020 | 4132.465 |  |  | 0.01 | 0.01091 | 9.12
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4359.24 | 146540 | 4359.24 |  |  | 0.01 | 0.01151 | 15.11
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4343.202 | 143908 | 4343.202 |  |  | 0.01 | 0.01147 | 14.69
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.39 | 4469.724 | 147677 | 4469.724 |  |  | 0.01 | 0.0118 | 18.03
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3699.702 | 122764 | 3699.702 |  |  | 0.01 | 0.00977 | -2.31
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3889.072 | 130869 | 3889.072 |  |  | 0.01 | 0.01027 | 2.69
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.38 | 3796.014 | 125474 | 3796.014 |  |  | 0.01 | 0.01002 | 0.24
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.38 | 4043.019 | 135627 | 4043.019 |  |  | 0.01 | 0.01068 | 6.76
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.38 | 4362.105 | 140819 | 4362.105 |  |  | 0.01 | 0.01152 | 15.18
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.39 | 4184.41 | 141733 | 4184.41 |  |  | 0.01 | 0.01105 | 10.49
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.39 | 4314.678 | 143586 | 4314.678 |  |  | 0.01 | 0.01139 | 13.93
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3810.099 | 126252 | 3810.099 |  |  | 0.01 | 0.01006 | 0.61
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3767.749 | 123504 | 3767.749 |  |  | 0.01 | 0.00995 | -0.51
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3839.881 | 128310 | 3839.881 |  |  | 0.01 | 0.01014 | 1.4
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.38 | 3970.605 | 133154 | 3970.605 |  |  | 0.01 | 0.01048 | 4.85
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4099.792 | 136853 | 4099.792 |  |  | 0.01 | 0.01083 | 8.26
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4219.208 | 141065 | 4219.208 |  |  | 0.01 | 0.01114 | 11.41
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.39 | 4151.266 | 138285 | 4151.266 |  |  | 0.01 | 0.01096 | 9.62
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3518.784 | 116780 | 3518.784 |  |  | 0.01 | 0.00929 | -7.08
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3848.652 | 125059 | 3848.652 |  |  | 0.01 | 0.01016 | 1.63
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3659.798 | 123427 | 3659.798 |  |  | 0.01 | 0.00966 | -3.36
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.38 | 3911.216 | 128306 | 3911.216 |  |  | 0.01 | 0.01033 | 3.28
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.38 | 4000.38 | 132484 | 4000.38 |  |  | 0.01 | 0.01056 | 5.63
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.39 | 4015.649 | 133746 | 4015.649 |  |  | 0.01 | 0.0106 | 6.04
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.39 | 4116.965 | 134986 | 4116.965 |  |  | 0.01 | 0.01087 | 8.71
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3525.893 | 117373 | 3525.893 |  |  | 0.01 | 0.00931 | -6.9
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3778.806 | 126983 | 3778.806 |  |  | 0.01 | 0.00998 | -0.22
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3857.974 | 129229 | 3857.974 |  |  | 0.01 | 0.01019 | 1.87
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.38 | 3887.929 | 128011 | 3887.929 |  |  | 0.01 | 0.01027 | 2.66
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.38 | 4255.308 | 140860 | 4255.308 |  |  | 0.01 | 0.01124 | 12.36
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.39 | 3982.236 | 132684 | 3982.236 |  |  | 0.01 | 0.01052 | 5.15
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.39 | 4224.377 | 140772 | 4224.377 |  |  | 0.01 | 0.01115 | 11.55
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.38 | 3792.97 | 124818 | 3792.97 |  |  | 0.01 | 0.01002 | 0.16
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.38 | 4020.333 | 132056 | 4020.333 |  |  | 0.01 | 0.01062 | 6.16
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.38 | 3544.214 | 120506 | 3544.214 |  |  | 0.01 | 0.00936 | -6.41
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.38 | 3883.1 | 128767 | 3883.1 |  |  | 0.01 | 0.01025 | 2.54
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.39 | 3472.039 | 115185 | 3472.039 |  |  | 0.01 | 0.00917 | -8.32
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.38 | 3891.241 | 125654 | 3891.241 |  |  | 0.01 | 0.01028 | 2.75
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.38 | 3420.079 | 112099 | 3420.079 |  |  | 0.01 | 0.00903 | -9.69
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.38 | 3876.859 | 129562 | 3876.859 |  |  | 0.01 | 0.01024 | 2.37
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.38 | 3621.232 | 118412 | 3621.232 |  |  | 0.01 | 0.00956 | -4.38
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.38 | 3949.588 | 127568 | 3949.588 |  |  | 0.01 | 0.01043 | 4.29
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.38 | 3567.718 | 114506 | 3567.718 |  |  | 0.01 | 0.00942 | -5.79
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.38 | 3972.074 | 130228 | 3972.074 |  |  | 0.01 | 0.01049 | 4.89
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.39 | 3390.249 | 108665 | 3390.249 |  |  | 0.01 | 0.00895 | -10.48
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.38 | 3703.594 | 118502 | 3703.594 |  |  | 0.01 | 0.00978 | -2.2
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.39 | 3215.925 | 98917 | 3215.925 |  |  | 0.01 | 0.00849 | -15.08
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.39 | 3471.925 | 109171 | 3471.925 |  |  | 0.01 | 0.00917 | -8.32
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.38 | 2816.201 | 87491 | 2816.201 |  |  | 0.01 | 0.00744 | -25.64
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.38 | 3256.591 | 101282 | 3256.591 |  |  | 0.01 | 0.0086 | -14.01
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.38 | 2776.837 | 86453 | 2776.837 |  |  | 0.01 | 0.00733 | -26.68
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.38 | 2948.698 | 91680 | 2948.698 |  |  | 0.01 | 0.00779 | -22.14
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.38 | 2500.137 | 75944 | 2500.137 |  |  | 0.01 | 0.0066 | -33.98
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.38 | 2836.917 | 86662 | 2836.917 |  |  | 0.01 | 0.00749 | -25.09
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.38 | 3612.967 | 119716 | 3612.967 |  |  | 0.01 | 0.00954 | -4.6
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.38 | 3638.662 | 120770 | 3638.662 |  |  | 0.01 | 0.00961 | -3.92
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.38 | 3744.692 | 124131 | 3744.692 |  |  | 0.01 | 0.00989 | -1.12
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.38 | 3766.95 | 124606 | 3766.95 |  |  | 0.01 | 0.00995 | -0.53
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.39 | 4093.351 | 136601 | 4093.351 |  |  | 0.01 | 0.01081 | 8.09
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.39 | 4037.25 | 135475 | 4037.25 |  |  | 0.01 | 0.01066 | 6.61
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.38 | 4145.597 | 135397 | 4145.597 |  |  | 0.01 | 0.01095 | 9.47
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.38 | 4047.1 | 136288 | 4047.1 |  |  | 0.01 | 0.01069 | 6.87
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.39 | 1428.834 | 48845 | 1428.834 |  |  | 0.01 | 0.00377 | -62.27
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.39 | 3600.895 | 118613 | 3600.895 |  |  | 0.01 | 0.00951 | -4.92
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.38 | 3799.566 | 128031 | 3799.566 |  |  | 0.01 | 0.01003 | 0.33
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.38 | 3990.492 | 131836 | 3990.492 |  |  | 0.01 | 0.01054 | 5.37
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.38 | 4658.844 | 155544 | 4658.844 |  |  | 0.01 | 0.0123 | 23.02
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.38 | 3898.205 | 129684 | 3898.205 |  |  | 0.01 | 0.01029 | 2.94
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.39 | 4262.813 | 140994 | 4262.813 |  |  | 0.01 | 0.01126 | 12.56
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.38 | 3650.252 | 120728 | 3650.252 |  |  | 0.01 | 0.00964 | -3.61
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.39 | 3777.862 | 121730 | 3777.862 |  |  | 0.01 | 0.00998 | -0.24
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.38 | 3968.314 | 132324 | 3968.314 |  |  | 0.01 | 0.01048 | 4.79
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.38 | 4331.905 | 143142 | 4331.905 |  |  | 0.01 | 0.01144 | 14.39
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.38 | 3953.513 | 132694 | 3953.513 |  |  | 0.01 | 0.01044 | 4.4
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.38 | 4149.038 | 136420 | 4149.038 |  |  | 0.01 | 0.01096 | 9.56
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.38 | 4039.311 | 135371 | 4039.311 |  |  | 0.01 | 0.01067 | 6.66
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.38 | 3948.045 | 134012 | 3948.045 |  |  | 0.01 | 0.01043 | 4.25
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.38 | 4110.244 | 134922 | 4110.244 |  |  | 0.01 | 0.01085 | 8.53
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.38 | 4134.708 | 137008 | 4134.708 |  |  | 0.01 | 0.01092 | 9.18
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.38 | 3806.312 | 126396 | 3806.312 |  |  | 0.01 | 0.01005 | 0.51
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.39 | 3714.114 | 124654 | 3714.114 |  |  | 0.01 | 0.00981 | -1.93
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.38 | 3808.625 | 124230 | 3808.625 |  |  | 0.01 | 0.01006 | 0.57
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.38 | 3752.886 | 126474 | 3752.886 |  |  | 0.01 | 0.00991 | -0.9
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.39 | 3651.146 | 121856 | 3651.146 |  |  | 0.01 | 0.00964 | -3.59
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.38 | 4014.833 | 133225 | 4014.833 |  |  | 0.01 | 0.0106 | 6.01
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.38 | 3989.71 | 134886 | 3989.71 |  |  | 0.01 | 0.01054 | 5.35
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.38 | 2913.662 | 91176 | 2913.662 |  |  | 0.01 | 0.00769 | -23.06
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.38 | 2855.056 | 90020 | 2855.056 |  |  | 0.01 | 0.00754 | -24.61
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.38 | 2979.928 | 97678 | 2979.928 |  |  | 0.01 | 0.00787 | -21.31
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.38 | 4423.004 | 146329 | 4423.004 |  |  | 0.01 | 0.01168 | 16.79
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.38 | 3863.019 | 128981 | 3863.019 |  |  | 0.01 | 0.0102 | 2.01
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.38 | 3884.237 | 128731 | 3884.237 |  |  | 0.01 | 0.01026 | 2.57
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.38 | 3750.135 | 123730 | 3750.135 |  |  | 0.01 | 0.0099 | -0.97
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.38 | 3862.314 | 129248 | 3862.314 |  |  | 0.01 | 0.0102 | 1.99
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.38 | 3867.536 | 128089 | 3867.536 |  |  | 0.01 | 0.01021 | 2.13
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.38 | 3048.485 | 98100 | 3048.485 |  |  | 0.01 | 0.00805 | -19.5
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.38 | 2956.842 | 92421 | 2956.842 |  |  | 0.01 | 0.00781 | -21.92
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.38 | 3010.562 | 94271 | 3010.562 |  |  | 0.01 | 0.00795 | -20.5

## DTXSID50375114
*Analyte in Mix 2 | [8 empty] | PFPE-6
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.65 | 921.642 | 27945 | 32447.529 | 0 | 0.997 | 0.17 | 0.10827 | -36.31
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1010.817 | 29652 | 32869.164 | 0 | 0.997 | 0.17 | 0.12227 | -28.08
UC_AF_Mix2_A | 0.55238 | 17.6762 | 0.0016817 | 0.0017667 | 0.00015993 | 9.053% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1031.095 | 29137 | 34714.773 | 0 | 0.997 | 0.17 | 0.116 | -31.76
UC_AF_Mix2_B | 0.55467 | 17.7494 | 0.0016671 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1015.728 | 29053 | 33756.629 | 0 | 0.997 | 0.17 | 0.11832 | -30.4
UC_AF_Mix2_C | 0.65526 | 20.9683 | 0.0019511 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1045.793 | 32103 | 36119.102 | 0 | 0.997 | 0.17 | 0.11154 | -34.39
UC_T1hr_Mix2_A | 135.1066 | 10808.5312 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.65 | 983.628 | 29409 | 35462.613 | 0 | 0.997 | 0.17 | 0.10429 | -38.65
UC_T1hr_Mix2_B | 128.4911 | 10279.284 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.65 | 1105.723 | 32191 | 35970.531 | 0 | 0.997 | 0.17 | 0.12219 | -28.12
UC_T1hr_Mix2_C | 136.5417 | 10923.3344 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1892.717 | 55740 | 32267.967 | 0.001 | 0.997 | 0.28 | 0.28866 | 3.09
UC_T5hr_Mix2_A | 131.3825 | 10510.596 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1967.408 | 59010 | 33731.746 | 0.001 | 0.997 | 0.28 | 0.28669 | 2.39
UC_T5hr_Mix2_B | 133.0863 | 10646.9056 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2055.06 | 59931 | 36595.254 | 0.001 | 0.997 | 0.28 | 0.27375 | -2.23
UC_T5hr_Mix2_C | 134.3334 | 10746.6696 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.65 | 1956.784 | 56560 | 36770.707 | 0.001 | 0.997 | 0.28 | 0.25622 | -8.49
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2080.034 | 60858 | 37791.871 | 0.001 | 0.997 | 0.28 | 0.26709 | -4.61
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2147.191 | 62832 | 37730.301 | 0.001 | 0.997 | 0.28 | 0.27824 | -0.63
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.65 | 2290.525 | 67744 | 37755.715 | 0.001 | 0.997 | 0.28 | 0.30065 | 7.38
20201123_PFAS_PPB_UC_Sample008 | CC2 | 0.28 | 0.28866 | 0.27876 | 0.014774 | 5.3% | 0.046436 | -0.4439% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3094.552 | 90538 | 33209.082 | 0.001 | 0.997 | 0.44 | 0.49462 | 12.41
20201123_PFAS_PPB_UC_Sample033 | CC2 | 0.28 | 0.28669 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3335.631 | 98158 | 35506.813 | 0.001 | 0.997 | 0.44 | 0.49915 | 13.44
20201123_PFAS_PPB_UC_Sample040 | CC2 | 0.28 | 0.27375 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3398.498 | 97760 | 34187.922 | 0.001 | 0.997 | 0.44 | 0.53175 | 20.85
20201123_PFAS_PPB_UC_Sample080 | CC2 | 0.28 | 0.25622 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3448.573 | 102044 | 36297.711 | 0.001 | 0.997 | 0.44 | 0.5055 | 14.89
20201123_PFAS_PPB_UC_Sample106 | CC2 | 0.28 | 0.26709 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3597.861 | 106337 | 37101.148 | 0.001 | 0.997 | 0.44 | 0.51724 | 17.55
20201123_PFAS_PPB_UC_Sample113 | CC2 | 0.28 | 0.27824 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3533.678 | 107018 | 37717.949 | 0.001 | 0.997 | 0.44 | 0.49762 | 13.1
20201123_PFAS_PPB_UC_Sample120 | CC2 | 0.28 | 0.30065 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.65 | 3538.376 | 103766 | 36696.27 | 0.001 | 0.997 | 0.44 | 0.51395 | 16.81
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4415.33 | 131556 | 33202.188 | 0.001 | 0.997 | 0.71 | 0.73212 | 3.12
Blanks should be less than… | 0.023218 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4474.458 | 129596 | 35507.883 | 0.001 | 0.997 | 0.71 | 0.69052 | -2.74
eLOQ | 0.28 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4496.663 | 133073 | 34225.648 | 0.001 | 0.997 | 0.71 | 0.72257 | 1.77
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4881.735 | 145085 | 36058.262 | 0.001 | 0.997 | 0.71 | 0.74645 | 5.13
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4702.09 | 135939 | 36060.398 | 0.001 | 0.997 | 0.71 | 0.71667 | 0.94
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4677.835 | 139603 | 36561.789 | 0.001 | 0.997 | 0.71 | 0.70204 | -1.12
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.65 | 4899.534 | 144733 | 36844.078 | 0.001 | 0.997 | 0.71 | 0.7321 | 3.11
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.65 | 6632.528 | 197162 | 32022.859 | 0.002 | 0.997 | 1.14 | 1.1748 | 3.05
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.65 | 6970.117 | 207434 | 33292.348 | 0.002 | 0.997 | 1.14 | 1.1882 | 4.22
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7280.397 | 213942 | 33805.305 | 0.002 | 0.997 | 1.14 | 1.224 | 7.37
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7141.836 | 210311 | 33890.82 | 0.002 | 0.997 | 1.14 | 1.1963 | 4.94
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7578.791 | 222935 | 35058.344 | 0.002 | 0.997 | 1.14 | 1.2289 | 7.8
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7474.03 | 220792 | 34807.703 | 0.002 | 0.997 | 1.14 | 1.2202 | 7.03
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.65 | 7780.425 | 229346 | 35534.512 | 0.002 | 0.997 | 1.14 | 1.2455 | 9.25
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.65 | 10836.208 | 318304 | 31705.016 | 0.003 | 0.997 | 1.82 | 1.9797 | 8.77
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11246.253 | 331176 | 33909.344 | 0.003 | 0.997 | 1.82 | 1.9191 | 5.45
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11218.586 | 341307 | 33640.234 | 0.003 | 0.997 | 1.82 | 1.9301 | 6.05
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11494.638 | 336124 | 34649.969 | 0.003 | 0.997 | 1.82 | 1.9196 | 5.47
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.65 | 12025.298 | 351434 | 35763.523 | 0.003 | 0.997 | 1.82 | 1.9466 | 6.95
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.65 | 11785.349 | 348689 | 35477.816 | 0.003 | 0.997 | 1.82 | 1.9223 | 5.62
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.65 | 12088.213 | 356663 | 36225.75 | 0.003 | 0.997 | 1.82 | 1.9313 | 6.11
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.65 | 15626.126 | 461167 | 34305.039 | 0.005 | 0.997 | 2.91 | 2.6603 | -8.58
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.65 | 17015.816 | 503700 | 36403.238 | 0.005 | 0.997 | 2.91 | 2.7317 | -6.13
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.65 | 25204.344 | 744851 | 33209.145 | 0.008 | 0.997 | 4.66 | 4.4804 | -3.85
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.65 | 26487.211 | 777943 | 35900.168 | 0.007 | 0.997 | 4.66 | 4.3533 | -6.58
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.65 | 39542.523 | 1165600 | 30390.211 | 0.013 | 0.997 | 7.45 | 7.7468 | 3.98
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.65 | 43067.457 | 1273534 | 33082.625 | 0.013 | 0.997 | 7.45 | 7.7507 | 4.04
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.65 | 70232.273 | 2073122 | 32738.229 | 0.021 | 0.997 | 11.92 | 12.8693 | 7.96
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.65 | 75099.82 | 2225947 | 34432.25 | 0.022 | 0.997 | 11.92 | 13.0877 | 9.8
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.65 | 125018.797 | 3684048 | 34820.98 | 0.036 | 0.997 | 19.07 | 21.7474 | 14.04
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.65 | 133696.859 | 3914060 | 36870.191 | 0.036 | 0.997 | 19.07 | 21.9692 | 15.2
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.65 | 181213.188 | 5323364 | 35522.934 | 0.051 | 0.997 | 30.52 | 31.1834 | 2.17
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.65 | 189827.453 | 5621227 | 37336.027 | 0.051 | 0.997 | 30.52 | 31.0763 | 1.82
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.65 | 272351.188 | 7932075 | 34067.25 | 0.08 | 0.997 | 48.83 | 49.7182 | 1.82
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.65 | 289306.313 | 8436817 | 36169.313 | 0.08 | 0.997 | 48.83 | 49.7453 | 1.87
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.65 | 416492.813 | 12135428 | 34067.969 | 0.122 | 0.997 | 78.13 | 78.0479 | -0.11
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.65 | 433477.563 | 12554088 | 34693.082 | 0.125 | 0.997 | 78.13 | 79.9052 | 2.27
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.65 | 584937.688 | 16983996 | 32972.277 | 0.177 | 0.997 | 125 | 117.5629 | -5.95
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.65 | 621589.5 | 17811714 | 36845.703 | 0.169 | 0.997 | 125 | 111.1073 | -11.11
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.65 | 749184.313 | 21513114 | 33862.289 | 0.221 | 0.997 | 156.25 | 151.556 | -3
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.65 | 799396.875 | 22652376 | 36388.051 | 0.22 | 0.997 | 156.25 | 150.3022 | -3.81
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.65 | 1160101.125 | 32674774 | 34733.418 | 0.334 | 0.997 | 250 | 254.8213 | 1.93
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.65 | 1213751.25 | 33633596 | 35195.254 | 0.345 | 0.997 | 250 | 266.5475 | 6.62
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.65 | 15.691 | 361 |  |  | 0.997
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.67 | 6.266 | 160 |  |  | 0.997
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.67 | 15.721 | 454 |  |  | 0.997
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.68 | 9.299 | 203 |  |  | 0.997
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.66 | 40.602 | 772 | 31803.217 | 0 | 0.997
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.65 | 63.036 | 1855 | 32830.207 | 0 | 0.997
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.65 | 65.394 | 1925 | 33975.277 | 0 | 0.997
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.65 | 59.23 | 1431 | 35277.617 | 0 | 0.997
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.65 | 212.124 | 6384 | 35834.434 | 0 | 0.997
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.65 | 54.949 | 1671 | 36040.789 | 0 | 0.997
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.65 | 122.007 | 3292 | 37023.516 | 0 | 0.997
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.65 | 66.616 | 1643 | 37009.66 | 0 | 0.997
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.65 | 22.222 | 484 | 12664.324 | 0 | 0.997
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.65 | 3654.858 | 107228 | 32111.395 | 0.001 | 0.997 | 0.63 | 0.61774 | -1.95
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.65 | 3979.43 | 116367 | 35430.559 | 0.001 | 0.997 | 0.63 | 0.60878 | -3.37
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.65 | 17056.02 | 495244 | 35946.789 | 0.005 | 0.997 | 2.5 | 2.774 | 10.96
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.65 | 19007.379 | 562991 | 39598.043 | 0.005 | 0.997 | 2.5 | 2.8071 | 12.28
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.65 | 39337.004 | 1152457 | 34698.945 | 0.011 | 0.997 | 6.25 | 6.7358 | 7.77
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.65 | 42555.238 | 1260814 | 38674.035 | 0.011 | 0.997 | 6.25 | 6.5349 | 4.56
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.65 | 150432.547 | 4406268 | 35005.285 | 0.043 | 0.997 | 25 | 26.1433 | 4.57
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.65 | 156770.422 | 4610806 | 35777.035 | 0.044 | 0.997 | 25 | 26.6706 | 6.68
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.67 | 7.649 | 267 |  |  | 0.997
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.64 | 20.492 | 490 |  |  | 0.997
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.71 | 14.427 | 319 |  |  | 0.997
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.61 | 5.261 | 129 |  |  | 0.997
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.65 | 139.596 | 3672 |  |  | 0.997
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.65 | 23.374 | 439 |  |  | 0.997
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.64 | 2.933 | 160 |  |  | 0.997
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.65 | 7.274 | 181 |  |  | 0.997
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.75 | 29.395 | 394 |  |  | 0.997
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.65 | 190.765 | 5700 |  |  | 0.997
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.66 | 48.535 | 646 |  |  | 0.997
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [5 empty] | 0.997
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.66 | 13.691 | 211 |  |  | 0.997
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 2.75 | 49.989 | 547 |  |  | 0.997
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.65 | 2.424 | 74 |  |  | 0.997
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.65 | 20.206 | 751 | 35815.477 | 0 | 0.997
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.65 | 48.589 | 954 | 38437.145 | 0 | 0.997
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.65 | 36.054 | 960 | 36556.488 | 0 | 0.997
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.65 | 3861.719 | 116503 | 37541.629 | 0.001 | 0.997 |  | 0.55238
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.65 | 3670.254 | 108089 | 35548.031 | 0.001 | 0.997 |  | 0.55467
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.65 | 4291.371 | 125681 | 35730.238 | 0.001 | 0.997 |  | 0.65526
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.66 | 46.087 | 738 | 37072.477 | 0 | 0.997
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.66 | 22.878 | 332 | 38052.293 | 0 | 0.997
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.65 | 35.243 | 1118 | 34841.289 | 0 | 0.997
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.65 | 37.985 | 1123 | 33680.504 | 0 | 0.997
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.66 | 20.323 | 481 | 33313.039 | 0 | 0.997
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.65 | 75.389 | 1965 | 35218.301 | 0 | 0.997
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.65 | 637827.313 | 18381228 | 31820.521 | 0.2 | 0.997 |  | 135.1066
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.65 | 689029.875 | 19863128 | 35913.09 | 0.192 | 0.997 |  | 128.4911
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.65 | 696401.188 | 19752984 | 34425.738 | 0.202 | 0.997 |  | 136.5417
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.65 | 272.514 | 7109 | 33026.402 | 0 | 0.997
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.65 | 84.635 | 2296 | 33342.656 | 0 | 0.997
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.65 | 103.472 | 2750 | 35495.684 | 0 | 0.997
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.65 | 70.968 | 1289 | 41065.688 | 0 | 0.997
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.65 | 99.183 | 2854 | 36334.883 | 0 | 0.997
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.65 | 40.615 | 1160 | 33962.336 | 0 | 0.997
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.65 | 662656.688 | 18908058 | 33873.305 | 0.196 | 0.997 |  | 131.3825
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.65 | 686166 | 19567960 | 34683.414 | 0.198 | 0.997 |  | 133.0863
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.65 | 685958.188 | 19722810 | 34392.777 | 0.199 | 0.997 |  | 134.3334
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.65 | 243.215 | 6252 | 34104.559 | 0 | 0.997
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.65 | 56.921 | 1696 | 33587.531 | 0 | 0.997
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.65 | 104.988 | 2933 | 34354.273 | 0 | 0.997
M7PFUdA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.58 | 32447.529 | 974791 | 32447.529 |  |  | 0.01 | 0.00928 | -7.17
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.58 | 32869.164 | 971478 | 32869.164 |  |  | 0.01 | 0.0094 | -5.96
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.58 | 34714.773 | 1029571 | 34714.773 |  |  | 0.01 | 0.00993 | -0.68
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.58 | 33756.629 | 983178 | 33756.629 |  |  | 0.01 | 0.00966 | -3.42
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.58 | 36119.102 | 1081464 | 36119.102 |  |  | 0.01 | 0.01033 | 3.34
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.58 | 35462.613 | 1070005 | 35462.613 |  |  | 0.01 | 0.01015 | 1.46
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.58 | 35970.531 | 1074156 | 35970.531 |  |  | 0.01 | 0.01029 | 2.91
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.58 | 32267.967 | 986301 | 32267.967 |  |  | 0.01 | 0.00923 | -7.68
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.58 | 33731.746 | 994147 | 33731.746 |  |  | 0.01 | 0.00965 | -3.49
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.58 | 36595.254 | 1075579 | 36595.254 |  |  | 0.01 | 0.01047 | 4.7
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.58 | 36770.707 | 1084336 | 36770.707 |  |  | 0.01 | 0.01052 | 5.2
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37791.871 | 1131169 | 37791.871 |  |  | 0.01 | 0.01081 | 8.12
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37730.301 | 1124057 | 37730.301 |  |  | 0.01 | 0.01079 | 7.95
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.58 | 37755.715 | 1135042 | 37755.715 |  |  | 0.01 | 0.0108 | 8.02
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.58 | 33209.082 | 973811 | 33209.082 |  |  | 0.01 | 0.0095 | -4.99
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.58 | 35506.813 | 1053187 | 35506.813 |  |  | 0.01 | 0.01016 | 1.59
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.58 | 34187.922 | 1004147 | 34187.922 |  |  | 0.01 | 0.00978 | -2.19
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.58 | 36297.711 | 1072373 | 36297.711 |  |  | 0.01 | 0.01038 | 3.85
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.58 | 37101.148 | 1109658 | 37101.148 |  |  | 0.01 | 0.01061 | 6.15
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.58 | 37717.949 | 1137986 | 37717.949 |  |  | 0.01 | 0.01079 | 7.91
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.58 | 36696.27 | 1092910 | 36696.27 |  |  | 0.01 | 0.0105 | 4.99
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.58 | 33202.188 | 1009091 | 33202.188 |  |  | 0.01 | 0.0095 | -5.01
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.58 | 35507.883 | 1037444 | 35507.883 |  |  | 0.01 | 0.01016 | 1.59
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.58 | 34225.648 | 1002659 | 34225.648 |  |  | 0.01 | 0.00979 | -2.08
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36058.262 | 1061800 | 36058.262 |  |  | 0.01 | 0.01032 | 3.16
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36060.398 | 1072029 | 36060.398 |  |  | 0.01 | 0.01032 | 3.17
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36561.789 | 1080162 | 36561.789 |  |  | 0.01 | 0.01046 | 4.6
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.58 | 36844.078 | 1099248 | 36844.078 |  |  | 0.01 | 0.01054 | 5.41
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.58 | 32022.859 | 951242 | 32022.859 |  |  | 0.01 | 0.00916 | -8.38
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33292.348 | 989585 | 33292.348 |  |  | 0.01 | 0.00953 | -4.75
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33805.305 | 980715 | 33805.305 |  |  | 0.01 | 0.00967 | -3.28
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.58 | 33890.82 | 986507 | 33890.82 |  |  | 0.01 | 0.0097 | -3.04
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.58 | 35058.344 | 1046315 | 35058.344 |  |  | 0.01 | 0.01003 | 0.3
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.58 | 34807.703 | 1039396 | 34807.703 |  |  | 0.01 | 0.00996 | -0.41
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.58 | 35534.512 | 1064559 | 35534.512 |  |  | 0.01 | 0.01017 | 1.67
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.58 | 31705.016 | 963116 | 31705.016 |  |  | 0.01 | 0.00907 | -9.29
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.58 | 33909.344 | 1001427 | 33909.344 |  |  | 0.01 | 0.0097 | -2.98
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.58 | 33640.234 | 1009063 | 33640.234 |  |  | 0.01 | 0.00962 | -3.75
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.58 | 34649.969 | 1022090 | 34649.969 |  |  | 0.01 | 0.00991 | -0.87
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.58 | 35763.523 | 1070814 | 35763.523 |  |  | 0.01 | 0.01023 | 2.32
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.58 | 35477.816 | 1062062 | 35477.816 |  |  | 0.01 | 0.01015 | 1.5
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.58 | 36225.75 | 1082349 | 36225.75 |  |  | 0.01 | 0.01036 | 3.64
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.58 | 34305.039 | 1017728 | 34305.039 |  |  | 0.01 | 0.00981 | -1.85
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.58 | 36403.238 | 1080267 | 36403.238 |  |  | 0.01 | 0.01042 | 4.15
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.58 | 33209.145 | 983289 | 33209.145 |  |  | 0.01 | 0.0095 | -4.99
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.58 | 35900.168 | 1066387 | 35900.168 |  |  | 0.01 | 0.01027 | 2.71
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.58 | 30390.211 | 927562 | 30390.211 |  |  | 0.01 | 0.00869 | -13.05
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.58 | 33082.625 | 981951 | 33082.625 |  |  | 0.01 | 0.00947 | -5.35
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.58 | 32738.229 | 977888 | 32738.229 |  |  | 0.01 | 0.00937 | -6.33
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.58 | 34432.25 | 1023852 | 34432.25 |  |  | 0.01 | 0.00985 | -1.49
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.58 | 34820.98 | 1040287 | 34820.98 |  |  | 0.01 | 0.00996 | -0.38
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.58 | 36870.191 | 1095480 | 36870.191 |  |  | 0.01 | 0.01055 | 5.49
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.58 | 35522.934 | 1047819 | 35522.934 |  |  | 0.01 | 0.01016 | 1.63
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.58 | 37336.027 | 1108892 | 37336.027 |  |  | 0.01 | 0.01068 | 6.82
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.58 | 34067.25 | 1018919 | 34067.25 |  |  | 0.01 | 0.00975 | -2.53
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.58 | 36169.313 | 1077802 | 36169.313 |  |  | 0.01 | 0.01035 | 3.48
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.58 | 34067.969 | 1023156 | 34067.969 |  |  | 0.01 | 0.00975 | -2.53
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.58 | 34693.082 | 1046635 | 34693.082 |  |  | 0.01 | 0.00993 | -0.74
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.58 | 32972.277 | 988333 | 32972.277 |  |  | 0.01 | 0.00943 | -5.67
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.58 | 36845.703 | 1098532 | 36845.703 |  |  | 0.01 | 0.01054 | 5.42
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.58 | 33862.289 | 1002519 | 33862.289 |  |  | 0.01 | 0.00969 | -3.12
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.58 | 36388.051 | 1073869 | 36388.051 |  |  | 0.01 | 0.01041 | 4.11
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.58 | 34733.418 | 1034220 | 34733.418 |  |  | 0.01 | 0.00994 | -0.63
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.58 | 35195.254 | 1041674 | 35195.254 |  |  | 0.01 | 0.01007 | 0.69
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.58 | 31803.217 | 947983 | 31803.217 |  |  | 0.01 | 0.0091 | -9.01
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.58 | 32830.207 | 975030 | 32830.207 |  |  | 0.01 | 0.00939 | -6.07
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.58 | 33975.277 | 1013740 | 33975.277 |  |  | 0.01 | 0.00972 | -2.8
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.58 | 35277.617 | 1031353 | 35277.617 |  |  | 0.01 | 0.01009 | 0.93
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.58 | 35834.434 | 1074706 | 35834.434 |  |  | 0.01 | 0.01025 | 2.52
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.58 | 36040.789 | 1079855 | 36040.789 |  |  | 0.01 | 0.01031 | 3.11
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.58 | 37023.516 | 1096703 | 37023.516 |  |  | 0.01 | 0.01059 | 5.93
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.58 | 37009.66 | 1102895 | 37009.66 |  |  | 0.01 | 0.01059 | 5.89
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.58 | 12664.324 | 375482 | 12664.324 |  |  | 0.01 | 0.00362 | -63.77
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.58 | 32111.395 | 978162 | 32111.395 |  |  | 0.01 | 0.00919 | -8.13
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.58 | 35430.559 | 1047127 | 35430.559 |  |  | 0.01 | 0.01014 | 1.37
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.58 | 35946.789 | 1062522 | 35946.789 |  |  | 0.01 | 0.01028 | 2.85
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.58 | 39598.043 | 1175237 | 39598.043 |  |  | 0.01 | 0.01133 | 13.29
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.58 | 34698.945 | 1020781 | 34698.945 |  |  | 0.01 | 0.00993 | -0.72
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.58 | 38674.035 | 1156652 | 38674.035 |  |  | 0.01 | 0.01106 | 10.65
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.58 | 35005.285 | 1033252 | 35005.285 |  |  | 0.01 | 0.01002 | 0.15
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.58 | 35777.035 | 1055289 | 35777.035 |  |  | 0.01 | 0.01024 | 2.36
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.58 | 35815.477 | 1046068 | 35815.477 |  |  | 0.01 | 0.01025 | 2.47
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.58 | 38437.145 | 1124788 | 38437.145 |  |  | 0.01 | 0.011 | 9.97
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.58 | 36556.488 | 1079330 | 36556.488 |  |  | 0.01 | 0.01046 | 4.59
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.58 | 37541.629 | 1112351 | 37541.629 |  |  | 0.01 | 0.01074 | 7.41
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.58 | 35548.031 | 1050255 | 35548.031 |  |  | 0.01 | 0.01017 | 1.7
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.58 | 35730.238 | 1047871 | 35730.238 |  |  | 0.01 | 0.01022 | 2.23
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.58 | 37072.477 | 1098487 | 37072.477 |  |  | 0.01 | 0.01061 | 6.07
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.58 | 38052.293 | 1127983 | 38052.293 |  |  | 0.01 | 0.01089 | 8.87
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.58 | 34841.289 | 1025977 | 34841.289 |  |  | 0.01 | 0.00997 | -0.32
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.58 | 33680.504 | 1011129 | 33680.504 |  |  | 0.01 | 0.00964 | -3.64
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.58 | 33313.039 | 994208 | 33313.039 |  |  | 0.01 | 0.00953 | -4.69
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.58 | 35218.301 | 1035739 | 35218.301 |  |  | 0.01 | 0.01008 | 0.76
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.58 | 31820.521 | 953848 | 31820.521 |  |  | 0.01 | 0.0091 | -8.96
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.58 | 35913.09 | 1059917 | 35913.09 |  |  | 0.01 | 0.01027 | 2.75
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.58 | 34425.738 | 1012072 | 34425.738 |  |  | 0.01 | 0.00985 | -1.51
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.58 | 33026.402 | 977107 | 33026.402 |  |  | 0.01 | 0.00945 | -5.51
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.58 | 33342.656 | 982047 | 33342.656 |  |  | 0.01 | 0.00954 | -4.61
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.58 | 35495.684 | 1043413 | 35495.684 |  |  | 0.01 | 0.01016 | 1.55
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.58 | 41065.688 | 1199133 | 41065.688 |  |  | 0.01 | 0.01175 | 17.49
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.58 | 36334.883 | 1074823 | 36334.883 |  |  | 0.01 | 0.0104 | 3.96
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.58 | 33962.336 | 1005645 | 33962.336 |  |  | 0.01 | 0.00972 | -2.83
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.58 | 33873.305 | 1004033 | 33873.305 |  |  | 0.01 | 0.00969 | -3.09
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.58 | 34683.414 | 1029481 | 34683.414 |  |  | 0.01 | 0.00992 | -0.77
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.58 | 34392.777 | 1023789 | 34392.777 |  |  | 0.01 | 0.00984 | -1.6
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.58 | 34104.559 | 997654 | 34104.559 |  |  | 0.01 | 0.00976 | -2.43
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.58 | 33587.531 | 992023 | 33587.531 |  |  | 0.01 | 0.00961 | -3.9
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.58 | 34354.273 | 1023908 | 34354.273 |  |  | 0.01 | 0.00983 | -1.71

## DTXSID90868151
*Analyte in Mix 2 | [8 empty] | PFTrDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.25 | 361.462 | 9725 | 9730.571 | 0 | 0.994 | 0.17 | 0.11196 | -34.14
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.25 | 305.135 | 7545 | 11255 | 0 | 0.994 | 0.17 | 0.05845 | -65.62
UC_AF_Mix2_A | 15.8488 | 507.1606 | 0.048324 | 0.043263 | 0.0054205 | 12.53% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.25 | 332.151 | 8955 | 11228.034 | 0 | 0.994 | 0.17 | 0.07162 | -57.87
UC_AF_Mix2_B | 14.2689 | 456.6042 | 0.037543 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.25 | 334.842 | 8353 | 10426.968 | 0 | 0.994 | 0.17 | 0.08512 | -49.93
UC_AF_Mix2_C | 14.8474 | 475.1165 | 0.043922 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.25 | 390.738 | 9523 | 9798.542 | 0 | 0.994 | 0.17 | 0.12651 | -25.58
UC_T1hr_Mix2_A | 135.7308 | 10858.4648 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.25 | 291.109 | 7256 | 9861.838 | 0 | 0.994 | 0.17 | 0.07129 | -58.07
UC_T1hr_Mix2_B | 123.2462 | 9859.6992 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.25 | 369.467 | 9542 | 8900.187 | 0 | 0.994 | 0.17 | 0.13523 | -20.45
UC_T1hr_Mix2_C | 137.4478 | 10995.8264 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.25 | 566.751 | 15043 | 11120.537 | 0.001 | 0.994 | 0.28 | 0.18563 | -33.7
UC_T5hr_Mix2_A | 131.1884 | 10495.0744 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.25 | 629.304 | 16713 | 12190.335 | 0.001 | 0.994 | 0.28 | 0.18914 | -32.45
UC_T5hr_Mix2_B | 152.0277 | 12162.216 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.25 | 518.194 | 14739 | 11829.166 | 0 | 0.994 | 0.28 | 0.14746 | -47.33
UC_T5hr_Mix2_C | 135.2173 | 10817.3848 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.25 | 582.102 | 14804 | 11531.981 | 0.001 | 0.994 | 0.28 | 0.18303 | -34.63
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.25 | 564.932 | 15445 | 10690.253 | 0.001 | 0.994 | 0.28 | 0.19566 | -30.12
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.25 | 589.088 | 15622 | 10913.814 | 0.001 | 0.994 | 0.28 | 0.20169 | -27.97
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.25 | 530.259 | 14582 | 9899.175 | 0.001 | 0.994 | 0.28 | 0.1995 | -28.75
20201123_PFAS_PPB_UC_Sample010 | CC4 | 0.71 | 0.54778 | 0.57883 | 0.05306 | 9.167% | 0.16677 | -18.47% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1490.313 | 38142 | 10876.152 | 0.001 | 0.994 | 0.44 | 0.64466 | 46.51
20201123_PFAS_PPB_UC_Sample035 | CC4 | 0.71 | 0.5029 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1584.214 | 41103 | 11718.064 | 0.001 | 0.994 | 0.44 | 0.63489 | 44.29
20201123_PFAS_PPB_UC_Sample042 | CC4 | 0.71 | 0.54509 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1553.966 | 40896 | 11542.465 | 0.001 | 0.994 | 0.44 | 0.63188 | 43.61
20201123_PFAS_PPB_UC_Sample082 | CC4 | 0.71 | 0.56766 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1453.744 | 41948 | 10908.972 | 0.001 | 0.994 | 0.44 | 0.62458 | 41.95
20201123_PFAS_PPB_UC_Sample108 | CC4 | 0.71 | 0.62475 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1654.718 | 43523 | 10546.275 | 0.002 | 0.994 | 0.44 | 0.7507 | 70.61
20201123_PFAS_PPB_UC_Sample115 | CC4 | 0.71 | 0.65605 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1572.219 | 38941 | 9692.394 | 0.002 | 0.994 | 0.44 | 0.77904 | 77.06
20201123_PFAS_PPB_UC_Sample122 | CC4 | 0.71 | 0.60761 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.25 | 1511.87 | 41215 | 9065.441 | 0.002 | 0.994 | 0.44 | 0.80338 | 82.59
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1231.507 | 33415 | 10360.407 | 0.001 | 0.994 | 0.71 | 0.54778 | -22.85
Blanks should be less than… | 0.083385 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1250.094 | 32864 | 11317.867 | 0.001 | 0.994 | 0.71 | 0.5029 | -29.17
eLOQ | 0.71 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1403.119 | 37945 | 11854.473 | 0.001 | 0.994 | 0.71 | 0.54509 | -23.23
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1271.26 | 34273 | 10369.82 | 0.001 | 0.994 | 0.71 | 0.56766 | -20.05
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1374.593 | 37593 | 10312.553 | 0.001 | 0.994 | 0.71 | 0.62475 | -12.01
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1355.567 | 35995 | 9741.055 | 0.001 | 0.994 | 0.71 | 0.65605 | -7.6
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.25 | 1248.689 | 32255 | 9599.27 | 0.001 | 0.994 | 0.71 | 0.60761 | -14.42
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2790.359 | 74985 | 10282.522 | 0.003 | 0.994 | 1.14 | 1.3617 | 19.45
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2686.308 | 72115 | 10528.761 | 0.003 | 0.994 | 1.14 | 1.2751 | 11.85
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2752.255 | 76585 | 10775.113 | 0.003 | 0.994 | 1.14 | 1.2766 | 11.98
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2877.709 | 74907 | 9679.873 | 0.003 | 0.994 | 1.14 | 1.5001 | 31.59
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2912.181 | 77185 | 9162.68 | 0.003 | 0.994 | 1.14 | 1.6099 | 41.22
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2749.095 | 72310 | 8925.449 | 0.003 | 0.994 | 1.14 | 1.5574 | 36.61
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.25 | 2588.253 | 67985 | 8059.229 | 0.003 | 0.994 | 1.14 | 1.6276 | 42.77
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3754.326 | 101025 | 10162.122 | 0.004 | 0.994 | 1.82 | 1.8856 | 3.6
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3925.793 | 105378 | 11242.785 | 0.003 | 0.994 | 1.82 | 1.7773 | -2.34
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3868.403 | 104825 | 10874.418 | 0.004 | 0.994 | 1.82 | 1.8123 | -0.42
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.25 | 4052.725 | 110355 | 10265.066 | 0.004 | 0.994 | 1.82 | 2.0211 | 11.05
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3771.753 | 100105 | 9531.285 | 0.004 | 0.994 | 1.82 | 2.026 | 11.32
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3863.057 | 105763 | 9877.485 | 0.004 | 0.994 | 1.82 | 2.0013 | 9.96
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.25 | 3943.559 | 106325 | 8519.921 | 0.005 | 0.994 | 1.82 | 2.3849 | 31.04
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.25 | 4767.658 | 128384 | 10490.147 | 0.005 | 0.994 | 2.91 | 2.3401 | -19.58
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.25 | 4873.014 | 136983 | 11516.313 | 0.004 | 0.994 | 2.91 | 2.1725 | -25.34
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.25 | 10289.728 | 279833 | 9708.633 | 0.011 | 0.994 | 4.66 | 5.5828 | 19.8
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.25 | 11089.063 | 302811 | 9848.446 | 0.011 | 0.994 | 4.66 | 5.9377 | 27.42
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.25 | 14118.135 | 384250 | 9783.861 | 0.014 | 0.994 | 7.45 | 7.6417 | 2.57
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.25 | 15005.965 | 402023 | 10084.969 | 0.015 | 0.994 | 7.45 | 7.8836 | 5.82
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.25 | 21835.18 | 595031 | 11453.623 | 0.019 | 0.994 | 11.92 | 10.1387 | -14.94
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.25 | 22863.998 | 623795 | 10773.109 | 0.021 | 0.994 | 11.92 | 11.3048 | -5.16
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.25 | 44740.945 | 1217712 | 10865.012 | 0.041 | 0.994 | 19.07 | 22.161 | 16.21
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.25 | 45678.59 | 1249665 | 11209.78 | 0.041 | 0.994 | 19.07 | 21.9256 | 14.97
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.25 | 60740.434 | 1636760 | 11101.751 | 0.055 | 0.994 | 30.52 | 29.6069 | -2.99
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.25 | 63296.691 | 1727377 | 10679.77 | 0.059 | 0.994 | 30.52 | 32.1286 | 5.27
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.25 | 92595.516 | 2486691 | 10508.111 | 0.088 | 0.994 | 48.83 | 48.2876 | -1.11
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.25 | 92582.195 | 2506564 | 10697.722 | 0.087 | 0.994 | 48.83 | 47.3973 | -2.93
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.25 | 133000.719 | 3594597 | 9989.264 | 0.133 | 0.994 | 78.13 | 74.1923 | -5.04
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.25 | 143420.453 | 3838232 | 10158.705 | 0.141 | 0.994 | 78.13 | 78.9092 | 1
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.25 | 208385.203 | 5623033 | 10623.775 | 0.196 | 0.994 | 125 | 111.9995 | -10.4
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.25 | 218987.391 | 5837530 | 10158.639 | 0.216 | 0.994 | 125 | 124.0568 | -0.75
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.25 | 278826.625 | 7470329 | 10923.09 | 0.255 | 0.994 | 156.25 | 149.3655 | -4.41
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.25 | 278772.781 | 7373114 | 9865.15 | 0.283 | 0.994 | 156.25 | 167.3413 | 7.1
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.25 | 386971.281 | 10065610 | 9730.678 | 0.398 | 0.994 | 250 | 249.1041 | -0.36
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.25 | 401981.25 | 10638402 | 9832.22 | 0.409 | 0.994 | 250 | 257.6472 | 3.06
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 3.3 | 43.052 | 544 |  |  | 0.994
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 3.25 | 43.496 | 761 |  |  | 0.994
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 3.21 | 6.538 | 147 |  |  | 0.994
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank |  |  |  | 3.524 |  | 0.994
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.25 | 18.166 | 557 | 10173.215 | 0 | 0.994
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.25 | 58.275 | 1024 | 11173.664 | 0 | 0.994
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.24 | 8.5 | 344 | 12048.315 | 0 | 0.994
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.25 | 30.159 | 885 | 11975.843 | 0 | 0.994
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.25 | 81.038 | 2420 | 10773.584 | 0 | 0.994
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank |  |  |  | 10459.238 |  | 0.994
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.31 | 12.677 | 214 | 10166.878 | 0 | 0.994
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.24 | 116.384 | 1307 | 9387.197 | 0 | 0.994
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.26 | 32.978 | 469 | 4318.83 | 0 | 0.994
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.25 | 1071.004 | 29043 | 10475.948 | 0.001 | 0.994 | 0.63 | 0.45906 | -27.13
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.25 | 1148.222 | 30281 | 10532.542 | 0.001 | 0.994 | 0.63 | 0.49524 | -21.39
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.25 | 7202.33 | 195015 | 11635.099 | 0.006 | 0.994 | 2.5 | 3.2202 | 28.81
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.25 | 5734.268 | 154740 | 6178.086 | 0.009 | 0.994 | 2.5 | 4.8763 | 95.05
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.25 | 13334.586 | 368465 | 12150.41 | 0.011 | 0.994 | 6.25 | 5.7846 | -7.45
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.25 | 12606.851 | 342443 | 10931.824 | 0.012 | 0.994 | 6.25 | 6.084 | -2.66
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.25 | 46874.168 | 1286494 | 10343.043 | 0.045 | 0.994 | 25 | 24.4318 | -2.27
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.25 | 47936.488 | 1305538 | 9113.911 | 0.053 | 0.994 | 25 | 28.4385 | 13.75
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 3.24 | 64.051 | 641 |  |  | 0.994
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 3.24 | 66.092 | 715 |  |  | 0.994
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 3.2 | 50.479 | 528 |  |  | 0.994
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 3.21 | 1.82 | 169 |  |  | 0.994
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 3.25 | 87.035 | 1492 |  |  | 0.994
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [5 empty] | 0.994
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 3.28 | 56.314 | 578 |  |  | 0.994
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [5 empty] | 0.994
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 3.19 | 91.734 | 574 |  |  | 0.994
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 3.25 | 191.709 | 1927 |  |  | 0.994
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 3.26 | 89.381 | 786 |  |  | 0.994
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 3.2 | 49.8 | 661 |  |  | 0.994
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 3.21 | 59.636 | 545 |  |  | 0.994
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 3.15 | 83.196 | 909 |  |  | 0.994
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 3.16 | 66.668 | 687 |  |  | 0.994
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.21 | 84.168 | 702 | 11440.534 | 0 | 0.994
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.25 | 133.335 | 1209 | 13256.198 | 0 | 0.994
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte |  |  |  | 11719.4 |  | 0.994
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.25 | 33506.41 | 915674 | 11316.626 | 0.03 | 0.994 |  | 15.8488
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.25 | 29692.799 | 819239 | 11121.704 | 0.027 | 0.994 |  | 14.2689
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.25 | 31607.668 | 866983 | 11384.287 | 0.028 | 0.994 |  | 14.8474
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.25 | 54.334 | 622 | 11676.331 | 0 | 0.994
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.23 | 42.634 | 734 | 11894.549 | 0 | 0.994
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.24 | 28.432 | 575 | 11048.28 | 0 | 0.994
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.25 | 159.608 | 2999 | 9862.112 | 0 | 0.994 |  | 0.00019
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.25 | 131.481 | 3142 | 9161.524 | 0 | 0.994
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.25 | 134.293 | 3308 | 9049.197 | 0 | 0.994
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.25 | 208258.063 | 5472016 | 8897.753 | 0.234 | 0.994 |  | 135.7308
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.25 | 231957.547 | 6248865 | 10825.349 | 0.214 | 0.994 |  | 123.2462
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.25 | 236618.672 | 6307770 | 9994.433 | 0.237 | 0.994 |  | 137.4478
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.25 | 147.506 | 3836 | 8190.83 | 0 | 0.994 |  | 0.00992
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.25 | 84.121 | 1180 | 9179.411 | 0 | 0.994
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.25 | 62.151 | 1381 | 10046.605 | 0 | 0.994
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.25 | 106.976 | 2961 | 11574.24 | 0 | 0.994
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.25 | 149.675 | 2844 | 8454.886 | 0 | 0.994 |  | 0.00829
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.25 | 128.604 | 2937 | 8997.701 | 0 | 0.994
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.25 | 214740.047 | 5692729 | 9464.119 | 0.227 | 0.994 |  | 131.1884
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.25 | 216414.156 | 5761683 | 8344.308 | 0.259 | 0.994 |  | 152.0277
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.25 | 209803.172 | 5591388 | 8994.774 | 0.233 | 0.994 |  | 135.2173
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.24 | 65.747 | 1251 | 9773.682 | 0 | 0.994
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.24 | 40.735 | 798 | 9471.696 | 0 | 0.994
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.25 | 164.247 | 1293 | 9756.347 | 0 | 0.994 |  | 0.00366
M2PFTeDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9730.571 | 329079 | 9730.571 |  |  | 0.01 | 0.00936 | -6.38
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11255 | 376966 | 11255 |  |  | 0.01 | 0.01083 | 8.29
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11228.034 | 376645 | 11228.034 |  |  | 0.01 | 0.0108 | 8.03
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.48 | 10426.968 | 352207 | 10426.968 |  |  | 0.01 | 0.01003 | 0.32
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9798.542 | 327493 | 9798.542 |  |  | 0.01 | 0.00943 | -5.72
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9861.838 | 334756 | 9861.838 |  |  | 0.01 | 0.00949 | -5.11
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.48 | 8900.187 | 298634 | 8900.187 |  |  | 0.01 | 0.00856 | -14.37
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11120.537 | 373820 | 11120.537 |  |  | 0.01 | 0.0107 | 7
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.48 | 12190.335 | 409431 | 12190.335 |  |  | 0.01 | 0.01173 | 17.29
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11829.166 | 397464 | 11829.166 |  |  | 0.01 | 0.01138 | 13.82
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11531.981 | 391568 | 11531.981 |  |  | 0.01 | 0.0111 | 10.96
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10690.253 | 359731 | 10690.253 |  |  | 0.01 | 0.01029 | 2.86
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10913.814 | 370373 | 10913.814 |  |  | 0.01 | 0.0105 | 5.01
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.48 | 9899.175 | 337103 | 9899.175 |  |  | 0.01 | 0.00952 | -4.75
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10876.152 | 369500 | 10876.152 |  |  | 0.01 | 0.01046 | 4.65
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11718.064 | 392785 | 11718.064 |  |  | 0.01 | 0.01127 | 12.75
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11542.465 | 386368 | 11542.465 |  |  | 0.01 | 0.01111 | 11.06
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10908.972 | 365733 | 10908.972 |  |  | 0.01 | 0.0105 | 4.96
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10546.275 | 355639 | 10546.275 |  |  | 0.01 | 0.01015 | 1.47
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9692.394 | 329243 | 9692.394 |  |  | 0.01 | 0.00933 | -6.74
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9065.441 | 305364 | 9065.441 |  |  | 0.01 | 0.00872 | -12.78
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10360.407 | 348348 | 10360.407 |  |  | 0.01 | 0.00997 | -0.32
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11317.867 | 377809 | 11317.867 |  |  | 0.01 | 0.01089 | 8.9
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11854.473 | 394417 | 11854.473 |  |  | 0.01 | 0.01141 | 14.06
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10369.82 | 347172 | 10369.82 |  |  | 0.01 | 0.00998 | -0.23
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10312.553 | 346986 | 10312.553 |  |  | 0.01 | 0.00992 | -0.78
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9741.055 | 331371 | 9741.055 |  |  | 0.01 | 0.00937 | -6.28
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9599.27 | 327251 | 9599.27 |  |  | 0.01 | 0.00924 | -7.64
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10282.522 | 349144 | 10282.522 |  |  | 0.01 | 0.00989 | -1.07
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10528.761 | 357528 | 10528.761 |  |  | 0.01 | 0.01013 | 1.3
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10775.113 | 363685 | 10775.113 |  |  | 0.01 | 0.01037 | 3.67
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9679.873 | 327049 | 9679.873 |  |  | 0.01 | 0.00931 | -6.86
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9162.68 | 310580 | 9162.68 |  |  | 0.01 | 0.00882 | -11.84
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8925.449 | 297809 | 8925.449 |  |  | 0.01 | 0.00859 | -14.12
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8059.229 | 274308 | 8059.229 |  |  | 0.01 | 0.00775 | -22.46
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10162.122 | 345524 | 10162.122 |  |  | 0.01 | 0.00978 | -2.22
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.48 | 11242.785 | 374717 | 11242.785 |  |  | 0.01 | 0.01082 | 8.17
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10874.418 | 363355 | 10874.418 |  |  | 0.01 | 0.01046 | 4.63
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10265.066 | 348695 | 10265.066 |  |  | 0.01 | 0.00988 | -1.23
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9531.285 | 324174 | 9531.285 |  |  | 0.01 | 0.00917 | -8.29
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9877.485 | 335723 | 9877.485 |  |  | 0.01 | 0.0095 | -4.96
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.48 | 8519.921 | 288423 | 8519.921 |  |  | 0.01 | 0.0082 | -18.02
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.48 | 10490.147 | 354799 | 10490.147 |  |  | 0.01 | 0.01009 | 0.93
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.48 | 11516.313 | 392960 | 11516.313 |  |  | 0.01 | 0.01108 | 10.81
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9708.633 | 329742 | 9708.633 |  |  | 0.01 | 0.00934 | -6.59
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9848.446 | 330014 | 9848.446 |  |  | 0.01 | 0.00948 | -5.24
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.48 | 9783.861 | 332235 | 9783.861 |  |  | 0.01 | 0.00941 | -5.86
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.48 | 10084.969 | 342025 | 10084.969 |  |  | 0.01 | 0.0097 | -2.97
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.48 | 11453.623 | 387651 | 11453.623 |  |  | 0.01 | 0.01102 | 10.2
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.48 | 10773.109 | 363168 | 10773.109 |  |  | 0.01 | 0.01037 | 3.65
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.48 | 10865.012 | 367049 | 10865.012 |  |  | 0.01 | 0.01045 | 4.54
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.48 | 11209.78 | 378373 | 11209.78 |  |  | 0.01 | 0.01079 | 7.86
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.48 | 11101.751 | 374089 | 11101.751 |  |  | 0.01 | 0.01068 | 6.82
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.48 | 10679.77 | 353296 | 10679.77 |  |  | 0.01 | 0.01028 | 2.76
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10508.111 | 351061 | 10508.111 |  |  | 0.01 | 0.01011 | 1.1
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10697.722 | 357232 | 10697.722 |  |  | 0.01 | 0.01029 | 2.93
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.48 | 9989.264 | 334986 | 9989.264 |  |  | 0.01 | 0.00961 | -3.89
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.48 | 10158.705 | 334343 | 10158.705 |  |  | 0.01 | 0.00977 | -2.26
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10623.775 | 354798 | 10623.775 |  |  | 0.01 | 0.01022 | 2.22
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10158.639 | 338366 | 10158.639 |  |  | 0.01 | 0.00977 | -2.26
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.48 | 10923.09 | 361685 | 10923.09 |  |  | 0.01 | 0.01051 | 5.1
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.48 | 9865.15 | 326284 | 9865.15 |  |  | 0.01 | 0.00949 | -5.08
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9730.678 | 316707 | 9730.678 |  |  | 0.01 | 0.00936 | -6.38
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9832.22 | 321241 | 9832.22 |  |  | 0.01 | 0.00946 | -5.4
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 3.49 | 3.524 | 114 | 3.524 |  |  | 0.01 | 0 | -99.97
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.48 | 10173.215 | 343416 | 10173.215 |  |  | 0.01 | 0.00979 | -2.12
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.48 | 11173.664 | 376859 | 11173.664 |  |  | 0.01 | 0.01075 | 7.51
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.48 | 12048.315 | 404321 | 12048.315 |  |  | 0.01 | 0.01159 | 15.92
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.48 | 11975.843 | 401621 | 11975.843 |  |  | 0.01 | 0.01152 | 15.23
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.48 | 10773.584 | 362156 | 10773.584 |  |  | 0.01 | 0.01037 | 3.66
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 3.48 | 10459.238 | 354868 | 10459.238 |  |  | 0.01 | 0.01006 | 0.63
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.48 | 10166.878 | 344723 | 10166.878 |  |  | 0.01 | 0.00978 | -2.18
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.48 | 9387.197 | 317571 | 9387.197 |  |  | 0.01 | 0.00903 | -9.68
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.48 | 4318.83 | 148626 | 4318.83 |  |  | 0.01 | 0.00416 | -58.45
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.48 | 10475.948 | 355039 | 10475.948 |  |  | 0.01 | 0.01008 | 0.8
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.48 | 10532.542 | 355947 | 10532.542 |  |  | 0.01 | 0.01013 | 1.34
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.48 | 11635.099 | 388325 | 11635.099 |  |  | 0.01 | 0.01119 | 11.95
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.48 | 6178.086 | 209790 | 6178.086 |  |  | 0.01 | 0.00594 | -40.56
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.48 | 12150.41 | 410208 | 12150.41 |  |  | 0.01 | 0.01169 | 16.91
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.48 | 10931.824 | 367413 | 10931.824 |  |  | 0.01 | 0.01052 | 5.18
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.48 | 10343.043 | 345742 | 10343.043 |  |  | 0.01 | 0.00995 | -0.48
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.48 | 9113.911 | 304569 | 9113.911 |  |  | 0.01 | 0.00877 | -12.31
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.48 | 11440.534 | 382724 | 11440.534 |  |  | 0.01 | 0.01101 | 10.08
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.48 | 13256.198 | 439493 | 13256.198 |  |  | 0.01 | 0.01275 | 27.55
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 3.48 | 11719.4 | 391810 | 11719.4 |  |  | 0.01 | 0.01128 | 12.76
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.48 | 11316.626 | 377070 | 11316.626 |  |  | 0.01 | 0.01089 | 8.88
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.48 | 11121.704 | 373796 | 11121.704 |  |  | 0.01 | 0.0107 | 7.01
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.48 | 11384.287 | 381386 | 11384.287 |  |  | 0.01 | 0.01095 | 9.53
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.48 | 11676.331 | 390146 | 11676.331 |  |  | 0.01 | 0.01123 | 12.34
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.48 | 11894.549 | 401063 | 11894.549 |  |  | 0.01 | 0.01144 | 14.44
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.48 | 11048.28 | 371884 | 11048.28 |  |  | 0.01 | 0.01063 | 6.3
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.48 | 9862.112 | 325453 | 9862.112 |  |  | 0.01 | 0.00949 | -5.11
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.48 | 9161.524 | 302491 | 9161.524 |  |  | 0.01 | 0.00881 | -11.85
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.48 | 9049.197 | 303067 | 9049.197 |  |  | 0.01 | 0.00871 | -12.93
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.48 | 8897.753 | 300237 | 8897.753 |  |  | 0.01 | 0.00856 | -14.39
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.48 | 10825.349 | 365981 | 10825.349 |  |  | 0.01 | 0.01042 | 4.16
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.48 | 9994.433 | 339009 | 9994.433 |  |  | 0.01 | 0.00962 | -3.84
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.48 | 8190.83 | 272286 | 8190.83 |  |  | 0.01 | 0.00788 | -21.19
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.48 | 9179.411 | 308907 | 9179.411 |  |  | 0.01 | 0.00883 | -11.68
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.48 | 10046.605 | 340119 | 10046.605 |  |  | 0.01 | 0.00967 | -3.34
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.48 | 11574.24 | 389337 | 11574.24 |  |  | 0.01 | 0.01114 | 11.36
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.48 | 8454.886 | 279574 | 8454.886 |  |  | 0.01 | 0.00813 | -18.65
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.48 | 8997.701 | 299688 | 8997.701 |  |  | 0.01 | 0.00866 | -13.43
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.48 | 9464.119 | 317767 | 9464.119 |  |  | 0.01 | 0.00911 | -8.94
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.48 | 8344.308 | 280834 | 8344.308 |  |  | 0.01 | 0.00803 | -19.71
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.48 | 8994.774 | 300266 | 8994.774 |  |  | 0.01 | 0.00865 | -13.46
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.48 | 9773.682 | 330647 | 9773.682 |  |  | 0.01 | 0.0094 | -5.96
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.48 | 9471.696 | 319932 | 9471.696 |  |  | 0.01 | 0.00911 | -8.87
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.48 | 9756.347 | 327921 | 9756.347 |  |  | 0.01 | 0.00939 | -6.13

## DTXSID3059921
*Analyte in Mix 1 | [8 empty] | PFTeDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.48 | 76.018 | 2541 | 9730.571 | 0 | 0.995 | 0.17 | 0.09251 | -45.58
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.48 | 96.457 | 3267 | 11255 | 0 | 0.995 | 0.17 | 0.11169 | -34.3
UC_AF_Mix1_A | 8.1092 | 259.495 | 0.028889 | 0.027055 | 0.0033392 | 12.34% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.48 | 88.639 | 3102 | 11228.034 | 0 | 0.995 | 0.17 | 0.09459 | -44.36
UC_AF_Mix1_B | 9.0871 | 290.7878 | 0.029075 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.48 | 93.003 | 2974 | 10426.968 | 0 | 0.995 | 0.17 | 0.12053 | -29.1
UC_AF_Mix1_C | 7.2619 | 232.3824 | 0.023201 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.48 | 89.666 | 2956 | 9798.542 | 0 | 0.995 | 0.17 | 0.12639 | -25.65
UC_T1hr_Mix1_A | 112.1514 | 8972.1104 | [6 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.48 | 87.609 | 2972 | 9861.838 | 0 | 0.995 | 0.17 | 0.11962 | -29.63
UC_T1hr_Mix1_B | 121.0757 | 9686.0552 | [6 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.48 | 66.422 | 2167 | 8900.187 | 0 | 0.995 | 0.17 | 0.08367 | -50.78
UC_T1hr_Mix1_C | 120.4974 | 9639.7944 | [6 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.48 | 150.396 | 5053 | 11120.537 | 0 | 0.995 | 0.28 | 0.23707 | -15.33
UC_T5hr_Mix1_A | 112.2807 | 8982.4552 | [6 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.48 | 160.949 | 5460 | 12190.335 | 0 | 0.995 | 0.28 | 0.22894 | -18.24
UC_T5hr_Mix1_B | 125.0164 | 10001.3128 | [6 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.48 | 149.478 | 5371 | 11829.166 | 0 | 0.995 | 0.28 | 0.2146 | -23.36
UC_T5hr_Mix1_C | 125.2025 | 10016.2 | [6 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.48 | 132.244 | 4290 | 11531.981 | 0 | 0.995 | 0.28 | 0.18502 | -33.92
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.48 | 126.176 | 4428 | 10690.253 | 0 | 0.995 | 0.28 | 0.19351 | -30.89
MDL Determination | [8 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.48 | 133.241 | 4606 | 10913.814 | 0 | 0.995 | 0.28 | 0.20377 | -27.22
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.48 | 134.751 | 4468 | 9899.175 | 0 | 0.995 | 0.28 | 0.2393 | -14.54
20201123_PFAS_PPB_UC_Sample011 | CC5 | 1.14 | 1.194 | 1.2416 | 0.047929 | 3.86% | 0.15064 | 8.91% | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.48 | 400.484 | 13499 | 10876.152 | 0 | 0.995 | 0.44 | 0.82675 | 87.9
20201123_PFAS_PPB_UC_Sample036 | CC5 | 1.14 | 1.1968 | [5 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.48 | 419.594 | 14218 | 11718.064 | 0 | 0.995 | 0.44 | 0.80107 | 82.06
20201123_PFAS_PPB_UC_Sample043 | CC5 | 1.14 | 1.2393 | [5 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.48 | 397.977 | 13732 | 11542.465 | 0 | 0.995 | 0.44 | 0.76745 | 74.42
20201123_PFAS_PPB_UC_Sample083 | CC5 | 1.14 | 1.329 | [5 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.48 | 397.124 | 13495 | 10908.972 | 0 | 0.995 | 0.44 | 0.81615 | 85.49
20201123_PFAS_PPB_UC_Sample109 | CC5 | 1.14 | 1.2723 | [5 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.48 | 359.254 | 12358 | 10546.275 | 0 | 0.995 | 0.44 | 0.75695 | 72.03
20201123_PFAS_PPB_UC_Sample116 | CC5 | 1.14 | 1.2118 | [5 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.48 | 346.95 | 11689 | 9692.394 | 0 | 0.995 | 0.44 | 0.80078 | 82
20201123_PFAS_PPB_UC_Sample123 | CC5 | 1.14 | 1.2478 | [5 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.48 | 338.576 | 11860 | 9065.441 | 0 | 0.995 | 0.44 | 0.84006 | 90.92
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.48 | 269.788 | 9343 | 10360.407 | 0 | 0.995 | 0.71 | 0.55385 | -21.99
Blanks should be less than… | 0.07532 | nM | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.48 | 286.428 | 9994 | 11317.867 | 0 | 0.995 | 0.71 | 0.53531 | -24.6
eLOQ | 1.14 | nM | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.48 | 280.302 | 9542 | 11854.473 | 0 | 0.995 | 0.71 | 0.49323 | -30.53
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.48 | 269.659 | 9183 | 10369.82 | 0 | 0.995 | 0.71 | 0.55294 | -22.12
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.48 | 270.067 | 9229 | 10312.553 | 0 | 0.995 | 0.71 | 0.55759 | -21.47
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.48 | 241.777 | 8413 | 9741.055 | 0 | 0.995 | 0.71 | 0.52298 | -26.34
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.48 | 252.29 | 8791 | 9599.27 | 0 | 0.995 | 0.71 | 0.55997 | -21.13
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.48 | 527.814 | 18224 | 10282.522 | 0.001 | 0.995 | 1.14 | 1.194 | 4.74
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.48 | 541.634 | 18738 | 10528.761 | 0.001 | 0.995 | 1.14 | 1.1968 | 4.99
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.48 | 572.384 | 19665 | 10775.113 | 0.001 | 0.995 | 1.14 | 1.2393 | 8.71
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.48 | 548.507 | 18901 | 9679.873 | 0.001 | 0.995 | 1.14 | 1.329 | 16.58
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.48 | 498.669 | 17122 | 9162.68 | 0.001 | 0.995 | 1.14 | 1.2723 | 11.6
QC FLAGS | [8 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.48 | 464.422 | 16113 | 8925.449 | 0.001 | 0.995 | 1.14 | 1.2118 | 6.3
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.48 | 430.803 | 15089 | 8059.229 | 0.001 | 0.995 | 1.14 | 1.2478 | 9.45
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.735 | 24808 | 10162.122 | 0.001 | 0.995 | 1.82 | 1.6949 | -6.87
Blanks |  | >1/2 MDL of analytes | [6 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.48 | 827.981 | 28603 | 11242.785 | 0.001 | 0.995 | 1.82 | 1.7589 | -3.36
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.48 | 759.402 | 25597 | 10874.418 | 0.001 | 0.995 | 1.82 | 1.6624 | -8.66
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.994 | 24870 | 10265.066 | 0.001 | 0.995 | 1.82 | 1.6775 | -7.83
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.48 | 691.849 | 23128 | 9531.285 | 0.001 | 0.995 | 1.82 | 1.7321 | -4.83
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.48 | 722.285 | 24130 | 9877.485 | 0.001 | 0.995 | 1.82 | 1.7457 | -4.08
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.48 | 631.31 | 21509 | 8519.921 | 0.001 | 0.995 | 1.82 | 1.7703 | -2.73
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.48 | 961.143 | 32731 | 10490.147 | 0.001 | 0.995 | 2.91 | 2.214 | -23.92
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.48 | 1026.99 | 34671 | 11516.313 | 0.001 | 0.995 | 2.91 | 2.1521 | -26.05
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.48 | 2133.045 | 71385 | 9708.633 | 0.002 | 0.995 | 4.66 | 5.4574 | 17.11
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.48 | 2207.919 | 71839 | 9848.446 | 0.002 | 0.995 | 4.66 | 5.571 | 19.55
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.48 | 2800.844 | 93515 | 9783.861 | 0.003 | 0.995 | 7.45 | 7.1437 | -4.11
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.48 | 2886.225 | 93590 | 10084.969 | 0.003 | 0.995 | 7.45 | 7.1416 | -4.14
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.48 | 4996.432 | 166370 | 11453.623 | 0.004 | 0.995 | 11.92 | 10.944 | -8.19
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.48 | 4827.986 | 156305 | 10773.109 | 0.004 | 0.995 | 11.92 | 11.2462 | -5.65
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.48 | 9734.145 | 318777 | 10865.012 | 0.009 | 0.995 | 19.07 | 22.607 | 18.55
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.48 | 10000.793 | 325065 | 11209.78 | 0.009 | 0.995 | 19.07 | 22.5113 | 18.05
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.48 | 13404.201 | 440392 | 11101.751 | 0.012 | 0.995 | 30.52 | 30.5213 | 0
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.48 | 13072.131 | 428255 | 10679.77 | 0.012 | 0.995 | 30.52 | 30.9437 | 1.39
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.48 | 19948.268 | 645349 | 10508.111 | 0.019 | 0.995 | 48.83 | 48.1116 | -1.47
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.48 | 19642.855 | 633581 | 10697.722 | 0.018 | 0.995 | 48.83 | 46.5263 | -4.72
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.48 | 30695.457 | 997521 | 9989.264 | 0.031 | 0.995 | 78.13 | 78.1173 | -0.02
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.48 | 31198.275 | 1010198 | 10158.705 | 0.031 | 0.995 | 78.13 | 78.0724 | -0.07
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.48 | 49368.879 | 1625311 | 10623.775 | 0.046 | 0.995 | 125 | 118.5485 | -5.16
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.48 | 47129.316 | 1509683 | 10158.639 | 0.046 | 0.995 | 125 | 118.3505 | -5.32
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.48 | 67692.227 | 2185550 | 10923.09 | 0.062 | 0.995 | 156.25 | 158.6059 | 1.51
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.48 | 63054.848 | 2014088 | 9865.15 | 0.064 | 0.995 | 156.25 | 163.6495 | 4.74
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.48 | 97595.797 | 3123183 | 9730.678 | 0.1 | 0.995 | 250 | 258.7078 | 3.48
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.48 | 92332.867 | 2901464 | 9832.22 | 0.094 | 0.995 | 250 | 241.9142 | -3.23
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 3.6 | 15.792 | 173 |  |  | 0.995
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 3.53 | 5.236 | 227 |  |  | 0.995
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 3.53 | 1.747 | 81 |  |  | 0.995
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 3.55 | 1.865 | 73 | 3.524 | 0.005 | 0.995 |  | 13.3019
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.48 | 19.907 | 704 | 10173.215 | 0 | 0.995
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.48 | 33.321 | 1070 | 11173.664 | 0 | 0.995
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.48 | 32.465 | 1024 | 12048.315 | 0 | 0.995
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.48 | 28.652 | 771 | 11975.843 | 0 | 0.995
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.48 | 24.97 | 652 | 10773.584 | 0 | 0.995
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 3.48 | 26.431 | 815 | 10459.238 | 0 | 0.995
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.48 | 32.773 | 1001 | 10166.878 | 0 | 0.995
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.48 | 24.036 | 656 | 9387.197 | 0 | 0.995
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.48 | 20.635 | 411 | 4318.83 | 0 | 0.995 |  | 0.01571
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.48 | 234.138 | 7814 | 10475.948 | 0 | 0.995 | 0.63 | 0.46045 | -26.91
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.48 | 226.375 | 7665 | 10532.542 | 0 | 0.995 | 0.63 | 0.43876 | -30.36
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.48 | 1541.29 | 52945 | 11635.099 | 0.001 | 0.995 | 2.5 | 3.2481 | 29.93
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.48 | 824.583 | 27864 | 6178.086 | 0.001 | 0.995 | 2.5 | 3.2735 | 30.94
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.48 | 2821.402 | 93332 | 12150.41 | 0.002 | 0.995 | 6.25 | 5.774 | -7.62
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.48 | 2458.629 | 82647 | 10931.824 | 0.002 | 0.995 | 6.25 | 5.5891 | -10.57
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.48 | 10033.756 | 330538 | 10343.043 | 0.01 | 0.995 | 25 | 24.491 | -2.04
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.48 | 9092.92 | 297083 | 9113.911 | 0.01 | 0.995 | 25 | 25.1921 | 0.77
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 3.48 | 0.826 | 51 |  |  | 0.995
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 3.54 | 9.21 | 111 |  |  | 0.995
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 3.4 | 2.286 | 77 |  |  | 0.995
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 3.49 | 2.49 | 101 |  |  | 0.995
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 3.48 | 20.559 | 696 |  |  | 0.995
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 3.48 | 2.372 | 126 |  |  | 0.995
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 3.43 | 10.672 | 237 |  |  | 0.995
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 3.48 | 19.075 | 513 |  |  | 0.995
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 3.49 | 5.293 | 146 |  |  | 0.995
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 3.48 | 40.174 | 694 |  |  | 0.995
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 3.44 | 5.467 | 68 |  |  | 0.995
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 3.57 | 2.565 | 101 |  |  | 0.995
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 3.53 | 2.172 | 67 |  |  | 0.995
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | 3.53 | 10.459 | 145 |  |  | 0.995
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 3.55 | 1.004 | 55 |  |  | 0.995
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.48 | 3711.082 | 121322 | 11440.534 | 0.003 | 0.995 |  | 8.1092
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.48 | 4811.606 | 159244 | 13256.198 | 0.004 | 0.995 |  | 9.0871
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 3.48 | 3409.649 | 110851 | 11719.4 | 0.003 | 0.995 |  | 7.2619
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.48 | 40.495 | 1302 | 11316.626 | 0 | 0.995
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.48 | 29.886 | 940 | 11121.704 | 0 | 0.995
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.48 | 33.682 | 1221 | 11384.287 | 0 | 0.995
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.48 | 24.934 | 939 | 11676.331 | 0 | 0.995
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.48 | 26.242 | 767 | 11894.549 | 0 | 0.995
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.48 | 24.269 | 516 | 11048.28 | 0 | 0.995
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.48 | 43379.383 | 1400961 | 9862.112 | 0.044 | 0.995 |  | 112.1514
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.48 | 43472.313 | 1406561 | 9161.524 | 0.047 | 0.995 |  | 121.0757
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.48 | 42736.262 | 1384160 | 9049.197 | 0.047 | 0.995 |  | 120.4974
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.48 | 77.155 | 1940 | 8897.753 | 0 | 0.995 |  | 0.11425
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.48 | 73.03 | 2627 | 10825.349 | 0 | 0.995 |  | 0.06553
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.48 | 58.565 | 1986 | 9994.433 | 0 | 0.995 |  | 0.04309
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.48 | 26.115 | 857 | 8190.83 | 0 | 0.995
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.48 | 23.96 | 691 | 9179.411 | 0 | 0.995
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.48 | 25.058 | 756 | 10046.605 | 0 | 0.995
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.48 | 50968.484 | 1646832 | 11574.24 | 0.044 | 0.995 |  | 112.2807
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.48 | 41411.648 | 1337742 | 8454.886 | 0.049 | 0.995 |  | 125.0164
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.48 | 44135.258 | 1418305 | 8997.701 | 0.049 | 0.995 |  | 125.2025
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.48 | 81.135 | 2930 | 9464.119 | 0 | 0.995 |  | 0.11176
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.48 | 50.627 | 1825 | 8344.308 | 0 | 0.995 |  | 0.04834
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.48 | 52.271 | 1924 | 8994.774 | 0 | 0.995 |  | 0.04186
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.48 | 22.398 | 811 | 9773.682 | 0 | 0.995
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.48 | 30.817 | 960 | 9471.696 | 0 | 0.995
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.48 | 24.207 | 564 | 9756.347 | 0 | 0.995
M2PFTeDA
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9730.571 | 329079 | 9730.571 |  |  | 0.01 | 0.00936 | -6.38
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11255 | 376966 | 11255 |  |  | 0.01 | 0.01083 | 8.29
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 3.48 | 11228.034 | 376645 | 11228.034 |  |  | 0.01 | 0.0108 | 8.03
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 3.48 | 10426.968 | 352207 | 10426.968 |  |  | 0.01 | 0.01003 | 0.32
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9798.542 | 327493 | 9798.542 |  |  | 0.01 | 0.00943 | -5.72
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 3.48 | 9861.838 | 334756 | 9861.838 |  |  | 0.01 | 0.00949 | -5.11
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 3.48 | 8900.187 | 298634 | 8900.187 |  |  | 0.01 | 0.00856 | -14.37
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11120.537 | 373820 | 11120.537 |  |  | 0.01 | 0.0107 | 7
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 3.48 | 12190.335 | 409431 | 12190.335 |  |  | 0.01 | 0.01173 | 17.29
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11829.166 | 397464 | 11829.166 |  |  | 0.01 | 0.01138 | 13.82
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 3.48 | 11531.981 | 391568 | 11531.981 |  |  | 0.01 | 0.0111 | 10.96
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10690.253 | 359731 | 10690.253 |  |  | 0.01 | 0.01029 | 2.86
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 3.48 | 10913.814 | 370373 | 10913.814 |  |  | 0.01 | 0.0105 | 5.01
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 3.48 | 9899.175 | 337103 | 9899.175 |  |  | 0.01 | 0.00952 | -4.75
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10876.152 | 369500 | 10876.152 |  |  | 0.01 | 0.01046 | 4.65
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11718.064 | 392785 | 11718.064 |  |  | 0.01 | 0.01127 | 12.75
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 3.48 | 11542.465 | 386368 | 11542.465 |  |  | 0.01 | 0.01111 | 11.06
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10908.972 | 365733 | 10908.972 |  |  | 0.01 | 0.0105 | 4.96
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 3.48 | 10546.275 | 355639 | 10546.275 |  |  | 0.01 | 0.01015 | 1.47
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9692.394 | 329243 | 9692.394 |  |  | 0.01 | 0.00933 | -6.74
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 3.48 | 9065.441 | 305364 | 9065.441 |  |  | 0.01 | 0.00872 | -12.78
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10360.407 | 348348 | 10360.407 |  |  | 0.01 | 0.00997 | -0.32
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11317.867 | 377809 | 11317.867 |  |  | 0.01 | 0.01089 | 8.9
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 3.48 | 11854.473 | 394417 | 11854.473 |  |  | 0.01 | 0.01141 | 14.06
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10369.82 | 347172 | 10369.82 |  |  | 0.01 | 0.00998 | -0.23
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 3.48 | 10312.553 | 346986 | 10312.553 |  |  | 0.01 | 0.00992 | -0.78
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9741.055 | 331371 | 9741.055 |  |  | 0.01 | 0.00937 | -6.28
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 3.48 | 9599.27 | 327251 | 9599.27 |  |  | 0.01 | 0.00924 | -7.64
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10282.522 | 349144 | 10282.522 |  |  | 0.01 | 0.00989 | -1.07
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10528.761 | 357528 | 10528.761 |  |  | 0.01 | 0.01013 | 1.3
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 3.48 | 10775.113 | 363685 | 10775.113 |  |  | 0.01 | 0.01037 | 3.67
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9679.873 | 327049 | 9679.873 |  |  | 0.01 | 0.00931 | -6.86
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 3.48 | 9162.68 | 310580 | 9162.68 |  |  | 0.01 | 0.00882 | -11.84
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8925.449 | 297809 | 8925.449 |  |  | 0.01 | 0.00859 | -14.12
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 3.48 | 8059.229 | 274308 | 8059.229 |  |  | 0.01 | 0.00775 | -22.46
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10162.122 | 345524 | 10162.122 |  |  | 0.01 | 0.00978 | -2.22
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 3.48 | 11242.785 | 374717 | 11242.785 |  |  | 0.01 | 0.01082 | 8.17
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10874.418 | 363355 | 10874.418 |  |  | 0.01 | 0.01046 | 4.63
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 3.48 | 10265.066 | 348695 | 10265.066 |  |  | 0.01 | 0.00988 | -1.23
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9531.285 | 324174 | 9531.285 |  |  | 0.01 | 0.00917 | -8.29
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 3.48 | 9877.485 | 335723 | 9877.485 |  |  | 0.01 | 0.0095 | -4.96
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 3.48 | 8519.921 | 288423 | 8519.921 |  |  | 0.01 | 0.0082 | -18.02
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 3.48 | 10490.147 | 354799 | 10490.147 |  |  | 0.01 | 0.01009 | 0.93
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 3.48 | 11516.313 | 392960 | 11516.313 |  |  | 0.01 | 0.01108 | 10.81
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9708.633 | 329742 | 9708.633 |  |  | 0.01 | 0.00934 | -6.59
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 3.48 | 9848.446 | 330014 | 9848.446 |  |  | 0.01 | 0.00948 | -5.24
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 3.48 | 9783.861 | 332235 | 9783.861 |  |  | 0.01 | 0.00941 | -5.86
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 3.48 | 10084.969 | 342025 | 10084.969 |  |  | 0.01 | 0.0097 | -2.97
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 3.48 | 11453.623 | 387651 | 11453.623 |  |  | 0.01 | 0.01102 | 10.2
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 3.48 | 10773.109 | 363168 | 10773.109 |  |  | 0.01 | 0.01037 | 3.65
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 3.48 | 10865.012 | 367049 | 10865.012 |  |  | 0.01 | 0.01045 | 4.54
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 3.48 | 11209.78 | 378373 | 11209.78 |  |  | 0.01 | 0.01079 | 7.86
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 3.48 | 11101.751 | 374089 | 11101.751 |  |  | 0.01 | 0.01068 | 6.82
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 3.48 | 10679.77 | 353296 | 10679.77 |  |  | 0.01 | 0.01028 | 2.76
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10508.111 | 351061 | 10508.111 |  |  | 0.01 | 0.01011 | 1.1
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 3.48 | 10697.722 | 357232 | 10697.722 |  |  | 0.01 | 0.01029 | 2.93
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 3.48 | 9989.264 | 334986 | 9989.264 |  |  | 0.01 | 0.00961 | -3.89
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 3.48 | 10158.705 | 334343 | 10158.705 |  |  | 0.01 | 0.00977 | -2.26
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10623.775 | 354798 | 10623.775 |  |  | 0.01 | 0.01022 | 2.22
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 3.48 | 10158.639 | 338366 | 10158.639 |  |  | 0.01 | 0.00977 | -2.26
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 3.48 | 10923.09 | 361685 | 10923.09 |  |  | 0.01 | 0.01051 | 5.1
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 3.48 | 9865.15 | 326284 | 9865.15 |  |  | 0.01 | 0.00949 | -5.08
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9730.678 | 316707 | 9730.678 |  |  | 0.01 | 0.00936 | -6.38
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 3.48 | 9832.22 | 321241 | 9832.22 |  |  | 0.01 | 0.00946 | -5.4
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 3.49 | 3.524 | 114 | 3.524 |  |  | 0.01 | 0 | -99.97
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 3.48 | 10173.215 | 343416 | 10173.215 |  |  | 0.01 | 0.00979 | -2.12
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 3.48 | 11173.664 | 376859 | 11173.664 |  |  | 0.01 | 0.01075 | 7.51
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 3.48 | 12048.315 | 404321 | 12048.315 |  |  | 0.01 | 0.01159 | 15.92
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 3.48 | 11975.843 | 401621 | 11975.843 |  |  | 0.01 | 0.01152 | 15.23
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 3.48 | 10773.584 | 362156 | 10773.584 |  |  | 0.01 | 0.01037 | 3.66
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 3.48 | 10459.238 | 354868 | 10459.238 |  |  | 0.01 | 0.01006 | 0.63
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 3.48 | 10166.878 | 344723 | 10166.878 |  |  | 0.01 | 0.00978 | -2.18
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 3.48 | 9387.197 | 317571 | 9387.197 |  |  | 0.01 | 0.00903 | -9.68
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 3.48 | 4318.83 | 148626 | 4318.83 |  |  | 0.01 | 0.00416 | -58.45
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 3.48 | 10475.948 | 355039 | 10475.948 |  |  | 0.01 | 0.01008 | 0.8
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 3.48 | 10532.542 | 355947 | 10532.542 |  |  | 0.01 | 0.01013 | 1.34
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 3.48 | 11635.099 | 388325 | 11635.099 |  |  | 0.01 | 0.01119 | 11.95
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 3.48 | 6178.086 | 209790 | 6178.086 |  |  | 0.01 | 0.00594 | -40.56
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 3.48 | 12150.41 | 410208 | 12150.41 |  |  | 0.01 | 0.01169 | 16.91
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 3.48 | 10931.824 | 367413 | 10931.824 |  |  | 0.01 | 0.01052 | 5.18
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 3.48 | 10343.043 | 345742 | 10343.043 |  |  | 0.01 | 0.00995 | -0.48
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 3.48 | 9113.911 | 304569 | 9113.911 |  |  | 0.01 | 0.00877 | -12.31
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 3.48 | 11440.534 | 382724 | 11440.534 |  |  | 0.01 | 0.01101 | 10.08
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 3.48 | 13256.198 | 439493 | 13256.198 |  |  | 0.01 | 0.01275 | 27.55
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 3.48 | 11719.4 | 391810 | 11719.4 |  |  | 0.01 | 0.01128 | 12.76
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 3.48 | 11316.626 | 377070 | 11316.626 |  |  | 0.01 | 0.01089 | 8.88
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 3.48 | 11121.704 | 373796 | 11121.704 |  |  | 0.01 | 0.0107 | 7.01
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 3.48 | 11384.287 | 381386 | 11384.287 |  |  | 0.01 | 0.01095 | 9.53
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 3.48 | 11676.331 | 390146 | 11676.331 |  |  | 0.01 | 0.01123 | 12.34
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 3.48 | 11894.549 | 401063 | 11894.549 |  |  | 0.01 | 0.01144 | 14.44
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 3.48 | 11048.28 | 371884 | 11048.28 |  |  | 0.01 | 0.01063 | 6.3
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 3.48 | 9862.112 | 325453 | 9862.112 |  |  | 0.01 | 0.00949 | -5.11
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 3.48 | 9161.524 | 302491 | 9161.524 |  |  | 0.01 | 0.00881 | -11.85
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 3.48 | 9049.197 | 303067 | 9049.197 |  |  | 0.01 | 0.00871 | -12.93
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 3.48 | 8897.753 | 300237 | 8897.753 |  |  | 0.01 | 0.00856 | -14.39
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 3.48 | 10825.349 | 365981 | 10825.349 |  |  | 0.01 | 0.01042 | 4.16
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 3.48 | 9994.433 | 339009 | 9994.433 |  |  | 0.01 | 0.00962 | -3.84
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 3.48 | 8190.83 | 272286 | 8190.83 |  |  | 0.01 | 0.00788 | -21.19
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 3.48 | 9179.411 | 308907 | 9179.411 |  |  | 0.01 | 0.00883 | -11.68
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 3.48 | 10046.605 | 340119 | 10046.605 |  |  | 0.01 | 0.00967 | -3.34
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 3.48 | 11574.24 | 389337 | 11574.24 |  |  | 0.01 | 0.01114 | 11.36
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 3.48 | 8454.886 | 279574 | 8454.886 |  |  | 0.01 | 0.00813 | -18.65
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 3.48 | 8997.701 | 299688 | 8997.701 |  |  | 0.01 | 0.00866 | -13.43
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 3.48 | 9464.119 | 317767 | 9464.119 |  |  | 0.01 | 0.00911 | -8.94
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 3.48 | 8344.308 | 280834 | 8344.308 |  |  | 0.01 | 0.00803 | -19.71
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 3.48 | 8994.774 | 300266 | 8994.774 |  |  | 0.01 | 0.00865 | -13.46
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 3.48 | 9773.682 | 330647 | 9773.682 |  |  | 0.01 | 0.0094 | -5.96
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 3.48 | 9471.696 | 319932 | 9471.696 |  |  | 0.01 | 0.00911 | -8.87
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 3.48 | 9756.347 | 327921 | 9756.347 |  |  | 0.01 | 0.00939 | -6.13

## DTXSID3020209
*Analyte in Mix 1, 2, 3 | [8 empty] | n-Butylparaben
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Coeff. Of Determination | Std. Conc (nM) | Actual Conc (nM) | %Dev
Fraction Unbound Determination | [8 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.49 | 275.932 | 7944 | 19211.244 | 0 | 0.977 | 0.17
Sample Text | Actual Conc (nM) | Corrected Conc (nM) | Fu | Avg Fu | Std Dev Fu | CV |  |  | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.49 | 344.988 | 9416 | 19356.43 | 0 | 0.977 | 0.17
UC_AF_Mix1_A | 4.6712 | 149.48 | 0.019927 | 0.020912 | 0.00077163 | 3.69% |  |  | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.49 | 255.55 | 6633 | 19171.268 | 0 | 0.977 | 0.17
UC_AF_Mix1_B | 5.1616 | 165.1712 | 0.020097 | [5 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.49 | 299.582 | 7591 | 20027.883 | 0 | 0.977 | 0.17
UC_AF_Mix1_C | 5.3156 | 170.0992 | 0.02079 | [5 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.49 | 324.78 | 8199 | 21216.219 | 0 | 0.977 | 0.17
UC_AF_Mix2_A | 5.4438 | 174.2019 | 0.020585 | [5 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.49 | 361.179 | 9824 | 20622.34 | 0 | 0.977 | 0.17
UC_AF_Mix2_B | 5.6609 | 181.1498 | 0.0211 | [5 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.49 | 358.748 | 10079 | 20838.236 | 0 | 0.977 | 0.17
UC_AF_Mix2_C | 5.7687 | 184.5968 | 0.021415 | [5 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1093.283 | 29519 | 19452.498 | 0.001 | 0.977 | 0.28 | 0.5328 | 90.29
UC_AF_Mix3_A | 4.9908 | 159.7056 | 0.020303 | [5 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1084.972 | 29756 | 20514.73 | 0.001 | 0.977 | 0.28 | 0.46425 | 65.8
UC_AF_Mix3_B | 5.5341 | 177.0912 | 0.022037 | [5 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1221.837 | 33205 | 21051.086 | 0.001 | 0.977 | 0.28 | 0.57082 | 103.87
UC_AF_Mix3_C | 5.4886 | 175.6349 | 0.021951 | [5 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1179.964 | 32207 | 21530.92 | 0.001 | 0.977 | 0.28 | 0.50386 | 79.95
UC_T1hr_Mix1_A | 115.3903 | 9231.224 | [6 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1213.55 | 32660 | 21769.814 | 0.001 | 0.977 | 0.28 | 0.52333 | 86.9
UC_T1hr_Mix1_B | 124.424 | 9953.92 | [6 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1313.061 | 35458 | 22279.238 | 0.001 | 0.977 | 0.28 | 0.58933 | 110.47
UC_T1hr_Mix1_C | 121.0162 | 9681.2968 | [6 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.49 | 1252.409 | 34936 | 22277.256 | 0.001 | 0.977 | 0.28 | 0.53314 | 90.41
UC_T1hr_Mix2_A | 132.7843 | 10622.7424 | [6 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2038.82 | 54724 | 19150.906 | 0.001 | 0.977 | 0.44 | 1.572 | 257.28
UC_T1hr_Mix2_B | 132.4023 | 10592.184 | [6 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2111.016 | 57924 | 20179.088 | 0.001 | 0.977 | 0.44 | 1.5338 | 248.6
UC_T1hr_Mix2_C | 135.2957 | 10823.6528 | [6 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2038.302 | 54126 | 20352.199 | 0.001 | 0.977 | 0.44 | 1.4416 | 227.63
UC_T1hr_Mix3_A | 123.297 | 9863.7608 | [6 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2052.207 | 59116 | 21003.221 | 0.001 | 0.977 | 0.44 | 1.3911 | 216.15
UC_T1hr_Mix3_B | 121.7901 | 9743.2104 | [6 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2281.747 | 61964 | 21732.258 | 0.001 | 0.977 | 0.44 | 1.5417 | 250.38
UC_T1hr_Mix3_C | 126.3165 | 10105.3192 | [6 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2139.417 | 57603 | 22107.174 | 0.001 | 0.977 | 0.44 | 1.3718 | 211.76
UC_T5hr_Mix1_A | 93.7681 | 7501.4456 | [6 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.49 | 2225.525 | 59572 | 22667.871 | 0.001 | 0.977 | 0.44 | 1.4008 | 218.36
UC_T5hr_Mix1_B | 102.7359 | 8218.8728 | [6 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1336.353 | 36085 | 19353 | 0.001 | 0.977 | 0.71 | 0.79849 | 12.46
UC_T5hr_Mix1_C | 102.2725 | 8181.8008 | [6 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1404.504 | 38192 | 20475.318 | 0.001 | 0.977 | 0.71 | 0.78905 | 11.13
UC_T5hr_Mix2_A | 105.7811 | 8462.492 | [6 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1503.632 | 39949 | 19914.854 | 0.001 | 0.977 | 0.71 | 0.93189 | 31.25
UC_T5hr_Mix2_B | 107.3146 | 8585.1672 | [6 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1596.819 | 42966 | 21356.199 | 0.001 | 0.977 | 0.71 | 0.91675 | 29.12
UC_T5hr_Mix2_C | 107.7522 | 8620.1768 | [6 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1593.851 | 44044 | 21173.01 | 0.001 | 0.977 | 0.71 | 0.92723 | 30.6
UC_T5hr_Mix3_A | 98.3282 | 7866.2552 | [6 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1595.844 | 43525 | 22507.645 | 0.001 | 0.977 | 0.71 | 0.83676 | 17.85
UC_T5hr_Mix3_B | 100.4497 | 8035.9744 | [6 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.49 | 1545.231 | 41500 | 21603.74 | 0.001 | 0.977 | 0.71 | 0.84966 | 19.67
UC_T5hr_Mix3_C | 100.0133 | 8001.064 | [6 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1132.776 | 30425 | 18654.168 | 0.001 | 0.977 | 1.14 | 0.62631 | -45.06
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1214.246 | 32244 | 19498.076 | 0.001 | 0.977 | 1.14 | 0.65837 | -42.25
MDL Determination | [8 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1165.913 | 32669 | 19468.383 | 0.001 | 0.977 | 1.14 | 0.60899 | -46.58
Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1243.231 | 34288 | 20342.879 | 0.001 | 0.977 | 1.14 | 0.63435 | -44.36
20201123_PFAS_PPB_UC_Sample012 | CC6 | 1.82 | 1.4423 | 1.3794 | 0.042396 | 3.074% | 0.13325 | -24.21% | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1309.173 | 35064 | 21096.451 | 0.001 | 0.977 | 1.14 | 0.65385 | -42.65
20201123_PFAS_PPB_UC_Sample037 | CC6 | 1.82 | 1.3659 | [5 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1263.111 | 35567 | 21428.309 | 0.001 | 0.977 | 1.14 | 0.58952 | -48.29
20201123_PFAS_PPB_UC_Sample044 | CC6 | 1.82 | 1.4255 | [5 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.49 | 1230.308 | 33452 | 21934.615 | 0.001 | 0.977 | 1.14 | 0.53046 | -53.47
20201123_PFAS_PPB_UC_Sample084 | CC6 | 1.82 | 1.33 | [5 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1858.92 | 49326 | 18554.959 | 0.001 | 0.977 | 1.82 | 1.4423 | -20.76
20201123_PFAS_PPB_UC_Sample110 | CC6 | 1.82 | 1.3927 | [5 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1892.037 | 49662 | 19607.979 | 0.001 | 0.977 | 1.82 | 1.3659 | -24.95
20201123_PFAS_PPB_UC_Sample117 | CC6 | 1.82 | 1.3575 | [5 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1999.524 | 53469 | 20120.816 | 0.001 | 0.977 | 1.82 | 1.4255 | -21.67
20201123_PFAS_PPB_UC_Sample124 | CC6 | 1.82 | 1.3417 | [5 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.49 | 1923.861 | 51393 | 20303.764 | 0.001 | 0.977 | 1.82 | 1.33 | -26.93
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2031.443 | 55841 | 20774.008 | 0.001 | 0.977 | 1.82 | 1.3927 | -23.48
Blanks should be less than… | 0.066626 | nM | [6 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2034.314 | 53494 | 21171.861 | 0.001 | 0.977 | 1.82 | 1.3575 | -25.41
eLOQ | 1.82 | nM | [6 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.49 | 2038.945 | 56389 | 21390.625 | 0.001 | 0.977 | 1.82 | 1.3417 | -26.28
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.49 | 3011.271 | 82197 | 19629.662 | 0.002 | 0.977 | 2.91 | 2.5427 | -12.62
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.49 | 3185.441 | 87389 | 20294.021 | 0.002 | 0.977 | 2.91 | 2.6163 | -10.09
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.49 | 5437.346 | 144185 | 18885.018 | 0.003 | 0.977 | 4.66 | 5.324 | 14.25
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.49 | 5724.029 | 158430 | 19831.742 | 0.003 | 0.977 | 4.66 | 5.3387 | 14.56
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.49 | 6844.65 | 186400 | 18374.332 | 0.004 | 0.977 | 7.45 | 7.073 | -5.06
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.49 | 7168.389 | 194432 | 20010.529 | 0.004 | 0.977 | 7.45 | 6.7777 | -9.02
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.49 | 11706.379 | 317014 | 19000.076 | 0.006 | 0.977 | 11.92 | 12.1094 | 1.59
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.49 | 12422.017 | 335605 | 20213.203 | 0.006 | 0.977 | 11.92 | 12.0769 | 1.32
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.49 | 23140.768 | 624188 | 20171.609 | 0.011 | 0.977 | 19.07 | 23.0868 | 21.06
QC FLAGS | [8 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.49 | 25001.709 | 672642 | 21640.061 | 0.012 | 0.977 | 19.07 | 23.2552 | 21.95
Calibration Curve Linearity |  | <7-point curve min and/or r2 < 0.98 | [6 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.49 | 30056.643 | 810873 | 19826.656 | 0.015 | 0.977 | 30.52 | 30.7079 | 0.62
Curve Precision Checks |  | <75 or >125% of expected | [6 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.49 | 31641.816 | 849207 | 21196.979 | 0.015 | 0.977 | 30.52 | 30.2281 | -0.96
Blanks |  | >1/2 MDL of analytes | [6 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.49 | 46749.152 | 1276742 | 19678.5 | 0.024 | 0.977 | 48.83 | 48.4689 | -0.74
Technical Replicate Assessment |  | Replicates > 25% of Sample Mean | [6 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.49 | 50444.727 | 1365900 | 20941.443 | 0.024 | 0.977 | 48.83 | 49.1547 | 0.66
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.49 | 73348.117 | 1995609 | 18895.691 | 0.039 | 0.977 | 78.13 | 79.568 | 1.84
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.49 | 77764 | 2124020 | 20097.592 | 0.039 | 0.977 | 78.13 | 79.3116 | 1.51
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.49 | 107727.188 | 2899075 | 18753.004 | 0.057 | 0.977 | 125 | 118.0027 | -5.6
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.49 | 116241.914 | 3148214 | 20623.277 | 0.056 | 0.977 | 125 | 115.7735 | -7.38
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.49 | 141221.516 | 3782837 | 19034.314 | 0.074 | 0.977 | 156.25 | 152.5302 | -2.38
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.49 | 150393.859 | 4087704 | 19677.9 | 0.076 | 0.977 | 156.25 | 157.1352 | 0.57
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.49 | 233946.781 | 6256998 | 19099.988 | 0.122 | 0.977 | 250 | 251.9417 | 0.78
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.49 | 253237.719 | 6808499 | 20306.045 | 0.125 | 0.977 | 250 | 256.5171 | 2.61
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | 2.49 | 12.328 | 184 |  |  | 0.977
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | 2.49 | 11.044 | 290 |  |  | 0.977
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | 2.47 | 11.231 | 317 |  |  | 0.977
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | 2.46 | 9.078 | 178 |  |  | 0.977
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.49 | 86.892 | 2244 | 18857.293 | 0 | 0.977
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.49 | 90.725 | 2272 | 19474.645 | 0 | 0.977
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.49 | 120.839 | 3122 | 19650.477 | 0 | 0.977
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.49 | 121.111 | 3289 | 20304.375 | 0 | 0.977
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.49 | 140.072 | 3686 | 20612.568 | 0 | 0.977
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.5 | 111.485 | 2843 | 20241.559 | 0 | 0.977
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.49 | 125.845 | 2872 | 21380.4 | 0 | 0.977
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.49 | 113.136 | 3323 | 21765.705 | 0 | 0.977
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.49 | 55.68 | 1125 | 7133.3 | 0 | 0.977
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.49 | 869.421 | 23752 | 19263.697 | 0 | 0.977 | 0.63 | 0.30389 | -51.76
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.49 | 885.964 | 23830 | 20563.285 | 0 | 0.977 | 0.63 | 0.26154 | -58.49
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.49 | 2826.188 | 75980 | 20758.027 | 0.001 | 0.977 | 2.5 | 2.1859 | -12.56
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.49 | 3360.3 | 91191 | 24810.557 | 0.001 | 0.977 | 2.5 | 2.1712 | -13.15
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.49 | 6402.18 | 171699 | 20168.148 | 0.003 | 0.977 | 6.25 | 5.9344 | -5.05
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.49 | 6762.778 | 184250 | 21870.688 | 0.003 | 0.977 | 6.25 | 5.7644 | -7.77
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.49 | 25493.525 | 686891 | 19404.105 | 0.013 | 0.977 | 25 | 26.5305 | 6.12
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.49 | 27081.482 | 739194 | 21461.814 | 0.013 | 0.977 | 25 | 25.4563 | 1.83
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | 2.52 | 1.753 | 72 |  |  | 0.977
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | 2.48 | 5.68 | 153 |  |  | 0.977
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | 2.46 | 1.126 | 41 |  |  | 0.977
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | 2.47 | 2.664 | 90 |  |  | 0.977
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | 2.47 | 50.034 | 746 |  |  | 0.977
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | 2.46 | 9.642 | 158 |  |  | 0.977
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | 2.6 | 4.744 | 112 |  |  | 0.977
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | 2.47 | 11.253 | 324 |  |  | 0.977
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | 2.47 | 7.267 | 249 |  |  | 0.977
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | 2.5 | 1.847 | 91 |  |  | 0.977
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | 2.52 | 2.993 | 128 |  |  | 0.977
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | 2.5 | 8.466 | 186 |  |  | 0.977
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | 2.47 | 1.671 | 69 |  |  | 0.977
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [5 empty] | 0.977
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | 2.43 | 3.952 | 157 |  |  | 0.977
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.49 | 5562.143 | 151564 | 21697.566 | 0.003 | 0.977 |  | 4.6712
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.49 | 6113.616 | 167782 | 21829.346 | 0.003 | 0.977 |  | 5.1616
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.49 | 5989.404 | 162539 | 20831.814 | 0.003 | 0.977 |  | 5.3156
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.49 | 6199.086 | 166477 | 21105.879 | 0.003 | 0.977 |  | 5.4438
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.49 | 6240.639 | 170442 | 20513.939 | 0.003 | 0.977 |  | 5.6609
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.49 | 6418.389 | 173326 | 20742.979 | 0.003 | 0.977 |  | 5.7687
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.49 | 5727.25 | 153109 | 21071.273 | 0.003 | 0.977 |  | 4.9908
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.49 | 6183.206 | 167058 | 20743.402 | 0.003 | 0.977 |  | 5.5341
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.49 | 6032.018 | 162409 | 20386.736 | 0.003 | 0.977 |  | 5.4886
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.49 | 107377.172 | 2903668 | 19113.52 | 0.056 | 0.977 |  | 115.3903
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.49 | 115516.563 | 3114122 | 19075.039 | 0.061 | 0.977 |  | 124.424
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.49 | 115451.219 | 3119983 | 19599.061 | 0.059 | 0.977 |  | 121.0162
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.49 | 120116.875 | 3285115 | 18590.02 | 0.065 | 0.977 |  | 132.7843
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.49 | 130066.008 | 3487232 | 20187.695 | 0.064 | 0.977 |  | 132.4023
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.49 | 130411.273 | 3512177 | 19809.779 | 0.066 | 0.977 |  | 135.2957
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.49 | 114309.352 | 3075295 | 19047.59 | 0.06 | 0.977 |  | 123.297
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.49 | 116265.188 | 3146169 | 19612.295 | 0.059 | 0.977 |  | 121.7901
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.49 | 121336.031 | 3245035 | 19736.887 | 0.061 | 0.977 |  | 126.3165
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.49 | 104149.32 | 2813923 | 22790.889 | 0.046 | 0.977 |  | 93.7681
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.49 | 100502.93 | 2698115 | 20082.863 | 0.05 | 0.977 |  | 102.7359
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.49 | 100130.891 | 2728622 | 20098.73 | 0.05 | 0.977 |  | 102.2725
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.49 | 101750.938 | 2724955 | 19749.709 | 0.052 | 0.977 |  | 105.7811
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.49 | 104112.445 | 2802601 | 19920.66 | 0.052 | 0.977 |  | 107.3146
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.49 | 101660.32 | 2720809 | 19372.84 | 0.052 | 0.977 |  | 107.7522
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.49 | 99647.773 | 2684259 | 20799.916 | 0.048 | 0.977 |  | 98.3282
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.49 | 98506.859 | 2675978 | 20129.732 | 0.049 | 0.977 |  | 100.4497
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.49 | 98968.75 | 2698679 | 20311.908 | 0.049 | 0.977 |  | 100.0133
13C6 n-Butylparaben
[9 empty] | # | Name | Sample Text | Type | RT | Area | Height | Response |  |  | Std. Conc (µM) | Actual Conc (µM) | %Dev
[9 empty] | 7 | 20201123_PFAS_PPB_UC_Sample007 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19211.244 | 512622 | 19211.244 |  |  | 0.01
[9 empty] | 32 | 20201123_PFAS_PPB_UC_Sample032 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19356.43 | 523968 | 19356.43 |  |  | 0.01
[9 empty] | 39 | 20201123_PFAS_PPB_UC_Sample039 | CC1 - 0.053 pg/uL | Standard | 2.49 | 19171.268 | 515644 | 19171.268 |  |  | 0.01
[9 empty] | 79 | 20201123_PFAS_PPB_UC_Sample079 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20027.883 | 552934 | 20027.883 |  |  | 0.01
[9 empty] | 105 | 20201123_PFAS_PPB_UC_Sample105 | CC1 - 0.053 pg/uL | Standard | 2.49 | 21216.219 | 576033 | 21216.219 |  |  | 0.01 | 0.00926 | -7.41
[9 empty] | 112 | 20201123_PFAS_PPB_UC_Sample112 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20622.34 | 561597 | 20622.34 |  |  | 0.01
[9 empty] | 119 | 20201123_PFAS_PPB_UC_Sample119 | CC1 - 0.053 pg/uL | Standard | 2.49 | 20838.236 | 563891 | 20838.236 |  |  | 0.01 | 0.00943 | -5.67
[9 empty] | 8 | 20201123_PFAS_PPB_UC_Sample008 | CC2 - 0.085 pg/uL | Standard | 2.49 | 19452.498 | 523086 | 19452.498 |  |  | 0.01 | 0.00935 | -6.54
[9 empty] | 33 | 20201123_PFAS_PPB_UC_Sample033 | CC2 - 0.085 pg/uL | Standard | 2.49 | 20514.73 | 557477 | 20514.73 |  |  | 0.01 | 0.00938 | -6.22
[9 empty] | 40 | 20201123_PFAS_PPB_UC_Sample040 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21051.086 | 566469 | 21051.086 |  |  | 0.01
[9 empty] | 80 | 20201123_PFAS_PPB_UC_Sample080 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21530.92 | 588493 | 21530.92 |  |  | 0.01 | 0.00946 | -5.41
[9 empty] | 106 | 20201123_PFAS_PPB_UC_Sample106 | CC2 - 0.085 pg/uL | Standard | 2.49 | 21769.814 | 595623 | 21769.814 |  |  | 0.01 | 0.01019 | 1.92
[9 empty] | 113 | 20201123_PFAS_PPB_UC_Sample113 | CC2 - 0.085 pg/uL | Standard | 2.49 | 22279.238 | 606721 | 22279.238 |  |  | 0.01 | 0.0099 | -0.97
[9 empty] | 120 | 20201123_PFAS_PPB_UC_Sample120 | CC2 - 0.085 pg/uL | Standard | 2.49 | 22277.256 | 608110 | 22277.256 |  |  | 0.01 | 0.00953 | -4.72
[9 empty] | 9 | 20201123_PFAS_PPB_UC_Sample009 | CC3 - 0.135 pg/uL | Standard | 2.49 | 19150.906 | 516256 | 19150.906 |  |  | 0.01
[9 empty] | 34 | 20201123_PFAS_PPB_UC_Sample034 | CC3 - 0.135 pg/uL | Standard | 2.49 | 20179.088 | 541603 | 20179.088 |  |  | 0.01 | 0.0095 | -4.96
[9 empty] | 41 | 20201123_PFAS_PPB_UC_Sample041 | CC3 - 0.135 pg/uL | Standard | 2.49 | 20352.199 | 553407 | 20352.199 |  |  | 0.01 | 0.01007 | 0.73
[9 empty] | 81 | 20201123_PFAS_PPB_UC_Sample081 | CC3 - 0.135 pg/uL | Standard | 2.49 | 21003.221 | 577334 | 21003.221 |  |  | 0.01 | 0.00991 | -0.92
[9 empty] | 107 | 20201123_PFAS_PPB_UC_Sample107 | CC3 - 0.135 pg/uL | Standard | 2.49 | 21732.258 | 586430 | 21732.258 |  |  | 0.01 | 0.01005 | 0.54
[9 empty] | 114 | 20201123_PFAS_PPB_UC_Sample114 | CC3 - 0.135 pg/uL | Standard | 2.49 | 22107.174 | 605748 | 22107.174 |  |  | 0.01 | 0.00957 | -4.26
[9 empty] | 121 | 20201123_PFAS_PPB_UC_Sample121 | CC3 - 0.135 pg/uL | Standard | 2.49 | 22667.871 | 620313 | 22667.871 |  |  | 0.01 | 0.00963 | -3.72
[9 empty] | 10 | 20201123_PFAS_PPB_UC_Sample010 | CC4 - 0.217 pg/uL | Standard | 2.49 | 19353 | 525200 | 19353 |  |  | 0.01 | 0.00965 | -3.51
[9 empty] | 35 | 20201123_PFAS_PPB_UC_Sample035 | CC4 - 0.217 pg/uL | Standard | 2.49 | 20475.318 | 551582 | 20475.318 |  |  | 0.01 | 0.00941 | -5.87
[9 empty] | 42 | 20201123_PFAS_PPB_UC_Sample042 | CC4 - 0.217 pg/uL | Standard | 2.49 | 19914.854 | 535562 | 19914.854 |  |  | 0.01 | 0.01034 | 3.36
[9 empty] | 82 | 20201123_PFAS_PPB_UC_Sample082 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21356.199 | 581177 | 21356.199 |  |  | 0.01 | 0.00999 | -0.07
[9 empty] | 108 | 20201123_PFAS_PPB_UC_Sample108 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21173.01 | 577698 | 21173.01 |  |  | 0.01 | 0.00978 | -2.22
[9 empty] | 115 | 20201123_PFAS_PPB_UC_Sample115 | CC4 - 0.217 pg/uL | Standard | 2.49 | 22507.645 | 615348 | 22507.645 |  |  | 0.01 | 0.00956 | -4.41
[9 empty] | 122 | 20201123_PFAS_PPB_UC_Sample122 | CC4 - 0.217 pg/uL | Standard | 2.49 | 21603.74 | 592242 | 21603.74 |  |  | 0.01 | 0.00988 | -1.2
[9 empty] | 11 | 20201123_PFAS_PPB_UC_Sample011 | CC5 - 0.347 pg/uL | Standard | 2.49 | 18654.168 | 506819 | 18654.168 |  |  | 0.01
[9 empty] | 36 | 20201123_PFAS_PPB_UC_Sample036 | CC5 - 0.347 pg/uL | Standard | 2.49 | 19498.076 | 523460 | 19498.076 |  |  | 0.01 | 0.01065 | 6.54
[9 empty] | 43 | 20201123_PFAS_PPB_UC_Sample043 | CC5 - 0.347 pg/uL | Standard | 2.49 | 19468.383 | 523251 | 19468.383 |  |  | 0.01 | 0.01072 | 7.19
[9 empty] | 83 | 20201123_PFAS_PPB_UC_Sample083 | CC5 - 0.347 pg/uL | Standard | 2.49 | 20342.879 | 549242 | 20342.879 |  |  | 0.01 | 0.01023 | 2.29
[9 empty] | 109 | 20201123_PFAS_PPB_UC_Sample109 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21096.451 | 577658 | 21096.451 |  |  | 0.01 | 0.01036 | 3.63
[9 empty] | 116 | 20201123_PFAS_PPB_UC_Sample116 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21428.309 | 584103 | 21428.309 |  |  | 0.01 | 0.01007 | 0.73
[9 empty] | 123 | 20201123_PFAS_PPB_UC_Sample123 | CC5 - 0.347 pg/uL | Standard | 2.49 | 21934.615 | 600256 | 21934.615 |  |  | 0.01 | 0.01019 | 1.85
[9 empty] | 12 | 20201123_PFAS_PPB_UC_Sample012 | CC6 - 0.555 pg/uL | Standard | 2.49 | 18554.959 | 509030 | 18554.959 |  |  | 0.01 | 0.00997 | -0.3
[9 empty] | 37 | 20201123_PFAS_PPB_UC_Sample037 | CC6 - 0.555 pg/uL | Standard | 2.49 | 19607.979 | 524201 | 19607.979 |  |  | 0.01 | 0.01035 | 3.46
[9 empty] | 44 | 20201123_PFAS_PPB_UC_Sample044 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20120.816 | 542770 | 20120.816 |  |  | 0.01 | 0.01019 | 1.85
[9 empty] | 84 | 20201123_PFAS_PPB_UC_Sample084 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20303.764 | 551555 | 20303.764 |  |  | 0.01 | 0.01001 | 0.1
[9 empty] | 110 | 20201123_PFAS_PPB_UC_Sample110 | CC6 - 0.555 pg/uL | Standard | 2.49 | 20774.008 | 565337 | 20774.008 |  |  | 0.01 | 0.00939 | -6.15
[9 empty] | 117 | 20201123_PFAS_PPB_UC_Sample117 | CC6 - 0.555 pg/uL | Standard | 2.49 | 21171.861 | 573511 | 21171.861 |  |  | 0.01 | 0.00937 | -6.34
[9 empty] | 124 | 20201123_PFAS_PPB_UC_Sample124 | CC6 - 0.555 pg/uL | Standard | 2.49 | 21390.625 | 584158 | 21390.625 |  |  | 0.01 | 0.00962 | -3.77
[9 empty] | 14 | 20201123_PFAS_PPB_UC_Sample014 | CC7 - 0.888 pg/uL | Standard | 2.49 | 19629.662 | 527429 | 19629.662 |  |  | 0.01
[9 empty] | 86 | 20201123_PFAS_PPB_UC_Sample086 | CC7 - 0.888 pg/uL | Standard | 2.49 | 20294.021 | 556258 | 20294.021 |  |  | 0.01 | 0.00913 | -8.72
[9 empty] | 15 | 20201123_PFAS_PPB_UC_Sample015 | CC8 - 1.420 pg/uL | Standard | 2.49 | 18885.018 | 512800 | 18885.018 |  |  | 0.01 | 0.00991 | -0.88
[9 empty] | 87 | 20201123_PFAS_PPB_UC_Sample087 | CC8 - 1.420 pg/uL | Standard | 2.49 | 19831.742 | 541739 | 19831.742 |  |  | 0.01 | 0.00973 | -2.73
[9 empty] | 16 | 20201123_PFAS_PPB_UC_Sample016 | CC9 - 2.272 pg/uL | Standard | 2.49 | 18374.332 | 490746 | 18374.332 |  |  | 0.01 | 0.00935 | -6.47
[9 empty] | 88 | 20201123_PFAS_PPB_UC_Sample088 | CC9 - 2.272 pg/uL | Standard | 2.49 | 20010.529 | 544568 | 20010.529 |  |  | 0.01 | 0.00963 | -3.7
[9 empty] | 17 | 20201123_PFAS_PPB_UC_Sample017 | CC10 - 3.636 pg/uL | Standard | 2.49 | 19000.076 | 514764 | 19000.076 |  |  | 0.01 | 0.00955 | -4.49
[9 empty] | 89 | 20201123_PFAS_PPB_UC_Sample089 | CC10 - 3.636 pg/uL | Standard | 2.49 | 20213.203 | 557415 | 20213.203 |  |  | 0.01 | 0.0094 | -5.97
[9 empty] | 18 | 20201123_PFAS_PPB_UC_Sample018 | CC11 - 5.817 pg/uL | Standard | 2.49 | 20171.609 | 542902 | 20171.609 |  |  | 0.01 | 0.0095 | -4.97
[9 empty] | 90 | 20201123_PFAS_PPB_UC_Sample090 | CC11 - 5.817 pg/uL | Standard | 2.49 | 21640.061 | 585142 | 21640.061 |  |  | 0.01 | 0.00916 | -8.41
[9 empty] | 19 | 20201123_PFAS_PPB_UC_Sample019 | CC12 - 9.308 pg/uL | Standard | 2.49 | 19826.656 | 531010 | 19826.656 |  |  | 0.01 | 0.00911 | -8.89
[9 empty] | 91 | 20201123_PFAS_PPB_UC_Sample091 | CC12 - 9.308 pg/uL | Standard | 2.49 | 21196.979 | 574966 | 21196.979 |  |  | 0.01
[9 empty] | 21 | 20201123_PFAS_PPB_UC_Sample021 | CC13 - 14.893 pg/uL | Standard | 2.49 | 19678.5 | 525681 | 19678.5 |  |  | 0.01 | 0.00964 | -3.62
[9 empty] | 93 | 20201123_PFAS_PPB_UC_Sample093 | CC13 - 14.893 pg/uL | Standard | 2.49 | 20941.443 | 568394 | 20941.443 |  |  | 0.01 | 0.00927 | -7.27
[9 empty] | 22 | 20201123_PFAS_PPB_UC_Sample022 | CC14 - 23.828 pg/uL | Standard | 2.49 | 18895.691 | 511302 | 18895.691 |  |  | 0.01 | 0.00902 | -9.78
[9 empty] | 94 | 20201123_PFAS_PPB_UC_Sample094 | CC14 - 23.828 pg/uL | Standard | 2.49 | 20097.592 | 543063 | 20097.592 |  |  | 0.01 | 0.00933 | -6.71
[9 empty] | 23 | 20201123_PFAS_PPB_UC_Sample023 | CC15 - 38.125 pg/uL | Standard | 2.49 | 18753.004 | 501592 | 18753.004 |  |  | 0.01 | 0.0099 | -0.95
[9 empty] | 95 | 20201123_PFAS_PPB_UC_Sample095 | CC15 - 38.125 pg/uL | Standard | 2.49 | 20623.277 | 568310 | 20623.277 |  |  | 0.01 | 0.00974 | -2.65
[9 empty] | 24 | 20201123_PFAS_PPB_UC_Sample024 | CC16 - 47.656 pg/uL | Standard | 2.49 | 19034.314 | 503536 | 19034.314 |  |  | 0.01 | 0.00956 | -4.38
[9 empty] | 96 | 20201123_PFAS_PPB_UC_Sample096 | CC16 - 47.656 pg/uL | Standard | 2.49 | 19677.9 | 548965 | 19677.9 |  |  | 0.01 | 0.00966 | -3.38
[9 empty] | 25 | 20201123_PFAS_PPB_UC_Sample025 | CC17 - 76.250 pg/uL | Standard | 2.49 | 19099.988 | 510694 | 19099.988 |  |  | 0.01 | 0.00928 | -7.22
[9 empty] | 97 | 20201123_PFAS_PPB_UC_Sample097 | CC17 - 76.250 pg/uL | Standard | 2.49 | 20306.045 | 546880 | 20306.045 |  |  | 0.01 | 0.00921 | -7.92
[9 empty] | 4 | 20201123_PFAS_PPB_UC_Sample004 | Crash Blank | Blank | [6 empty] | 0.01 | 0.00969 | -3.09
[9 empty] | 13 | 20201123_PFAS_PPB_UC_Sample013 | Crash Blank | Blank | [6 empty] | 0.01
[9 empty] | 66 | 20201123_PFAS_PPB_UC_Sample066 | Crash Blank | Blank | [6 empty] | 0.01 | 0.01119 | 11.91
[9 empty] | 85 | 20201123_PFAS_PPB_UC_Sample085 | Crash Blank | Blank | [6 empty] | 0.01 | 0.00986 | -1.39
[9 empty] | 5 | 20201123_PFAS_PPB_UC_Sample005 | Crash Mixed Matrix Blank | Blank | 2.49 | 18857.293 | 505995 | 18857.293 |  |  | 0.01 | 0.00987 | -1.31
[9 empty] | 20 | 20201123_PFAS_PPB_UC_Sample020 | Crash Mixed Matrix Blank | Blank | 2.49 | 19474.645 | 532376 | 19474.645 |  |  | 0.01 | 0.0097 | -3.03
[9 empty] | 38 | 20201123_PFAS_PPB_UC_Sample038 | Crash Mixed Matrix Blank | Blank | 2.49 | 19650.477 | 527645 | 19650.477 |  |  | 0.01 | 0.00978 | -2.19
[9 empty] | 52 | 20201123_PFAS_PPB_UC_Sample052 | Crash Mixed Matrix Blank | Blank | 2.49 | 20304.375 | 543610 | 20304.375 |  |  | 0.01 | 0.00951 | -4.88
[9 empty] | 73 | 20201123_PFAS_PPB_UC_Sample073 | Crash Mixed Matrix Blank | Blank | 2.49 | 20612.568 | 553657 | 20612.568 |  |  | 0.01 | 0.01012 | 1.21
[9 empty] | 77 | 20201123_PFAS_PPB_UC_Sample077 | Crash Mixed Matrix Blank | Blank | 2.49 | 20241.559 | 560304 | 20241.559 |  |  | 0.01 | 0.01021 | 2.13
[9 empty] | 92 | 20201123_PFAS_PPB_UC_Sample092 | Crash Mixed Matrix Blank | Blank | 2.49 | 21380.4 | 582255 | 21380.4 |  |  | 0.01 | 0.00988 | -1.16
[9 empty] | 104 | 20201123_PFAS_PPB_UC_Sample104 | Crash Mixed Matrix Blank | Blank | 2.49 | 21765.705 | 595826 | 21765.705 |  |  | 0.01 | 0.00997 | -0.27
[9 empty] | 111 | 20201123_PFAS_PPB_UC_Sample111 | Crash Mixed Matrix Blank | Blank | 2.49 | 7133.3 | 196191 | 7133.3 |  |  | 0.01 | 0.00994 | -0.61
[9 empty] | 27 | 20201123_PFAS_PPB_UC_Sample027 | QC-A | QC | 2.49 | 19263.697 | 517174 | 19263.697 |  |  | 0.01
[9 empty] | 99 | 20201123_PFAS_PPB_UC_Sample099 | QC-A | QC | 2.49 | 20563.285 | 569601 | 20563.285 |  |  | 0.01 | 0.00983 | -1.66
[9 empty] | 28 | 20201123_PFAS_PPB_UC_Sample028 | QC-B | QC | 2.49 | 20758.027 | 568945 | 20758.027 |  |  | 0.01 | 0.01057 | 5.72
[9 empty] | 100 | 20201123_PFAS_PPB_UC_Sample100 | QC-B | QC | 2.49 | 24810.557 | 681719 | 24810.557 |  |  | 0.01 | 0.01031 | 3.13
[9 empty] | 29 | 20201123_PFAS_PPB_UC_Sample029 | QC-C | QC | 2.49 | 20168.148 | 541891 | 20168.148 |  |  | 0.01 | 0.01049 | 4.86
[9 empty] | 101 | 20201123_PFAS_PPB_UC_Sample101 | QC-C | QC | 2.49 | 21870.688 | 595097 | 21870.688 |  |  | 0.01 | 0.00999 | -0.11
[9 empty] | 30 | 20201123_PFAS_PPB_UC_Sample030 | QC-D | QC | 2.49 | 19404.105 | 532543 | 19404.105 |  |  | 0.01 | 0.00997 | -0.31
[9 empty] | 102 | 20201123_PFAS_PPB_UC_Sample102 | QC-D | QC | 2.49 | 21461.814 | 589347 | 21461.814 |  |  | 0.01
[9 empty] | 1 | 20201123_PFAS_PPB_UC_Sample001 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00996 | -0.35
[9 empty] | 2 | 20201123_PFAS_PPB_UC_Sample002 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00974 | -2.62
[9 empty] | 3 | 20201123_PFAS_PPB_UC_Sample003 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00983 | -1.75
[9 empty] | 6 | 20201123_PFAS_PPB_UC_Sample006 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00992 | -0.75
[9 empty] | 26 | 20201123_PFAS_PPB_UC_Sample026 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.01063 | 6.26
[9 empty] | 31 | 20201123_PFAS_PPB_UC_Sample031 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.01041 | 4.08
[9 empty] | 45 | 20201123_PFAS_PPB_UC_Sample045 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.0105 | 4.98
[9 empty] | 59 | 20201123_PFAS_PPB_UC_Sample059 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.01028 | 2.83
[9 empty] | 78 | 20201123_PFAS_PPB_UC_Sample078 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00987 | -1.32
[9 empty] | 98 | 20201123_PFAS_PPB_UC_Sample098 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.01013 | 1.26
[9 empty] | 103 | 20201123_PFAS_PPB_UC_Sample103 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00966 | -3.38
[9 empty] | 118 | 20201123_PFAS_PPB_UC_Sample118 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.00997 | -0.29
[9 empty] | 125 | 20201123_PFAS_PPB_UC_Sample125 | Solvent Blank | Solvent | [6 empty] | 0.01
[9 empty] | 126 | 20201123_PFAS_PPB_UC_Sample126 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.0101 | 0.97
[9 empty] | 127 | 20201123_PFAS_PPB_UC_Sample127 | Solvent Blank | Solvent | [6 empty] | 0.01 | 0.01218 | 21.82
[9 empty] | 46 | 20201123_PFAS_PPB_UC_Sample046 | UC_AF_Mix1_A | Analyte | 2.49 | 21697.566 | 586996 | 21697.566 |  |  | 0.01 | 0.01074 | 7.39
[9 empty] | 47 | 20201123_PFAS_PPB_UC_Sample047 | UC_AF_Mix1_B | Analyte | 2.49 | 21829.346 | 586073 | 21829.346 |  |  | 0.01 | 0.01054 | 5.38
[9 empty] | 48 | 20201123_PFAS_PPB_UC_Sample048 | UC_AF_Mix1_C | Analyte | 2.49 | 20831.814 | 561981 | 20831.814 |  |  | 0.01
[9 empty] | 49 | 20201123_PFAS_PPB_UC_Sample049 | UC_AF_Mix2_A | Analyte | 2.49 | 21105.879 | 570134 | 21105.879 |  |  | 0.01 | 0.01069 | 6.87
[9 empty] | 50 | 20201123_PFAS_PPB_UC_Sample050 | UC_AF_Mix2_B | Analyte | 2.49 | 20513.939 | 559055 | 20513.939 |  |  | 0.01 | 0.01042 | 4.17
[9 empty] | 51 | 20201123_PFAS_PPB_UC_Sample051 | UC_AF_Mix2_C | Analyte | 2.49 | 20742.979 | 560428 | 20742.979 |  |  | 0.01 | 0.01069 | 6.89
[9 empty] | 53 | 20201123_PFAS_PPB_UC_Sample053 | UC_AF_Mix3_A | Analyte | 2.49 | 21071.273 | 570517 | 21071.273 |  |  | 0.01 | 0.01067 | 6.71
[9 empty] | 54 | 20201123_PFAS_PPB_UC_Sample054 | UC_AF_Mix3_B | Analyte | 2.49 | 20743.402 | 573332 | 20743.402 |  |  | 0.01 | 0.0104 | 3.96
[9 empty] | 55 | 20201123_PFAS_PPB_UC_Sample055 | UC_AF_Mix3_C | Analyte | 2.49 | 20386.736 | 551776 | 20386.736 |  |  | 0.01 | 0.01036 | 3.59
[9 empty] | 56 | 20201123_PFAS_PPB_UC_Sample056 | UC_T1hr_Mix1_A | Analyte | 2.49 | 19113.52 | 511307 | 19113.52 |  |  | 0.01 | 0.0102 | 2
[9 empty] | 57 | 20201123_PFAS_PPB_UC_Sample057 | UC_T1hr_Mix1_B | Analyte | 2.49 | 19075.039 | 510117 | 19075.039 |  |  | 0.01 | 0.0035 | -64.97
[9 empty] | 58 | 20201123_PFAS_PPB_UC_Sample058 | UC_T1hr_Mix1_C | Analyte | 2.49 | 19599.061 | 534441 | 19599.061 |  |  | 0.01 | 0.01013 | 1.26
[9 empty] | 60 | 20201123_PFAS_PPB_UC_Sample060 | UC_T1hr_Mix2_A | Analyte | 2.49 | 18590.02 | 500460 | 18590.02 |  |  | 0.01 | 0.01094 | 9.39
[9 empty] | 61 | 20201123_PFAS_PPB_UC_Sample061 | UC_T1hr_Mix2_B | Analyte | 2.49 | 20187.695 | 544215 | 20187.695 |  |  | 0.01 | 0.01085 | 8.55
[9 empty] | 62 | 20201123_PFAS_PPB_UC_Sample062 | UC_T1hr_Mix2_C | Analyte | 2.49 | 19809.779 | 537260 | 19809.779 |  |  | 0.01 | 0.01105 | 10.52
[9 empty] | 63 | 20201123_PFAS_PPB_UC_Sample063 | UC_T1hr_Mix3_A | Analyte | 2.49 | 19047.59 | 510104 | 19047.59 |  |  | 0.01 | 0.01052 | 5.22
[9 empty] | 64 | 20201123_PFAS_PPB_UC_Sample064 | UC_T1hr_Mix3_B | Analyte | 2.49 | 19612.295 | 525660 | 19612.295 |  |  | 0.01 | 0.0104 | 3.96
[9 empty] | 65 | 20201123_PFAS_PPB_UC_Sample065 | UC_T1hr_Mix3_C | Analyte | 2.49 | 19736.887 | 531419 | 19736.887 |  |  | 0.01
[9 empty] | 67 | 20201123_PFAS_PPB_UC_Sample067 | UC_T5hr_Mix1_A | Analyte | 2.49 | 22790.889 | 620955 | 22790.889 |  |  | 0.01 | 0.01023 | 2.32
[9 empty] | 68 | 20201123_PFAS_PPB_UC_Sample068 | UC_T5hr_Mix1_B | Analyte | 2.49 | 20082.863 | 540143 | 20082.863 |  |  | 0.01 | 0.01094 | 9.38
[9 empty] | 69 | 20201123_PFAS_PPB_UC_Sample069 | UC_T5hr_Mix1_C | Analyte | 2.49 | 20098.73 | 550011 | 20098.73 |  |  | 0.01 | 0.01113 | 11.3
[9 empty] | 70 | 20201123_PFAS_PPB_UC_Sample070 | UC_T5hr_Mix2_A | Analyte | 2.49 | 19749.709 | 542271 | 19749.709 |  |  | 0.01 | 0.01061 | 6.08
[9 empty] | 71 | 20201123_PFAS_PPB_UC_Sample071 | UC_T5hr_Mix2_B | Analyte | 2.49 | 19920.66 | 541444 | 19920.66 |  |  | 0.01 | 0.01077 | 7.7
[9 empty] | 72 | 20201123_PFAS_PPB_UC_Sample072 | UC_T5hr_Mix2_C | Analyte | 2.49 | 19372.84 | 527096 | 19372.84 |  |  | 0.01 | 0.0105 | 5.03
[9 empty] | 74 | 20201123_PFAS_PPB_UC_Sample074 | UC_T5hr_Mix3_A | Analyte | 2.49 | 20799.916 | 564585 | 20799.916 |  |  | 0.01
[9 empty] | 75 | 20201123_PFAS_PPB_UC_Sample075 | UC_T5hr_Mix3_B | Analyte | 2.49 | 20129.732 | 544708 | 20129.732 |  |  | 0.01
[9 empty] | 76 | 20201123_PFAS_PPB_UC_Sample076 | UC_T5hr_Mix3_C | Analyte | 2.49 | 20311.908 | 546886 | 20311.908 |  |  | 0.01